ABL80 CO-OX | pO2 high on system cycle. (93 on Solution 3, range 27-77). Moved SC to other analyzer, reading 63. New SC doesn't pass external QC. Replaced SP. SC on first analyzer won't cal due to SP. Replaced SP. SC still low on level 1 QC. Second SC installed. Customer needs replacement of two SP's and SC's as they now have none. Numerous QC's used during troubleshooting and need replaced. | New SC and SP corrected problem. SC, SP and QC on order. | D13D47FB52246139 | No | pO2 high on system cycle. (93 on Solution 3, range 27-77). Moved SC to other analyzer, reading 63. New SC doesn't pass external QC. Replaced SP. SC on first analyzer won't cal due to SP. Replaced SP. SC still low on level 1 QC. Second SC installed. Customer needs replacement of two SP's and SC's as they now have none. Numerous QC's used during troubleshooting and need replaced. | New SC and SP corrected problem. SC, SP and QC on order.
ABL90 | Found Inlet Gasket not installed correctly. Pulled main module Inspected fluid path for leaks. No leaks found. Located loose 25-pole WSCPCB to SensorInterface PCB cable. | Secured cable to Wet Section Bd. | D13CFA6D5259862E | No | Found Inlet Gasket not installed correctly. Pulled main module Inspected fluid path for leaks. No leaks found. Located loose 25-pole WSCPCB to SensorInterface PCB cable. | Secured cable to Wet Section Bd. | User error | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Abgebrochene Einlassdichtung aus Halterung gezogen. Neue Einlassdichtung und Einlasssonde eingesetzt. Einlassanschlussdichtung erneuert, da verstopft. Funktion Einlassmodul o. K. Durch längeres Ausschalten hat Sensorkassette Schaden genommen. Die Metabolite haben Kalibrierfehler. Kalibrationen und Qualitätskontrolle o. K. | Abgebrochene Einlassdichtung aus Halterung gezogen. Neue Einlassdichtung und Einlasssonde eingesetzt. Einlassanschlussdichtung erneuert, da verstopft. Funktion Einlassmodul o. K. Durch längeres Ausschalten hat Sensorkassette Schaden genommen. Die Metabolite haben Kalibrierfehler. Kalibrationen und Qualitätskontrolle o. K. | SWO-125835 | No | Aborted inlet seal pulled from bracket. New inlet seal and inlet probe inserted. Inlet connection seal renewed, there clogged. Function Inlet module o. K. The sensor cartridge has been damaged by prolonged switching off. The metabolites have calibration errors. Calibrations and quality control o.K. | Aborted inlet seal pulled from bracket. New inlet seal and inlet probe inserted. Inlet connection seal renewed, there clogged. Function Inlet module o. K. The sensor cartridge has been damaged by prolonged switching off. The metabolites have calibration errors. Calibrations and quality control o.K. | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | Nettoyage de la section fluidique e.o Remplacé le kit de maintenance e.o Vérification valve et hémolyseur de l'oxymétrie e.o Calibration e.o contrôle de qualité e.o | Nettoyage de la section fluidique e.o Remplacé le kit de maintenance e.o Vérification valve et hémolyseur de l'oxymétrie e.o Calibration e.o contrôle de qualité e.o | SWO-125797 | No | Cleaning the fluidic section e.o Replaced the maintenance kit e.o Checking valve and hemolyseur of the oxymetry e.o Calibration e.o quality control e.o | Cleaning the fluidic section e.o Replaced the maintenance kit e.o Checking valve and hemolyseur of the oxymetry e.o Calibration e.o quality control e.o | Defect | HW | PM Kit | PM Kit
ABL90 FLEX PLUS | Removed wet section and flushed, clearing blockages in small sections. Replaced inlet probe. Performed fluid transport checks and tests. Replaced sensor cassette. Analyser is ready for use. | Removed wet section and flushed, clearing blockages in small sections. Replaced inlet probe. Performed fluid transport checks and tests. Replaced sensor cassette. Analyser is ready for use. | SWO-125972 | No | Removed wet section and flushed, clearing blockages in small sections. Replaced inlet probe. Performed fluid transport checks and tests. Replaced sensor cassette. Analyser is ready for use. | Removed wet section and flushed, clearing blockages in small sections. Replaced inlet probe. Performed fluid transport checks and tests. Replaced sensor cassette. Analyser is ready for use. | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX PLUS | Einlassmodul im Service entfernt. Einlasssonde ist in Dichtung blockiert. Lässt sich nur unter Kraftaufwand betätigen. Einlasssonde mit Blut verstopft, erneuert auch Einlassdichtung. Einlassmodul neu kalibriert. Anmeldungen durch Anwender gescheitert. In letzter Zeit sehr offt, aber immer mal wieder. Kalibrationen und Qualitätskontrolle o. K. | Einlassmodul im Service entfernt. Einlasssonde ist in Dichtung blockiert. Lässt sich nur unter Kraftaufwand betätigen. Einlasssonde mit Blut verstopft, erneuert auch Einlassdichtung. Einlassmodul neu kalibriert. Anmeldungen durch Anwender gescheitert. In letzter Zeit sehr offt, aber immer mal wieder. Kalibrationen und Qualitätskontrolle o. K. | SWO-125837 | No | Removed the inlet module in the service. Inlet probe is blocked in seal. Can only be operated under force. Inlet probe clogged with blood, also renews inlet seal. inlet module recalibrated. Logins by users failed. Lately very high, but again and again. Calibrations and quality control o.K. | Removed the inlet module in the service. Inlet probe is blocked in seal. Can only be operated under force. Inlet probe clogged with blood, also renews inlet seal. inlet module recalibrated. Logins by users failed. Lately very high, but again and again. Calibrations and quality control o.K. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
TCM CombiM |  | B | SWO-125623 | No
TCM4 |  | B | SWO-125621 | No
AQT90 FLEX |  | 更换了910-343后未修复，需要再借件电源板 | SWO-125666 | No |  | Unrepaired after replacement of 910-343, need to borrow a second power board | Revisit required | HW | Power supply | Power supply
AQT90 FLEX | 更换了910-343后未修复，需要再借件电源板 | 更换了910-343后未修复，需要再借件电源板 | SWO-125294 | No | Unrepaired after replacement of 910-343, need to borrow a second power board | Unrepaired after replacement of 910-343, need to borrow a second power board | Defect | HW | Power supply | Power supply
ABL800 BASIC | cleaned all PCB and checked the valves, all are working satisfactorily, now the analyzer working satisfactorily. | cleaned all PCB and checked the valves, all are working satisfactorily, now the analyzer working satisfactorily. | SWO-126808 | No | cleaned all PCB and checked the valves, all are working satisfactorily, now the analyzer working satisfactorily. | cleaned all PCB and checked the valves, all are working satisfactorily, now the analyzer working satisfactorily. | Cleaned | HW | Inlet module | Pre heater
ABL815 FLEX | Startphase QC und Kalibration. | Startphase QC und Kalibration. | SWO-125959 | No | Start phase QC and calibration. | Start phase QC and calibration. | Cleaned | HW | AutoCheck Module | Barcode reader
ABL800 BASIC | Replaced inlet probe, and done inlet calibration -ok. Cleaned measuring chamber and electrode jacket of pH electrode and ran deprotenizer x2. Replaced lactate membrane. | Replaced inlet probe, and done inlet calibration -ok. Cleaned measuring chamber and electrode jacket of pH electrode and ran deprotenizer x2. Replaced lactate membrane. | SWO-125876 | No | Replaced inlet probe, and done inlet calibration -ok. Cleaned measuring chamber and electrode jacket of pH electrode and ran deprotenizer x2. Replaced lactate membrane. | Replaced inlet probe, and done inlet calibration -ok. Cleaned measuring chamber and electrode jacket of pH electrode and ran deprotenizer x2. Replaced lactate membrane. | Defect | HW | Inlet module | Probe bend
ABL835 FLEX | New upper sealing ring fitted to pCO2 electrode. New reference membrane fitted to analyser, (the current one had been installed for 21 days). | New upper sealing ring fitted to pCO2 electrode. New reference membrane fitted to analyser, (the current one had been installed for 21 days). | SWO-125842 | No | New upper sealing ring fitted to pCO2 electrode. New reference membrane fitted to analyser, (the current one had been installed for 21 days). | New upper sealing ring fitted to pCO2 electrode. New reference membrane fitted to analyser, (the current one had been installed for 21 days). | Assisted Replacement | Electrode | Membrane | pCO2 | Sealing Ring
ABL800 BASIC | Visited the concern dept, found room temperature was high, AC was not working, meet the Biochemistry HOD, she told that they will repair the with in a hour, they tried but it was not working, it will take a month time, she told that next month they will call, Shout down the analyzer for time been. | Visited the concern dept, found room temperature was high, AC was not working, meet the Biochemistry HOD, she told that they will repair the with in a hour, they tried but it was not working, it will take a month time, she told that next month they will call, Shout down the analyzer for time been. | SWO-126822 | No | Visited the concern dept, found room temperature was high, AC was not working, meet the Biochemistry HOD, she told that they will repair the with in a hour, they tried but it was not working, it will take a month time, she told that next month they will call , Shout down the analyzer for time been. | Visited the concern dept, found room temperature was high, AC was not working, meet the Biochemistry HOD, she told that they will repair the with in a hour, they tried but it was not working, it will take a month time, she told that next month they will call , Shout down the analyzer for time been. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX |  | B | D13CD269524FA612 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Messkammer und Schlauchsystem geprüft und gereinigt, Ca. Elektrode neu membraniert, Funktionskontrolle ok | Messkammer und Schlauchsystem geprüft und gereinigt, Ca. Elektrode neu membraniert, Funktionskontrolle ok | SWO-125935 | No | Measuring chamber and hose system tested and cleaned, approx. electrode re-membraneed, function control ok | Measuring chamber and hose system tested and cleaned, approx. electrode re-membraneed, function control ok | Cleaned | Fluidic pathway | El/Met Module | Measuring chamber
ABL80 CO-OX | installed new SC and SP | performed zero cal and tested operation of analyzer. Original SC being shipped back to RMED as per Techline instructions. Customer is going to run qc and do a patient comparrison between the 80 and the ABl 8xx units at sunnybrook. customer will send me the results after a week of data colection | D13E0559C83B7CB9 | No | installed new SC and SP | performed zero cal and tested operation of analyzer. Original SC being shipped back to RMED as per Techline instructions. Customer is going to run qc and do a patient comparrison between the 80 and the ABl 8xx units at sunnybrook. customer will send me the results after a week of data colection
ABL80 CO-OX | Current version is 1.60. | Installed software version 3.14. Tested the valves and liquid sensor. I believe the fluid blockage problem was with the previous SP. | D13E893CC814057B | No | Current version is 1.60. | Installed software version 3.14. Tested the valves and liquid sensor. I believe the fluid blockage problem was with the previous SP.
ABL90 FLEX | Bezeichnung des Analysator vorhanden unter Analysator ID. Eine ältere gesicherte ASTM.ini eingespielt und gesichert. Weiterhin bleibt es bei der Kennung. Roche Mitarbeiter im Urlaub. Weiteres Vorgehen zu späterem Zeitpunkt klärbar. Kalibrationen und Qualitätskontrolle o. K. | Bezeichnung des Analysator vorhanden unter Analysator ID. Eine ältere gesicherte ASTM.ini eingespielt und gesichert. Weiterhin bleibt es bei der Kennung. Roche Mitarbeiter im Urlaub. Weiteres Vorgehen zu späterem Zeitpunkt klärbar. Kalibrationen und Qualitätskontrolle o. K. | SWO-126060 | No | Name of the analyzer exists under Analyzer ID. An older secured ASTM.ini is saved and secured. Furthermore, it remains with the identifier. Roche employees on vacation. Further action can be clarified at a later date. Calibrations and quality control o.K. | Name of the analyzer exists under Analyzer ID. An older secured ASTM.ini is saved and secured. Furthermore, it remains with the identifier. Roche employees on vacation. Further action can be clarified at a later date. Calibrations and quality control o.K. | Unknown | Revisit required | His/lis | His/lis | Roche
ABL90 FLEX | Kanäle mit dem Sensorsimulator überprüft, Leihgerät 090R0795N001 aufgestellt, EDV Anbindung geprüft | Kanäle mit dem Sensorsimulator überprüft, Leihgerät 090R0795N001 aufgestellt, EDV Anbindung geprüft | SWO-126020 | No | Channels checked with the sensor simulator, rental device 090R0795N001 set up, computer connection tested | Channels checked with the sensor simulator, rental device 090R0795N001 set up, computer connection tested | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | Tausch der Einlassdichtung und der Einlasssonde, da krum. Fließwege gespült, Keine Leckage erkannt. Fehler durch Verwendung von Fremdspritzen, Mikrogerinsel vorhanden. Analysator wieder Ohne Beanstandung nach Durchführung seiner Routinen. Da keine Einlassdichtung auf Lager, auf Wunsch der Anwender eine da gelassen. Unterer berichtbarer Bereich für Ca auf 0,1 mmol l, auf Kundenwunsch eingestellt und in Roma.ini programmiert. Kalibrationen und Qualitätskontrolle o. K. | Tausch der Einlassdichtung und der Einlasssonde, da krum. Fließwege gespült, Keine Leckage erkannt. Fehler durch Verwendung von Fremdspritzen, Mikrogerinsel vorhanden. Analysator wieder Ohne Beanstandung nach Durchführung seiner Routinen. Da keine Einlassdichtung auf Lager, auf Wunsch der Anwender eine da gelassen. Unterer berichtbarer Bereich für Ca auf 0,1 mmol l, auf Kundenwunsch eingestellt und in Roma.ini programmiert. Kalibrationen und Qualitätskontrolle o. K. | SWO-126118 | No | Replacement of the inlet seal and the inlet probe, since krum. Flow paths rinsed, no leakage detected. Error due to use of foreign syringes, microger island present. Analyzer again without complaint after carrying out his routines. Since no inlet seal in stock, at the request of the user one left there. Lower reportable range for Ca at 0.1 mmol l, adjusted to customer request and programmed in Roma.ini. Calibrations and quality control o.K. | Replacement of the inlet seal and the inlet probe, since krum. Flow paths rinsed, no leakage detected. Error due to use of foreign syringes, microger island present. Analyzer again without complaint after carrying out his routines. Since no inlet seal in stock, at the request of the user one left there. Lower reportable range for Ca at 0.1 mmol l, adjusted to customer request and programmed in Roma.ini. Calibrations and quality control o.K. | Replaced | HW | Inlet module | Probe bend
ABL90 FLEX PLUS | Treiber für Etikettendrucker neu installiert und konfiguriert, Patientenlayout auf Betriebssystemebene angepasst, Funktionskontrolle. | Treiber für Etikettendrucker neu installiert und konfiguriert, Patientenlayout auf Betriebssystemebene angepasst, Funktionskontrolle. | SWO-126014 | No | Drivers for label printers reinstalled and configured, patient layout customized at the operating system level, functional control. | Drivers for label printers reinstalled and configured, patient layout customized at the operating system level, functional control. | Setup | SW | Ext printer | Ext printer
ABL90 FLEX PLUS |  | B | D13E06BAC8607996 | No |  |  | Defect | HW | Pre heater | Pre heater
ABL825 FLEX |  | B | D13E3CD9C81A5C0D | Yes |  |  | Re classified RMED | Phone call
ABL825 FLEX | Error 0579: Halogen Lamp intensity too high Réglage de l'intensité de la lampe halogène e.o Test complet de l'oxymétrie e.o Calibration e.o CQ e.o | Error 0579: Halogen Lamp intensity too high Réglage de l'intensité de la lampe halogène e.o Test complet de l'oxymétrie e.o Calibration e.o CQ e.o | SWO-126262 | No | Error 0579: Halogen Lamp intensity too high Setting the intensity of the halogen lamp e.o Complete oxymetry test e.o Calibration e.o CQ e.o | Error 0579: Halogen Lamp intensity too high Setting the intensity of the halogen lamp e.o Complete oxymetry test e.o Calibration e.o CQ e.o | Adjustment | HW | Oxi module | Lamp unit
ABL805 FLEX | Flüssigkeitswege gespült ersetzt, tropft weiter, Friction detection Unit ersetzt, Neustart, Kalibration, pCO2 Empf. nur noch 70% daher neu membraniert...... , Kalibration 102,4% Kalibrationen.....,QC-Kontrolle, pO2 Fehler Dichtung wg. Verfärbung ersetzt, Neustart Taste steht nicht zur verfügung, Analysator neu gestartet......neuer AC-Deckel wird bestellt. pCO2 QC Verletzung, Kalibrationen.... pCO2 Empfindlichkeit bei 82,5% !!! Ein paar Kalibrationen späterQC-bestanden | Flüssigkeitswege gespült ersetzt, tropft weiter, Friction detection Unit ersetzt, Neustart, Kalibration, pCO2 Empf. nur noch 70% daher neu membraniert...... , Kalibration 102,4% Kalibrationen.....,QC-Kontrolle, pO2 Fehler Dichtung wg. Verfärbung ersetzt, Neustart Taste steht nicht zur verfügung, Analysator neu gestartet......neuer AC-Deckel wird bestellt. pCO2 QC Verletzung, Kalibrationen.... pCO2 Empfindlichkeit bei 82,5% !!! Ein paar Kalibrationen späterQC-bestanden | SWO-126094 | No | Liquid pathways rinsed replaced, further drips, Friction detection unit replaced, restart, calibration, pCO2 empf. only 70% therefore re-membraneed...... , calibration 102.4% calibrations.....,QC control, pO2 error seal wg. Discoloration replaced, restart button is not available, analyzer restarted...... new AC lid is ordered. pCO2 QC Violation, Calibrations.... pCO2 Sensitivity at 82.5% !!! A few calibrations laterQC passed | Liquid pathways rinsed replaced, further drips, Friction detection unit replaced, restart, calibration, pCO2 empf. only 70% therefore re-membraneed...... , calibration 102.4% calibrations.....,QC control, pO2 error seal wg. Discoloration replaced, restart button is not available, analyzer restarted...... new AC lid is ordered. pCO2 QC Violation, Calibrations.... pCO2 Sensitivity at 82.5% !!! A few calibrations laterQC passed | Defect | HW | Inlet module | Friction detector unit
ABL90 FLEX | Bei eintreffen war Gerät grün, alle QC i.O. (im oberen QC Bereich aber nicht grenzwertig), Kunde hat Lüfter vor das Gerät gestellt, mit Fr, Beyer die aktuellen SC & LP Problematik besprochen, QC Plan geändert (Level C wird jetzt 03.30Uhr gemessen ->vorher 14.00Uhr), | Bei eintreffen war Gerät grün, alle QC i.O. (im oberen QC Bereich aber nicht grenzwertig), Kunde hat Lüfter vor das Gerät gestellt, mit Fr, Beyer die aktuellen SC & LP Problematik besprochen, QC Plan geändert (Level C wird jetzt 03.30Uhr gemessen ->vorher 14.00Uhr), | SWO-126344 | No | When arrived device was green, all QC i.O. (in the upper QC range but not borderline), customer has put fan in front of the device, with Fr, Beyer discussed the current SC & LP problem, QC plan changed (Level C is now measured 3.30am ->before 14.00o), | When arrived device was green, all QC i.O. (in the upper QC range but not borderline), customer has put fan in front of the device, with Fr, Beyer discussed the current SC & LP problem, QC plan changed (Level C is now measured 3.30am ->before 14.00o), | Assisted Replacement | SP | SP | pCO2
ABL90 | BATTERY INSTALLED FAILED TO SWITCH BACK TO OUTLET POWER, ONCE THE BATTERY DIED THE SYSTEM COULD NOT BE RESTARTED. | AFTER 5 BATTERY INSTALLS THE 6TH WORKED NORMALLY | D1404116B41B40A2 | No | BATTERY INSTALLED FAILED TO SWITCH BACK TO OUTLET POWER, ONCE THE BATTERY DIED THE SYSTEM COULD NOT BE RESTARTED. | AFTER 5 BATTERY INSTALLS THE 6TH WORKED NORMALLY | Defect | HW | Battery module | Battery module
ABL90 FLEX PLUS | Hemolyzer und Spectrometer ersetzt, Neustart, Kalibration, tHb Kalibration, QC- Kontrolle. | Hemolyzer und Spectrometer ersetzt, Neustart, Kalibration, tHb Kalibration, QC- Kontrolle. | SWO-126574 | No | Hemolyzer and Spectrometer replaced, restart, calibration, tHb calibration, QC control. | Hemolyzer and Spectrometer replaced, restart, calibration, tHb calibration, QC control. | Defect | HW | Oxi module | Spectrometer
ABL90 FLEX | 更换测试卡和试剂包修复 | 更换测试卡和试剂包修复 | SWO-126644 | No | Replacement test card and reagent kit repair | Replacement test card and reagent kit repair
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-126173 | No | Repair after replacing a new test card | Repair after replacing a new test card
AQT90 FLEX | 检查机器，发现前面板盖上后。维修菜单中打开和关闭有时不能正常动作，由于近期经常发生此故障，重启后能正常使用，但是不定期出现此故障，所以更换903-224，903-028，580-404 后目前正常。byjesson | 检查机器，发现前面板盖上后。维修菜单中打开和关闭有时不能正常动作，由于近期经常发生此故障，重启后能正常使用，但是不定期出现此故障，所以更换903-224，903-028，580-404 后目前正常。byjesson | SWO-125994 | No | Check the machine and find the front panel cover behind. The service menu openands and closes sometimes can not normal lybecause of this recent failure, can be used normally after restart, but this fault occurs from time to time, so replace 903-224, 903-028, 580-404 is currently normal. byjesson | Check the machine and find the front panel cover behind. The service menu openands and closes sometimes can not normal lybecause of this recent failure, can be used normally after restart, but this fault occurs from time to time, so replace 903-224, 903-028, 580-404 is currently normal. byjesson | Defect | HW | Cover and Chassis | Front cover
AQT90 FLEX | 取出损坏测试卡后解决 | 取出损坏测试卡后解决 | SWO-126520 | No | Resolve the damaged test card after removing it | Resolve the damaged test card after removing it | Defect | Consumables | Test Kit | Cartridge | 167 (email)
ABL825 FLEX | Customer remembraned electrodes and replaced the reference electrode but still had the same problems. Checked the reference electrode connector and electrode channels. All checked okay. | Replaced the pH electrode and pH electrode seal ring. Remembraned the lactate electrode. | D13F83AD3E3B5F51 | No | Customer remembraned electrodes and replaced the reference electrode but still had the same problems. Checked the reference electrode connector and electrode channels. All checked okay. | Replaced the pH electrode and pH electrode seal ring. Remembraned the lactate electrode. | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | checked data logs for errors, checked autocheck operation, | removed Autocheck and replaced FPC cable, carousel not rotating, replaced Autocheck board, installed board software, | D13FCE353E3FBE19 | No | checked data logs for errors, checked autocheck operation, | removed Autocheck and replaced FPC cable, carousel not rotating, replaced Autocheck board, installed board software, | Defect | HW | Auto Check Module | PCB, Autocheck
ABL825 FLEX | Ablenkspiegel gereinigt , Laserschacht ausgeblasen, kein Erfolg. AQC Modul ausgebaut und entstaubt. Scanner ausgebaut und gereinigt. Ref. Glu. und Lac. wegen Fälligkeit membraniert. Einlaufverfahren durchgeführt. Danach alle Funktionen o.k. | Ablenkspiegel gereinigt , Laserschacht ausgeblasen, kein Erfolg. AQC Modul ausgebaut und entstaubt. Scanner ausgebaut und gereinigt. Ref. Glu. und Lac. wegen Fälligkeit membraniert. Einlaufverfahren durchgeführt. Danach alle Funktionen o.k. | SWO-126536 | No | Deflection mirror cleaned, laser shaft blown out, no success. AQC module removed and dusted. Scanner removed and cleaned. Ref. Glu. and Lac. because of maturity. run-in procedures. After that, all functions ok. | Deflection mirror cleaned, laser shaft blown out, no success. AQC module removed and dusted. Scanner removed and cleaned. Ref. Glu. and Lac. because of maturity. run-in procedures. After that, all functions ok. | Cleaned | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | Inlet kalibriert und den Neustart zu Ende geführt inkl. 2. Kal. Gerät nochmals neu gestartet. Jetzt ist das Gerät wieder einsatzbereit. | Inlet kalibriert und den Neustart zu Ende geführt inkl. 2. Kal. Gerät nochmals neu gestartet. Jetzt ist das Gerät wieder einsatzbereit. | SWO-126286 | No | Inlet calibrated and the restart finished incl. 2nd cal. restarts the device. Now the device is ready for use again. | Inlet calibrated and the restart finished incl. 2nd cal. restarts the device. Now the device is ready for use again. | Defect | SW | Analyzer SW | Analyzer SW | 6.16
ABL825 FLEX | Looked at data logs and found error message 251, halogen lamp intensity too high. | Adjusted lamp intensity, Did complete oximetry, Did THB in service, in user mode and as sample. | D13F60373E4535E9 | No | Looked at data logs and found error message 251, halogen lamp intensity too high. | Adjusted lamp intensity, Did complete oximetry, Did THB in service, in user mode and as sample. | Adjustment | HW | Oxi module | Lamp unit | 251
ABL80 FLEX CO-OX | Kundengespräch und Geräteeinstellungen durchgeführt. | Kundengespräch und Geräteeinstellungen durchgeführt. | SWO-126652 | No | customer call and device settings. | customer call and device settings.
ABL90 FLEX |  | B | SWO-126626 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | 更换试剂包修复 | 更换试剂包修复 | SWO-126641 | No | Replacement reagent kit repair | Replacement reagent kit repair
ABL827 FLEX | Service Dump und notwendige Daten gezogen und an den Technischen Support übertragen. Weiteres Vorgehen nach Rückmeldung vom Technischen Support. Kalibrationen und Qualitätskontrolle in Ordnung. | Service Dump und notwendige Daten gezogen und an den Technischen Support übertragen. Weiteres Vorgehen nach Rückmeldung vom Technischen Support. Kalibrationen und Qualitätskontrolle in Ordnung. | SWO-126765 | No | Service dump and necessary data pulled and transferred to technical support. Further action after feedback from Technical Support. Calibrations and quality control in order. | Service dump and necessary data pulled and transferred to technical support. Further action after feedback from Technical Support. Calibrations and quality control in order. | Unknown | SW | Met II | Crea | No result
ABL805 FLEX | Elektrode erneuert und mehrfach kalibriert. | Elektrode erneuert und mehrfach kalibriert. | SWO-126754 | No | electrode renewed and calibrated several times. | electrode renewed and calibrated several times. | Assisted Replacement | Electrode | Electrode | pCO2
ABL815 FLEX | AC Modul aus GS Bestand erneuert. Einlassnadel erneuert. | AC Modul aus GS Bestand erneuert. Einlassnadel erneuert. | SWO-126752 | No | AC module from GS stock renewed. Inlet needle renewed. | AC module from GS stock renewed. Inlet needle renewed. | Defect | HW | AutoCheck Module | Breaker arm | Friction detection??
ABL815 FLEX | Bildschirm komplett ersetzt. K+ Komma stellen erweitert. Kalibriert und QC durchgeführt. Mit dem Labor die Datenübertragung überprüft. | B | SWO-126749 | No | completely replaced on the screen. K+ comma pose extended. Calibrated and QC performed. Check the data transfer with the laboratory. |  | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX | 更换测试卡，进样支架正常。 | 更换测试卡，进样支架正常。 | SWO-126815 | No | Replace the test card and the sample bracket is OK. | Replace the test card and the sample bracket is OK. | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-126916 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | pCO2
ABL800 BASIC | cleaned all fluidic valves. observed moisture on PH BG upper pre heater. now analyser is working properly. | cleaned all fluidic valves. observed moisture on PH BG upper pre heater. now analyser is working properly. | SWO-127546 | No | cleaned all fluidic valves. observed moisture on PH BG upper pre heater. now analyser is working properly. | cleaned all fluidic valves. observed moisture on PH BG upper pre heater. now analyser is working properly. | Defect | HW | pH/BG Module | Tubing
ABL825 FLEX | Überprüfung der Ventile in den Messmodulen. Membran des Air Ventils im pH BG Modul ist gerissen, erneuert. Funktion in Ordnung. LS Justage Fehler noch vorhanden, LS Justage und Pumpenkalibration durchgeführt. Fehler behoben. Mehrmals kalibriert, Kalibrationswerte in Ordnung. Qualitätskontrolle in Ordnung. | Überprüfung der Ventile in den Messmodulen. Membran des Air Ventils im pH BG Modul ist gerissen, erneuert. Funktion in Ordnung. LS Justage Fehler noch vorhanden, LS Justage und Pumpenkalibration durchgeführt. Fehler behoben. Mehrmals kalibriert, Kalibrationswerte in Ordnung. Qualitätskontrolle in Ordnung. | SWO-126830 | No | Checking the valves in the measuring modules. Membrane of the air valve in the pH BG module is torn, renewed. Function in order. LS adjustment error still present, LS adjustment and pump calibration performed. Fixed a bug. Calibrated several times, calibration values are fine. Quality control in order. | Checking the valves in the measuring modules. Membrane of the air valve in the pH BG module is torn, renewed. Function in order. LS adjustment error still present, LS adjustment and pump calibration performed. Fixed a bug. Calibrated several times, calibration values are fine. Quality control in order. | Defect | HW | pH/BG Module | Valve, air | 653
ABL80 FLEX BASIC | センサーカセット交換 | B | SWO-126879 | No | Sensor cassette replacement
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-126864 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | 更换新测试卡后解决 | 更换新测试卡后解决 | SWO-126865 | No | Resolve after replacing a new test card | Resolve after replacing a new test card | Assisted Replacement | SC | SC | Reference
AQT90 FLEX | 更换了扫描器910-304后修复 by jesson | 更换了扫描器910-304后修复 by jesson | SWO-126819 | No | Repair edgy jesson after replacing scanner 910-304 | Repair edgy jesson after replacing scanner 910-304 | Defect | HW | Storage Module | Barcode reader, Storage wheel
ABL800 BASIC | Taken the Inlet module out and checked the mechanical alignment. Adjusted the Inlet Gasket again and did the alignment through software. Now Did the calibration, which has been passed. Provided the Operator training to newly joined staff- Mr. Ashok and lady staff. | B | SWO-128488 | No | Taken the Inlet module out and checked the mechanical alignment. Adjusted the Inlet Gasket again and did the alignment through software. Now Did the calibration, which has been passed. Provided the Operator training to newly joined staff- Mr. Ashok and lady staff. |  | User error | HW | Inlet module | Inlet Gasket
ABL80 FLEX Basic version | Se retiraron los consumibles, se limpiaron los conectores y se reinstalaron los mismos consumibles. Se comprobó correcto funcionamiento mediante QC. | B | D16160E1B02C07DC | No | Consumables were removed, connectors cleaned and the same consumables reinstalled. Proper operation was checked by QC.
ABL80 FLEX BASIC | Replace the Sensor board with new one. Again try to install SC but still it shows the error. So there is a problem with Signal Data Board. | Replace the SC with new one. After replace the SC unit shows V C failed error. So replace the SP, after replace SP still it shows the V C error. So there is a possibilities of failure of SC. | SWO-127954 | No | Replace the Sensor board with new one. Again try to install SC but still it shows the error. So there is a problem with Signal Data Board. | Replace the SC with new one. After replace the SC unit shows V C failed error. So replace the SP, after replace SP still it shows the V C error. So there is a possibilities of failure of SC.
ABL80 FLEX CO-OX | Gerät ausgeliefert, meßbereit aufgestellt, Anbindung geprüft Leihgerät 304909 gereinigt und mitgenommen | Gerät ausgeliefert, meßbereit aufgestellt, Anbindung geprüft Leihgerät 304909 gereinigt und mitgenommen | SWO-127332 | No | Device delivered, ready to be measured, connection tested loaner 304909 cleaned and taken with you | Device delivered, ready to be measured, connection tested loaner 304909 cleaned and taken with you
ABL90 | Confirmed problem. | Replaced the tubing. | D1446F6D1660FC75 | No | Confirmed problem. | Replaced the tubing. | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX | CPU Unit bestellt. | SW installiert, auf der Windows ebene alle Einstellungen überprüft und eingestellt. SC und LP eingesetzt Startphase QC und Kalibration. | SWO-127156 | No | CPU Unit. | SW installed, on the Windows level, all settings are checked and set. SC and LP used start phase QC and calibration. | Revisit required | HW | DMS_PC | CPU
ABL90 FLEX PLUS |  | B | D157842900598FE9 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX | 清洁进样口后修复(by andrew) | 清洁进样口后修复(by andrew) | SWO-127572 | No | Repair after cleaning the inlet (by andrew) | Repair after cleaning the inlet (by andrew) | Defect | HW | Inlet module | Inlet Gasket
TCM5 FLEX |  | B | D144434316335F64 | No
AQT90 FLEX |  | B | SWO-126934 | No |  |  | Defect | Consumables | Test Kit | Cartridge | Lot 11755
ABL815 FLEX | Ausbau des AutoCheck Moduls. Ventilator defekt, Lüfter abgefallen., erneuert. Funktion nach Inbetriebnahme wieder in Ordnung. Laut QC Plan werden an Wochenenden dreimal die AutoCheck Ampullen gemessen in den fünf Tagen der Woche, wie von der RiliBÄK nur zweimal am Tag gefordert. Laut Qualitätskontroll Historie ist ein Mehrverbrauch von AutoCheck Ampullen ersichtlich. Kalibrationen und Qualitätskontrolle in Ordnung. | Ausbau des AutoCheck Moduls. Ventilator defekt, Lüfter abgefallen., erneuert. Funktion nach Inbetriebnahme wieder in Ordnung. Laut QC Plan werden an Wochenenden dreimal die AutoCheck Ampullen gemessen in den fünf Tagen der Woche, wie von der RiliBÄK nur zweimal am Tag gefordert. Laut Qualitätskontroll Historie ist ein Mehrverbrauch von AutoCheck Ampullen ersichtlich. Kalibrationen und Qualitätskontrolle in Ordnung. | SWO-127176 | No | Removal of the AutoCheck module. Fan defective, fan dropped, renewed. Function is back in order after commissioning. According to the QC plan, on weekends the AutoCheck ampoules are measured three times in the five days of the week, as required by the RiliBÄK only twice a day. According to quality control history, an additional consumption of AutoCheck ampoules can be seen. Calibrations and quality control in order. | Removal of the AutoCheck module. Fan defective, fan dropped, renewed. Function is back in order after commissioning. According to the QC plan, on weekends the AutoCheck ampoules are measured three times in the five days of the week, as required by the RiliBÄK only twice a day. According to quality control history, an additional consumption of AutoCheck ampoules can be seen. Calibrations and quality control in order. | Defect | HW | AutoCheck Module | Fan | 763
ABL815 FLEX | Display durch Ersatzmodul(R1346N7104, Original: R1611N0637) ersetzt, Warmstart durchgeführt und Tochpanel kalibriert | B | SWO-127109 | No | Display by replacement module (R1346N7104; Original: R1611N0637) replaced, hot start performed and tochpanel calibrated |  | Defect | HW | DMS_PC | Display
ABL827 FLEX |  | B | D144530B160FD389 | No |  |  | Defect | HW | Oxi module | Lamp unit | Neon
ABL815 FLEX | Intensive Reinigung der pH-Elektrode, pO2-Elektrode membraniert, korrekte Kal abgewartet, QK ok, Messungen an Blut ok | Intensive Reinigung der pH-Elektrode, pO2-Elektrode membraniert, korrekte Kal abgewartet, QK ok, Messungen an Blut ok | SWO-127401 | No | Intensive cleaning of the pH electrode, pO2 electrode membraneed, correct kalwaited, QK ok, measurements on blood ok | Intensive cleaning of the pH electrode, pO2 electrode membraneed, correct kalwaited, QK ok, measurements on blood ok | Cleaned | Electrode | Electrode | pH
AQT90 FLEX | Cleaned the storage wheel and observed that inner lock of cartridge guide of storage wheel is not fixed properly. Clean the incubator wheel also, checked the motor performance via service program. | Guided to Mr. Manoj Sharma to Run the safe positions. Now Mr. Sharma was there he took the telephonic help of Mr. Deepak Chand and issue got resolved . Run the samples. | SWO-127263 | No | Cleaned the storage wheel and observed that inner lock of cartridge guide of storage wheel is not fixed properly. Clean the incubator wheel also, checked the motor performance via service program. | Guided to Mr. Manoj Sharma to Run the safe positions. Now Mr. Sharma was there he took the telephonic help of Mr. Deepak Chand and issue got resolved . Run the samples. | Defect | HW | Storage Module | Cartridge guide | SETSAFE 22
AQT90 FLEX | Antriebsriemen der Einlasstüren erneuert. Funktion iO. | Antriebsriemen der Einlasstüren erneuert. Funktion iO. | SWO-127003 | No | drive belt of the inlet doors. Function iO. | drive belt of the inlet doors. Function iO. | Defect | HW | Cover and Chassis | Inlet Door
ABL80 FLEX CO-OX | Überprüfung Dateiarchiv Kal und QC Kal und QC Level 2 ohne Probleme => QC Level 1 alle Parameter fehlerhaft (außer Bereich, kein Endpunkt , Kalkulationsfehler) Durchführung tHb-Kal i.O. => Fcuv = 0,983 Anschließende Kal2 i.O. => bei folgender QC Level 1 Messung kommt Meldung auf mögliche Luft in QC Lösung => Austausch LP Mit Hr. Kühn Reklamationmeldung fertig gemacht | Überprüfung Dateiarchiv Kal und QC Kal und QC Level 2 ohne Probleme => QC Level 1 alle Parameter fehlerhaft (außer Bereich; kein Endpunkt ; Kalkulationsfehler) Durchführung tHb-Kal i.O. => Fcuv = 0,983 Anschließende Kal2 i.O. => bei folgender QC Level 1 Messung kommt Meldung auf mögliche Luft in QC Lösung => Austausch LP Mit Hr. Kühn Reklamationmeldung fertig gemacht | SWO-127344 | No | Checking file archive Kal and QC Kal and QC Level 2 without problems => QC Level 1 all parameters faulty (except range; no endpoint ; Calculation error) Execution tHb-Kal i.O. => Fcuv = 0.983 Subsequent Kal2 i.O. => at the following QC Level 1 measurement comes notification on possible air in QC solution => Exchange LP With Mr. Kühn complaint report completed | Checking file archive Kal and QC Kal and QC Level 2 without problems => QC Level 1 all parameters faulty (except range; no endpoint ; Calculation error) Execution tHb-Kal i.O. => Fcuv = 0.983 Subsequent Kal2 i.O. => at the following QC Level 1 measurement comes notification on possible air in QC solution => Exchange LP With Mr. Kühn complaint report completed
ABL80 FLEX BASIC | Clot was trapped inside the SC, so flushed the SC using heparin, kept the SC for 20 mins in the refrigerator, reinstalled both SP and SC, after 2 point calibration run a sample, machine is working fine. | Clot was trapped inside the SC, so flushed the SC using heparin, kept the SC for 20 mins in the refrigerator, reinstalled both SP and SC, after 2 point calibration run a sample, machine is working fine. | SWO-128441 | No | Clot was trapped inside the SC, so flushed the SC using heparin, kept the SC for 20 mins in the refrigerator, reinstalled both SP and SC, after 2 point calibration run a sample, machine is working fine. | Clot was trapped inside the SC, so flushed the SC using heparin, kept the SC for 20 mins in the refrigerator, reinstalled both SP and SC, after 2 point calibration run a sample, machine is working fine.
ABL90 |  | B | D1457CA28C4B2C5B | No |  |  | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | SW installiert, auf der Windows ebene alle Einstellungen überprüft und eingestellt. SC und LP eingesetzt Startphase QC und Kalibration. | SW installiert, auf der Windows ebene alle Einstellungen überprüft und eingestellt. SC und LP eingesetzt Startphase QC und Kalibration. | SWO-127473 | No | SW installed, on the Windows level, all settings are checked and set. SC and LP used start phase QC and calibration. | SW installed, on the Windows level, all settings are checked and set. SC and LP used start phase QC and calibration. | Defect | HW | DMS_PC | CPU
ABL90 | REPLACED BATTERY PAC WITH NO RESOLUTION | REPLACED POWER SUPPLY | D1483C5778530235 | No | REPLACED BATTERY PAC WITH NO RESOLUTION | REPLACED POWER SUPPLY | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-127402 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | pO2
ABL825 FLEX | Oximetriemodul erneuert. | Oximetriemodul erneuert. | SWO-127505 | No | Oximetry module. | Oximetry module. | Defect | HW | Oxi module | Oxi module
ABL815 FLEX | OXI Modul PCB nicht verfügbar. OXI Modul erneuert. Gerät kalibriert und QC durchgeführt. | OXI Modul PCB nicht verfügbar. OXI Modul erneuert. Gerät kalibriert und QC durchgeführt. | SWO-128317 | No | OXI module PCB not available. OXI module renewed. calibrated and QC performed. | OXI module PCB not available. OXI module renewed. calibrated and QC performed. | Defect | HW | Oxi module | PCB, Oxi
ABL820 FLEX | customer stated inlet has been leaking out the inlet and now calibration of the inlet will not pass--will order inlet friction dector to have on hand if problem continues | replace inlet gasket no help-- replace probe still no help-- replace inlet nipple on the left do to clot -- pulled inlet probe all the way out and performed inlet calibration this enabled the inlet to pass calibration | D145F8578C13C2C7 | No | customer stated inlet has been leaking out the inlet and now calibration of the inlet will not pass--will order inlet friction dector to have on hand if problem continues | replace inlet gasket no help-- replace probe still no help-- replace inlet nipple on the left do to clot -- pulled inlet probe all the way out and performed inlet calibration this enabled the inlet to pass calibration | Defect | HW | Inlet module | Inlet cal. | Revisit required?
ABL825 FLEX | Confirmed issue when onsite. Observed salt crust built up on pH end of membrane jacket. This membrane jacket may not be sealing with measuring chamber completely, causing the leak or seepage. | Replaced inlet gasket, pump tubing, connector at inlet module and tubing between pH BG and EI Met modules. Performed pump cal and liquid sensor calibration which passed. | D145A6F48C1FF8C3 | No | Confirmed issue when onsite. Observed salt crust built up on pH end of membrane jacket. This membrane jacket may not be sealing with measuring chamber completely, causing the leak or seepage. | Replaced inlet gasket, pump tubing, connector at inlet module and tubing between pH BG and EI Met modules. Performed pump cal and liquid sensor calibration which passed. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
AQT90 FLEX | Replaced Load the TnT Cartridge. According to Data Log total 44 nos. samples has been aborted in AQT90 after installation. | Guided to Mr. Manoj Sharma to Run the safe positions. Now Mr. Sharma was there he took the telephonic help of Mr. Deepak Chand and issue got resolved . Run the samples. | SWO-127279 | No | Replaced Load the TnT Cartridge. According to Data Log total 44 nos. samples has been aborted in AQT90 after installation. | Guided to Mr. Manoj Sharma to Run the safe positions. Now Mr. Sharma was there he took the telephonic help of Mr. Deepak Chand and issue got resolved . Run the samples. | Defect | Consumables | Test Kit | Cartridge | Unknown
AQT90 FLEX |  | B | D145863D8C0F7ECE | No |  |  | Defect | HW | Analyzer Control Module | Display cable
ABL800 BASIC | Oxymentry PCB and Neon lamp defective,need to replace with new one. | B | SWO-128451 | No | Oxymentry PCB and Neon lamp defective,need to replace with new one. |  | Defect | HW | Oxi module | Lamp unit
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-127668 | No | Replacement test card repair | Replacement test card repair
ABL80 FLEX BASIC | 没有用备件修复(by andrew) | 没有用备件修复(by andrew) | SWO-127322 | No | No spare parts repair (by andrew) | No spare parts repair (by andrew)
ABL80 FLEX BASIC | Checked the solution pack, no any error found in the solution pack. Tried reinstalling both SP and SC with air detection ON and OFF, same problem was observed. Replaced the Sensor Cassette, after calibratio, run 2 samples. Results are OK. Machine is working fine. | Checked the solution pack, no any error found in the solution pack. Tried reinstalling both SP and SC with air detection ON and OFF, same problem was observed. Replaced the Sensor Cassette, after calibratio, run 2 samples. Results are OK. Machine is working fine. | SWO-128445 | No | Checked the solution pack, no any error found in the solution pack. Tried reinstalling both SP and SC with air detection ON and OFF, same problem was observed. Replaced the Sensor Cassette, after calibratio, run 2 samples. Results are OK. Machine is working fine. | Checked the solution pack, no any error found in the solution pack. Tried reinstalling both SP and SC with air detection ON and OFF, same problem was observed. Replaced the Sensor Cassette, after calibratio, run 2 samples. Results are OK. Machine is working fine.
ABL90 FLEX | Gerät geprüft. LowLevel Protokoll von TCP IP auf (TCP IP)(ASTM) umgestellt. - ABL sendet nun keine Daten. HighLevel Protokoll umgestellt. Diverse Telefonate mit der IT in Leverkusen geführt. | B | SWO-127880 | No | device checked. Changed the LowLevel protocol from TCP IP to (TCP IP) (ASTM). - ABL now does not send any data. HighLevel protocol. Various phone calls with IT in Leverkusen. |  | Setup | His/Lis | His/Lis | His/Lis
ABL90 FLEX PLUS | tHb Kalibration durchgeführt mit eigenem Kalibrator. Tausch Lösungspack. Weiterhin Grenzverletzung im Level blau. Tausch der Sensorkassette. Nach Konditionierung sind die pCO2 Werte wieder zielwertig. Beide Betriebsmittel reklamiert. Kalibrationen und Qualitätskontrolle in Ordnung. | tHb Kalibration durchgeführt mit eigenem Kalibrator. Tausch Lösungspack. Weiterhin Grenzverletzung im Level blau. Tausch der Sensorkassette. Nach Konditionierung sind die pCO2 Werte wieder zielwertig. Beide Betriebsmittel reklamiert. Kalibrationen und Qualitätskontrolle in Ordnung. | SWO-127852 | No | tHb calibration carried out with its own calibrator. Exchange solution pack. Further boundary violation in the level blue. Replacement of the sensor cassette. After conditioning, the pCO2 values are again targeted. Both equipment claims. Calibrations and quality control in order. | tHb calibration carried out with its own calibrator. Exchange solution pack. Further boundary violation in the level blue. Replacement of the sensor cassette. After conditioning, the pCO2 values are again targeted. Both equipment claims. Calibrations and quality control in order. | Assisted Replacement | SC | SC | pCO2
TCM5 FLEX |  | B | D14702D0025CC973 | No
AQT90 FLEX | 更换903-113后修复(by andrew) | 更换903-113后修复(by andrew) | SWO-127238 | No | Repair after replacement 903-113 (by andrew) | Repair after replacement 903-113 (by andrew) | Defect | HW | Aspiration Module | PCB, Pippette | 1999
ABL800 BASIC | Reset the connection. check the status. flap position detection issue.sample aborted. Need to replace the flap detector PCB. | B | SWO-128436 | No | Reset the connection. check the status. flap position detection issue.sample aborted. Need to replace the flap detector PCB. |  | Revisit required | HW | Inlet module | Flap PCB
ABL800 BASIC |  | B | SWO-127616 | No |  |  | Assisted Replacement | Electrode | Membrane | Lac | Cal.
ABL835 FLEX | . New large inlet nipple fitted. Reference membrane replaced. Lv's 3 & 4 QC were ran, the pCO2 QC was central on Lv3 and Lv 4 had improved but was still low. New pCO2 membrane fitted to the analyser. | New upper sealing ring fitted to pCO2 electrode. New reference membrane fitted to analyser, (the current one had been installed for 21 days). | SWO-127848 | No | . New large inlet nipple fitted. Reference membrane replaced. Lv's 3 & 4 QC were ran, the pCO2 QC was central on Lv3 and Lv 4 had improved but was still low. New pCO2 membrane fitted to the analyser. | New upper sealing ring fitted to pCO2 electrode. New reference membrane fitted to analyser, (the current one had been installed for 21 days). | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | ディスプレイユニット交換 910-320CPH16090001C | ディスプレイユニット交換 910-320CPH16090001C | SWO-128512 | No | Display Unit Replacement 910-320CPH16090001C | Display Unit Replacement 910-320CPH16090001C | Defect | HW | DMS_PC | Touch screen
ABL835 FLEX | All three levels reading low. Level 1 reading at or below low end of package insert. | Replaced membrane, changed cal1 and cal2. | D146D8E80202A13D | No | All three levels reading low. Level 1 reading at or below low end of package insert. | Replaced membrane, changed cal1 and cal2. | Assisted Replacement | Electrode | Membrane | Na+ | Low QC
ABL800 BASIC | BG側測定室剥離の為、交換 EL側電極ハウジング固定用穴 再加工 | B | SWO-128815 | No | Reprocessing hole for fixing replacement EL side electrode housing for BG side measuring chamber peeling |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | BG側測定室交換 | B | SWO-128816 | No | BG side measuring chamber replacement |  | Defect | HW | pH/BG Module | Measuring chamber | pCO2
ABL835 FLEX |  | B | D146FCC0021388D4 | Yes |  |  | Re classified RMED | Demo installation
ABL800 BASIC | Need to replace printer head. | Need to replace printer head. | SWO-127579 | No | Need to replace printer head. | Need to replace printer head. | Revisit required | HW | Printer module | Printer Head
ABL815 FLEX | Monitor ersetzt mit Funktionstest, OK. | Monitor ersetzt mit Funktionstest, OK. | SWO-127818 | No | Monitor replaced with function test, OK. | Monitor replaced with function test, OK. | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX |  | B | D1471FEC020AA047 | No |  |  | Replaced | Analyzer removed | Analyzer removed | Analyzer removed
AQT90 FLEX | SENSÖRLER TEMİZLENEREK SORUN GİDERİLDİ. | B | SWO-127774 | Yes | SENSORS HAVE BEEN SECURED AND THE PROBLEM HAS BEEN SENT. |  | Re classified RMED | RTUR need upSWO training
ABL80 FLEX BASIC | カセット交換実施、センサーボードに腐食があるので注意 | B | SWO-134051 | No | Cassette replacement, be careful because there is corrosion on the sensor board
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-127950 | No | Repair after replacing test card | Repair after replacing test card
ABL80 CO-OX | 581 error | i replaced the hemolyzer, ran thb cal | D148151578434558 | No | 581 error | i replaced the hemolyzer, ran thb cal
ABL80 FLEX BASIC | Done two point calibration and QC all results found within the range. compared 3 samples all results found with in the range. explained pre analytical errors to avoid sample errors. | Done two point calibration and QC all results found within the range. compared 3 samples all results found with in the range. explained pre analytical errors to avoid sample errors. | SWO-128708 | No | Done two point calibration and QC all results found within the range. compared 3 samples all results found with in the range. explained pre analytical errors to avoid sample errors. | Done two point calibration and QC all results found within the range. compared 3 samples all results found with in the range. explained pre analytical errors to avoid sample errors.
ABL90 FLEX | Removed & cleaned touch screen. Calibrated screen & committed calibration to software. | Removed & cleaned touch screen. Calibrated screen & committed calibration to software. | SWO-128129 | No | Removed & cleaned touch screen. Calibrated screen & committed calibration to software. | Removed & cleaned touch screen. Calibrated screen & committed calibration to software. | Adjustment | HW | DMS_PC | Touch screen
ABL90 FLEX | Ran all fluid transport tests. Failing SP check, clot detection & pressure test. SP already replaced. Back flushed analyser. Still failing pressure test. Removed SC & manually flushed. Re ran all fluid transport tests. All passed. | Ran all fluid transport tests. Failing SP check, clot detection & pressure test. SP already replaced. Back flushed analyser. Still failing pressure test. Removed SC & manually flushed. Re ran all fluid transport tests. All passed. | SWO-128128 | No | Ran all fluid transport tests. Failing SP check, clot detection & pressure test. SP already replaced. Back flushed analyser. Still failing pressure test. Removed SC & manually flushed. Re ran all fluid transport tests. All passed. | Ran all fluid transport tests. Failing SP check, clot detection & pressure test. SP already replaced. Back flushed analyser. Still failing pressure test. Removed SC & manually flushed. Re ran all fluid transport tests. All passed. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Collected SP & brought to Hermitage. Loaded SP & left to calibrate. | Collected SP & brought to Hermitage. Loaded SP & left to calibrate. | SWO-127983 | No | Collected SP & brought to Hermitage. Loaded SP & left to calibrate. | Collected SP & brought to Hermitage. Loaded SP & left to calibrate. | Assisted Replacement | SP | SP | Unknown
ABL90 FLEX PLUS | Remplacement du module OXI complet, mise en route, calibration tHb et contrôle : OK Le module va être réparé à la firme. | B | SWO-128121 | No | Replacement of the complete OXI module, start-up, tHb calibration and control: OK The module will be repaired at the firm. |  | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX PLUS | Removed screen and removed touch screen and cleaned round edges as dirty. Refitted touch screen and screen and test ok. | Customer has cleaned and it is now functioning correctly. Service visit required as it has happened twice in one day. Details passed to DF | SWO-128133 | No | Removed screen and removed touch screen and cleaned round edges as dirty. Refitted touch screen and screen and test ok. | Customer has cleaned and it is now functioning correctly. Service visit required as it has happened twice in one day. Details passed to DF | Adjustment | HW | DMS_PC | Touch screen
ABL90 FLEX PLUS | IP Adressen eingetragen. Speicherbefehl ausgeführt. Kontrolle der Datenübertragung im Labor. Übertragung iO. | IP Adressen eingetragen. Speicherbefehl ausgeführt. Kontrolle der Datenübertragung im Labor. Übertragung iO. | SWO-128146 | No | IP addresses. Memory command executed. Control of data transmission in the laboratory. Transfer iO. | IP addresses. Memory command executed. Control of data transmission in the laboratory. Transfer iO. | Setup | His/Lis | His/Lis | His/Lis
ABL90 FLEX PLUS |  | B | D148397D78532173 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-128526 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Glu/Lac
AQT90 FLEX | REVISIÓN DE HISTÓRICOS. LIMPIEZA INTERIOR, FILTROS, SENSORES ÓPTICOS, IMPRESORA Y ÓPTICA. REALIZAR SUSTITUCIÓN DE FLUÍDICA GASTADA (KIT MANTENIMIENTO ANUAL). TEST TÉCNICOS DE FLUÍDICA, MECÁNICA Y ÓPTICA. LIMPIEZA ESPECIAL. CONTROLES. | REVISIÓN DE HISTÓRICOS. LIMPIEZA INTERIOR, FILTROS, SENSORES ÓPTICOS, IMPRESORA Y ÓPTICA. REALIZAR SUSTITUCIÓN DE FLUÍDICA GASTADA (KIT MANTENIMIENTO ANUAL). TEST TÉCNICOS DE FLUÍDICA, MECÁNICA Y ÓPTICA. LIMPIEZA ESPECIAL. | D148459578409027 | Yes | HISTORICAL REVIEW. INTERIOR CLEANING, FILTERS, OPTICAL SENSORS, PRINTER AND OPTICAL. MAKE SUBSTITUTE OF GASTATed FLUIDIC (ANNUAL MAINTENANCE KIT). TECHNICAL TEST OF FLUDIC, MECHANICAL AND OPTICAL. SPECIAL CLEANING. Controls. | HISTORICAL REVIEW. INTERIOR CLEANING, FILTERS, OPTICAL SENSORS, PRINTER AND OPTICAL. MAKE SUBSTITUTE OF GASTATed FLUIDIC (ANNUAL MAINTENANCE KIT). TECHNICAL TEST OF FLUDIC, MECHANICAL AND OPTICAL. SPECIAL CLEANING. | Re classified RMED | PM
ABL825 FLEX | Eredeti hőírófej tisztítása, Új hőírófej beépítése, Tesztelés, Kalibrálás, QC mérés | B | SWO-128082 | No | Original thermal head cleaning, new burner head installation, testing, calibration, QC measurement |  | Defect | HW | Printer module | Printer Head
ABL837 FLEX | 消毒実施 | B | SWO-130588 | No | Disinfection implementation |  | Cleaned | Fluidic pathway | Met II | Lac | Disinfected
ABL825 FLEX | Neue Autocheckmodul 902-821R2558N007 eingebaut, Kalibrierung und QC erfolgreich | Neue Autocheckmodul 902-821R2558N007 eingebaut, Kalibrierung und QC erfolgreich | SWO-128078 | No | New autocheck module 902-821R2558N007 built-in, calibration and QC successful | New autocheck module 902-821R2558N007 built-in, calibration and QC successful | Defect | HW | AutoCheck Module | Breaker arm
ABL835 FLEX | The measuring chamber was inspected and re-sat. It was noted that depending on how the measuring chamber and housing block were held when tightening it down, (holding it in a position whilst tightening) that this can effect the position of the reference electrode and membrane. The trapped rinse has been minimised by positioning the chambers during tightening. A new liquid sensor was fitted to the pH BG module. LS adjust performed and passed. | New upper sealing ring fitted to pCO2 electrode. New reference membrane fitted to analyser, (the current one had been installed for 21 days). | SWO-128039 | Yes | The measuring chamber was inspected and re-sat. It was noted that depending on how the measuring chamber and housing block were held when tightening it down, (holding it in a position whilst tightening) that this can effect the position of the reference electrode and membrane. The trapped rinse has been minimised by positioning the chambers during tightening. A new liquid sensor was fitted to the pH BG module. LS adjust performed and passed. | New upper sealing ring fitted to pCO2 electrode. New reference membrane fitted to analyser, (the current one had been installed for 21 days). | Re classified RMED | Courtesy
ABL825 FLEX | Ca. Elektrode geprüft und neu membraniert, Funktionskontrolle ok | Ca. Elektrode geprüft und neu membraniert, Funktionskontrolle ok | SWO-128125 | No | Approx. electrode tested and re-membraneed, function control ok | Approx. electrode tested and re-membraneed, function control ok | Assisted Replacement | Electrode | Membrane | Ca+
ABL800 BASIC | Reconnected all connection of printer unit. Also applied greece to moving section of printer. Temporary unit is working ok. | Reconnected all connection of printer unit. Also applied greece to moving section of printer. Temporary unit is working ok. | SWO-128387 | No | Reconnected all connection of printer unit. Also applied greece to moving section of printer. Temporary unit is working ok. | Reconnected all connection of printer unit. Also applied greece to moving section of printer. Temporary unit is working ok. | Defect | HW | Printer module | Printer module
ABL825 FLEX | ph will not calibrate for sensitivity or status | found that the glass tip was broken on replacement electrode. | D1486CFF7848F986 | No | ph will not calibrate for sensitivity or status | found that the glass tip was broken on replacement electrode. | Assisted Replacement | Electrode | Electrode | pH | Glass Electrode broken
AQT90 FLEX | Incubationswheel ausgebaut. Es schien, als sei einmal Flüssigkeit ausgelaufen, da an der Unterseite Auskristallisierungen zu sehen und auch zu spüren waren. Reinigung durchgeführt und Gerät überprüft. 8 Messungen durchgeführt. | Incubationswheel ausgebaut. Es schien, als sei einmal Flüssigkeit ausgelaufen, da an der Unterseite Auskristallisierungen zu sehen und auch zu spüren waren. Reinigung durchgeführt und Gerät überprüft. 8 Messungen durchgeführt. | SWO-128105 | No | Incubationswheel. It seemed as if liquid had once leaked, as crystallizations could be seen and felt at the bottom. cleaning and checking the device. 8 measurements performed. | Incubationswheel. It seemed as if liquid had once leaked, as crystallizations could be seen and felt at the bottom. cleaning and checking the device. 8 measurements performed. | Cleaned | HW | Incubator module | Incubator wheel | 5663
AQT90 FLEX | CPU-Unit getauscht, Einlaufverfahren | CPU-Unit getauscht, Einlaufverfahren | SWO-128126 | No | CPU unit exchanged, inlet procedure | CPU unit exchanged, inlet procedure | Defect | HW | Analyzer Control Module | CPU
AQT90 FLEX | Ausbau des alten Mainboards und Einbau eines neuen. Einspielung der aktuellen Software ins Mainboard. Hochladen der special.reg Datei , klappte nicht auf Anhieb. Eine andere Datei wurde hochgeladen, die umgeschrieben wurde. Blankmessungen durchgeführt, beim ersten Durchgang wurden alle abgebrochen mit der Fehlermeldung 8130 Testflowwithoutdilution... . Nach Rücksprache mit Technischen Support, erneut die Blankmessungen durchgeführt, diesmal Meldung BNP Sensor defekt. Überprüfung ergab, Sensorkontakt nicht tief genug eingesteckt. Werte Blankmessungen in Ordnung. Qualitätskontrolle innerhalb Grenzwerte. | Ausbau des alten Mainboards und Einbau eines neuen. Einspielung der aktuellen Software ins Mainboard. Hochladen der special.reg Datei , klappte nicht auf Anhieb. Eine andere Datei wurde hochgeladen, die umgeschrieben wurde. Blankmessungen durchgeführt, beim ersten Durchgang wurden alle abgebrochen mit der Fehlermeldung 8130 Testflowwithoutdilution... . Nach Rücksprache mit Technischen Support, erneut die Blankmessungen durchgeführt, diesmal Meldung BNP Sensor defekt. Überprüfung ergab, Sensorkontakt nicht tief genug eingesteckt. Werte Blankmessungen in Ordnung. Qualitätskontrolle innerhalb Grenzwerte. | SWO-128131 | No | Removal of the old motherboard and installation of a new one. Insertthestandard of the current software into the motherboard. Uploading the special.reg file did not work immediately. Another file was uploaded that was rewritten. Blank measurements were made, during the first pass all were aborted with the error message 8130 Testflowwithoutdilution... . After consultation with technical support, again the blank measurements were performed, this time reporting BNP sensor defective. verification showed sensor contact not inserted deep enough. Values blank measurements in order. Quality control within limit values. | Removal of the old motherboard and installation of a new one. Insertthestandard of the current software into the motherboard. Uploading the special.reg file did not work immediately. Another file was uploaded that was rewritten. Blank measurements were made, during the first pass all were aborted with the error message 8130 Testflowwithoutdilution... . After consultation with technical support, again the blank measurements were performed, this time reporting BNP sensor defective. verification showed sensor contact not inserted deep enough. Values blank measurements in order. Quality control within limit values. | Defect | HW | Analyzer Control Module | PCB, Main board
ABL80 FLEX BASIC | 更换903-228修复 | 更换903-228修复 | SWO-128217 | No | Replacement 903-228 repair | Replacement 903-228 repair
ABL80 CO-OX | Metal latching plate was lose. | Tightened the metal latching plate. | D1492EECEE5A6713 | No | Metal latching plate was lose. | Tightened the metal latching plate.
ABL80 FLEX | Tried reinstalling the solution pack 2-3 times, same problem was observed. Also flushed the SC using heparin, tried reinstalling the SC, but same problem was observed. After repeating the same procedure, run 1 sample. Results are found OK. Machine is working fine | Tried reinstalling the solution pack 2-3 times, same problem was observed. Also flushed the SC using heparin, tried reinstalling the SC, but same problem was observed. After repeating the same procedure, run 1 sample. Results are found OK. Machine is working fine | SWO-128448 | No | Tried reinstalling the solution pack 2-3 times, same problem was observed. Also flushed the SC using heparin, tried reinstalling the SC, but same problem was observed. After repeating the same procedure, run 1 sample. Results are found OK. Machine is working fine | Tried reinstalling the solution pack 2-3 times, same problem was observed. Also flushed the SC using heparin, tried reinstalling the SC, but same problem was observed. After repeating the same procedure, run 1 sample. Results are found OK. Machine is working fine
ABL80 FLEX | Replaced SC with New one. | Replaced SC with New one. | SWO-128452 | No | Replaced SC with New one. | Replaced SC with New one.
ABL90 FLEX | Im Servicemenu die Temperaturen überprüft und Kalibrierung durchgeführt. Kontakte der Kassettenaufnahme gereinigt. Lösungspack erneuert da Fälligkeit erreicht. Gerät nochmals runter und wieder hochgefahren. Danach Kassette kalibriert, Kal und QC in Ordnung, Gerät viel aber wieder mit Temperatur im Substrat aus. Neue Kassette installiert, die aber einen pCO2 Empfindlichkeitsfehler hatte. Nochmals eine neue Kassette installiert und Gerät überprüft. Kassetten für den Kunden reklamiert. | Im Servicemenu die Temperaturen überprüft und Kalibrierung durchgeführt. Kontakte der Kassettenaufnahme gereinigt. Lösungspack erneuert da Fälligkeit erreicht. Gerät nochmals runter und wieder hochgefahren. Danach Kassette kalibriert, Kal und QC in Ordnung, Gerät viel aber wieder mit Temperatur im Substrat aus. Neue Kassette installiert, die aber einen pCO2 Empfindlichkeitsfehler hatte. Nochmals eine neue Kassette installiert und Gerät überprüft. Kassetten für den Kunden reklamiert. | SWO-128577 | No | In the service menu the temperatures are checked and calibration carried out. contacts of the cassette recording. Solution pack renews as maturity is reached. device again down and restarted. After that cassette calibrated, Kal and QC fine, device much but again with temperature in the substrate off. New cartridge installed, but it had a pCO2 sensitivity error. Again installed a new cassette and checked the device. cassettes for the customer. | In the service menu the temperatures are checked and calibration carried out. contacts of the cassette recording. Solution pack renews as maturity is reached. device again down and restarted. After that cassette calibrated, Kal and QC fine, device much but again with temperature in the substrate off. New cartridge installed, but it had a pCO2 sensitivity error. Again installed a new cassette and checked the device. cassettes for the customer. | Assisted Replacement | SC | SC | Temperature
ABL90 FLEX | Straightened bend in inlet probe. | Unable to resolve by phone advice, service call required. | SWO-128585 | No | Straightened bend in inlet probe. | Unable to resolve by phone advice, service call required. | User error | SP Activation | SP Activation | SP Activation
ABL90 FLEX | 更换903-158修复(by andrew) | 更换903-158修复(by andrew) | SWO-128169 | No | Replacement 903-158 repair (by andrew) | Replacement 903-158 repair (by andrew) | Assisted Replacement | SP | SP | Leakage current
ABL825 FLEX |  | B | D149594CEE41088E | Yes |  |  | Re classified RMED | Logistical upSWO
ABL815 FLEX | QC Modul ausgebaut und Lüfter ersetzt, | QC Modul ausgebaut und Lüfter ersetzt, | SWO-128454 | No | QC module removed and fan replaced, | QC module removed and fan replaced, | Defect | HW | Auto Check Module | Fan | Temp
ABL835 FLEX |  | B | D1495C5FEE4F88D7 | No |  |  | Defect | HW | Fluidic manifold | Pump rotor
ABL815 FLEX | O.g. Arbeiten, Display kalibriert, Leihmodul zurückgenommen | O.g. Arbeiten, Display kalibriert, Leihmodul zurückgenommen | SWO-128228 | No | O.g. working, display calibrated, loan module retracted | O.g. working, display calibrated, loan module retracted | Defect | HW | DMS_PC | Touch screen
ABL815 FLEX | O.g. Arbeiten, Display kalibriert, Leihmodul zurückgenommen | O.g. Arbeiten, Display kalibriert, Leihmodul zurückgenommen | SWO-128229 | No | O.g. working, display calibrated, loan module retracted | O.g. working, display calibrated, loan module retracted | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX | Desmontaje y sustitución del soporte de muestras. Revisión y comprobación del funcionamiento. | Desmontaje y sustitución del soporte de muestras. | D14998D8EE48F717 | No | Disassembly and replacement of the sample holder. Review and check the operation. | Disassembly and replacement of the sample holder. | Defect | HW | Aspiration Module | Tube holder
AQT90 FLEX | Desmontaje y sustitución del soporte de muestras. Revisión y comprobación del fucionamiento | Desmontaje y sustitución del soporte de muestras. Revisión y comprobación del fucionamiento | D14E6B78C6191C76 | No | Disassembly and replacement of the sample holder. Review and verification of the fucion | Disassembly and replacement of the sample holder. Review and verification of the fucion | Defect | HW | Aspiration Module | Tube holder | To be returned
ABL90 FLEX PLUS | Schlauchwege erneuert, Analysator gespült, LP (XB-06) aus Bestand der GS19 eingebaut, | Schlauchwege erneuert, Analysator gespült, LP (XB-06) aus Bestand der GS19 eingebaut, | SWO-128474 | No | Hose paths renewed, analyzer rinsed, LP (XB-06) built-in from stock of GS19, | Hose paths renewed, analyzer rinsed, LP (XB-06) built-in from stock of GS19, | Defect | HW | PM Kit | PM Kit
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-128529 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX CO-OX | Oxiventil erneuert. Kalibrationen und Qualitätskontrollen in Ordnung. | Oxiventil erneuert. Kalibrationen und Qualitätskontrollen in Ordnung. | SWO-128715 | No | Oxivalve renewed. calibrations and quality controls in order. | Oxivalve renewed. calibrations and quality controls in order.
ABL90 FLEX | Mechanik des Sensorfachs überprüft und justiert, Sensorkassette installiert, Initialisierungsphase abgewartet. | Mechanik des Sensorfachs überprüft und justiert, Sensorkassette installiert, Initialisierungsphase abgewartet. | SWO-128642 | No | Mechanics of the sensor compartment checked and adjusted, sensor cassette installed, initialization phase awaited. | Mechanics of the sensor compartment checked and adjusted, sensor cassette installed, initialization phase awaited. | Adjustment | HW | Cover and Chassis | Sensor cassette door
ABL90 FLEX PLUS | Touch screen cleaned and checked for damage ok. Unable to calibrate screen is still not responding. New touch screen on order. | Advised customer screen needs calibration . Engineer required . | SWO-128754 | No | Touch screen cleaned and checked for damage ok. Unable to calibrate screen is still not responding. New touch screen on order. | Advised customer screen needs calibration . Engineer required . | Revisit required | HW | DMS_PC | Touch screen
ABL90 FLEX PLUS | Replaced inlet probe. | Replaced inlet probe. | SWO-128847 | No | Replaced inlet probe. | Replaced inlet probe. | Replaced | HW | Auto Inlet module | Probe bend
AQT90 FLEX | Cartridge Kunststoff-Führung korrekt eingesetzt, Storagewheel kontrolliert, Probelauf | Cartridge Kunststoff-Führung korrekt eingesetzt, Storagewheel kontrolliert, Probelauf | SWO-128843 | No | Cartridge plastic guide correctly inserted, storagewheel controlled, test run | Cartridge plastic guide correctly inserted, storagewheel controlled, test run | Adjustment | HW | Storage Module | Storage wheel
ABL80 FLEX CO-OX | Zur Reparatur Überbrückung für ID 303980 aufgestellt | Zur Reparatur Überbrückung für ID 303980 aufgestellt | SWO-128922 | No | For repair bridging for ID 303980 set up | For repair bridging for ID 303980 set up
ABL80 FLEX Basic version |  | B | D14E6BD3C6185F79 | No
ABL80 CO-OX | Analyzer doesn't see SP. When customer tried to bring analyzer up, SP errors present. Customer had to use SP and SC with other analyzer. Removed SP 224021 and SC 221929. | Reset of SP deon't work, analyzer does not see SP at all. Replaced SP and SC. | D14E5880C65149F4 | No | Analyzer doesn't see SP. When customer tried to bring analyzer up, SP errors present. Customer had to use SP and SC with other analyzer. Removed SP 224021 and SC 221929. | Reset of SP deon't work, analyzer does not see SP at all. Replaced SP and SC.
ABL90 FLEX | ASTM Verbindung mit ASTM.net an den Geräten im Kreißsaal und ITS überprüft an beiden Geräten werden die Daten korrekt gesendet astm.ini im ITS-Gerät angepasst - weiterhin LIS KIS: Paket Sende-Fehler opus L Herr Müller wendet sich an Opus | ASTM Verbindung mit ASTM.net an den Geräten im Kreißsaal und ITS überprüft an beiden Geräten werden die Daten korrekt gesendet astm.ini im ITS-Gerät angepasst - weiterhin LIS KIS: Paket Sende-Fehler opus L Herr Müller wendet sich an Opus | SWO-129002 | No | ASTM connection with ASTM.net checked on the devices in the delivery room and ITS on both devices the data is sent correctly astm.ini in the ITS device adapted - continue LIS KIS: package send error opus L Mr. Müller turns to Opus | ASTM connection with ASTM.net checked on the devices in the delivery room and ITS on both devices the data is sent correctly astm.ini in the ITS device adapted - continue LIS KIS: package send error opus L Mr. Müller turns to Opus | Setup | SW | His/Lis | His/Lis
ABL90 FLEX PLUS | Neu gelieferten Lösungspack eingesetzt. Korrektur zurückgestellt. Alten Lösungspack reklamiert. | Neu gelieferten Lösungspack eingesetzt. Korrektur zurückgestellt. Alten Lösungspack reklamiert. | SWO-128968 | No | newly delivered solution pack. Correction deferred. old solution pack. | newly delivered solution pack. Correction deferred. old solution pack. | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX PLUS | J'ai vérifier le SC et fait tout les test dans le menu service ( oxi test et pO2 test) : tout fonctionne normal . Nettoyage du joint d'intro OK. Calibrations OK et tout les CQ sont bon (293 mmHg). Si le problème reste en doit remplacer le SC | B | SWO-129332 | No | I check the SC and do all the tests in the service menu (oxi test and pO2 test): everything works normal. Cleaning the intro joint OK. OK calibrations and all CQ are good (293 mmHg). If the problem remains must replace the SC |  | Cleaned | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Überprüfung Dateiarchiv Aktivitäten => gestern 20 Uhr Einlass geöffnet => seitdem mit Positionsfehler => 22:51 Uhr Wechsel der Einlassnadel => sonst Gerät im Modus Intervention erforderlich LP über Service entfernt Einlass über Service probiert anzusteuern => keine Aktivität Gerät geöffnet und Einlass händisch geöffnet und geschlossen => keine Blockierung in Mechanik Überprüfung Einlassdichtung und Nadel => alles i.O. LP im Gerät installiert, Einlass händisch geschlossen, Gerät neugestartet Startphase abgewartet Kal und QC i.O. | Überprüfung Dateiarchiv Aktivitäten => gestern 20 Uhr Einlass geöffnet => seitdem mit Positionsfehler => 22:51 Uhr Wechsel der Einlassnadel => sonst Gerät im Modus Intervention erforderlich LP über Service entfernt Einlass über Service probiert anzusteuern => keine Aktivität Gerät geöffnet und Einlass händisch geöffnet und geschlossen => keine Blockierung in Mechanik Überprüfung Einlassdichtung und Nadel => alles i.O. LP im Gerät installiert, Einlass händisch geschlossen, Gerät neugestartet Startphase abgewartet Kal und QC i.O. | SWO-128910 | No | Review File Archive Activities => Yesterday 8 pm Inlet opened => since then with position error => 22:51 o'clock Change of inlet pin => otherwise device in mode intervention required LP via service removed inlet via service tried => no activity Device opened and inlet manually opened and closed => no blocking in mechanics inspection inlet seal and needle => all i.O. LP installed in the device, inlet manually closed, device restarted start phase awaited Kal and QC i.O. | Review File Archive Activities => Yesterday 8 pm Inlet opened => since then with position error => 22:51 o'clock Change of inlet pin => otherwise device in mode intervention required LP via service removed inlet via service tried => no activity Device opened and inlet manually opened and closed => no blocking in mechanics inspection inlet seal and needle => all i.O. LP installed in the device, inlet manually closed, device restarted start phase awaited Kal and QC i.O. | Adjustment | HW | Auto Inlet module | Auto Inlet module
ABL90 FLEX PLUS |  | B | D14E9194C6119A95 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | clot in SC
ABL90 FLEX PLUS | Ran oxi complete test, failing on spectrometer noise and Hemolyzer cavitation absorb. Ran the test two more times, still failing with results varying wildly. | B | SWO-129759 | No | Ran oxi complete test, failing on spectrometer noise and Hemolyzer cavitation absorb. Ran the test two more times, still failing with results varying wildly. |  | Unknown | Revisit required | Revisit required | Fail Oxi test
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-129235 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | Reference | 1361
AQT90 FLEX | REVISIÓN DE HISTÓRICOS. LIMPIEZA DE TODOS LOS SENSORES, RODAMIENTOS Y SUPERFÍCIES. LUBRICACIÓN DE PARTES MOTORAS. CONSTRUCCIÓN DE UN NUEVO FELT PAD. REALIZACIÓN DE TESTS TÉCNICOS. COMPROBAR CORRECTO FUNCIONAMIENTO. CALIBRACIONES Y CONTROLES. | REVISIÓN DE HISTÓRICOS. LIMPIEZA DE TODOS LOS SENSORES, RODAMIENTOS Y SUPERFÍCIES. LUBRICACIÓN DE PARTES MOTORAS. CONSTRUCCIÓN DE UN NUEVO FELT PAD. REALIZACIÓN DE TESTS TÉCNICOS. COMPROBAR CORRECTO FUNCIONAMIENTO. | D14FC3273C3D3C7B | No | HISTORICAL REVIEW. CLEANING OF ALL SENSORS, BEARINGS AND SURFACES. LUBRICATION OF MOTOR PARTS. CONSTRUCTION OF A NEW FELT PAD. PERFORMING TECHNICAL TESTS. CHECK CORRECT OPERATION. CALIBRATIONS AND CONTROLS. | HISTORICAL REVIEW. CLEANING OF ALL SENSORS, BEARINGS AND SURFACES. LUBRICATION OF MOTOR PARTS. CONSTRUCTION OF A NEW FELT PAD. PERFORMING TECHNICAL TESTS. CHECK CORRECT OPERATION. | Cleaned | HW | Incubator module | Sensor, Incubator
ABL800 BASIC | オキシモジュール内液センサー交換及びエアライン洗浄を実施しました。 暫く様子見をお願い致します。 intensity = 13⇒26 | オキシモジュール内液センサー交換及びエアライン洗浄を実施しました。 暫く様子見をお願い致します。 intensity = 13⇒26 | SWO-128831 | No | The oxy module liquid sensor was replaced and airline cleaning was carried out. Please take a look for a while. intensity = 13⇒26 | The oxy module liquid sensor was replaced and airline cleaning was carried out. Please take a look for a while. intensity = 13⇒26 | Defect | HW | Oxi module | LS
ABL835 FLEX | Sealing o ring and protective sleeve missing from electrode. | Replaced electrode | D14E5D74C62B87AE | No | Sealing o ring and protective sleeve missing from electrode. | Replaced electrode | Assisted Replacement | Electrode | Electrode | Na+ | DOA
ABL835 FLEX | Gas 1 Regulator ersetzt. Kal durchgeführt. Gas1 Fehler tritt nicht mehr auf. Drift bei PCo2 PO2 und Gas2 wird bei voller Flasche als leer angezeigt. Valve G2 schaltet nicht. Valve und Kabel ok..24V bei G1 messbar.....0,0 V bei G2 Valve messbar. In GS anderes AU Board geholt und eingebaut. AU PCB ersetzt. Key | B | SWO-128926 | No | Gas 1 Regulator replaced. Kal. Gas1 error no longer occurs. Drift at PCo2 PO2 and Gas2 is shown as empty at full bottle. Valve G2 does not switch. Valve and cable ok.. 24V at G1 measurable..... 0.0 V measurable at G2 Valve. In GS other AU board fetched and installed. AU PCB. Key |  | Defect | HW | Gas Module | Flow regulator 1
ABL825 FLEX | Found spring in connector out of position. | Cleaned leakage. Reinstalled spring in electrode connector. | D14E674AC61CFDB5 | No | Found spring in connector out of position. | Cleaned leakage. Reinstalled spring in electrode connector. | Adjustment | HW | pH/BG Module | Connector | Spring
AQT90 FLEX | Stabilise Incubation temperature to normal range. passed safe positions, machine working ok. | Stabilise Incubation temperature to normal range. passed safe positions, machine working ok. | SWO-129935 | No | Stabilise Incubation temperature to normal range. passed safe positions, machine working ok. | Stabilise Incubation temperature to normal range. passed safe positions, machine working ok. | Power off/on | SW | Incubator module | Incubator module
ABL80 FLEX Basic version | Verificação do funcionamento das válvulas com a primeira a falhar, substituição da manifold unit, montagem do equipamento, testes de funcionamento sem problemas, establização e passagem do QC. | B | D14FB02C3C055881 | No | Check the operation of the valves with the first to fail, replacing the manifold unit, assembling the equipment, running tests smoothly, establishing and passing the QC.
ABL80 FLEX BASIC | After install the SP unit is working properly. | After install the SP unit is working properly. | SWO-129609 | No | After install the SP unit is working properly. | After install the SP unit is working properly.
ABL80 OSM | CONFIRMED ISSUE WHEN ONSITE. | OBSERVED FLUIDS COLLECTED IN TUBING, NOT DRAINING. UNPLUGGED AND RESEATED VALVE CABLE SO VALVE ELECTRICAL CONTACTS ARE CONNECTED TO MAIN PCB FOR PROPER OPENING CLOSING. | D14F7BEA3C2D702C | No | CONFIRMED ISSUE WHEN ONSITE. | OBSERVED FLUIDS COLLECTED IN TUBING, NOT DRAINING. UNPLUGGED AND RESEATED VALVE CABLE SO VALVE ELECTRICAL CONTACTS ARE CONNECTED TO MAIN PCB FOR PROPER OPENING CLOSING.
ABL90 FLEX | パック交換 | パック交換 | SWO-129221 | No | Pack Replacement | Pack Replacement | Assisted Replacement | SP | SP | Leak current
ABL90 FLEX PLUS | Fitted new touch screen and calibrate the new screen in control panel ok. Test key functions all ok. | Fitted new touch screen and calibrate the new screen in control panel ok. Test key functions all ok. | SWO-129227 | No | Fitted new touch screen and calibrate the new screen in control panel ok. Test key functions all ok. | Fitted new touch screen and calibrate the new screen in control panel ok. Test key functions all ok. | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX PLUS | Checked unit back flush done, flushed SC - pO2status fails. Replaced SC, initialized ok. | Checked unit back flush done, flushed SC - pO2status fails. Replaced SC, initialized ok. | SWO-129868 | No | Checked unit back flush done, flushed SC - pO2status fails. Replaced SC, initialized ok. | Checked unit back flush done, flushed SC - pO2status fails. Replaced SC, initialized ok. | Assisted Replacement | SC | SC | pO2 status
ABL90 FLEX PLUS | Re install SW V3.1MR8 | Re install SW V3.1MR8 | SWO-129613 | No | Re install SW V3.1MR8 | Re install SW V3.1MR8 | Defect | SW | SW | application not loading | Re install SW
TCM CombiM |  | B | SWO-129086 | No
ABL90 FLEX | 更换新试剂包后修复 | 更换新试剂包后修复 | SWO-129234 | No | Repair after replacing a new kit | Repair after replacing a new kit | Assisted Replacement | SP | SP | pH, Glu, Lac | QC
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-130013 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | Lac | 378
ABL835 FLEX | Comprobar el módulo oxi, cambiar al unidad hemolizadora. Revisión general del equipo, realización de un kit de servicio con cambio de tubos y accesorios contenidos en el kit, aspiración del polvo del interior, limpieza de la cámara de medida y salida del desperdicio, comprobación del módulo oxi, realización de operaciones de mantenimiento y prueba con controles. Actualizar el software a la versión 6.16.ES. | Cambiar la unidad hemolizadora. | D14FA6B33C25C755 | No | Check the oxy module, switch to the hemolysizing unit. General equipment review, completion of a service kit with change of tubes and accessories contained in the kit, suction of dust from the inside, cleaning of the measuring chamber and waste output, checking the oxy module, performing operations of maintenance and testing with controls. Update the software to the 6.16.ES version. | Change the hemolysizing unit. | Defect | HW | Oxi module | Hemolyzer
ABL800 BASIC | Cleaned the fluidic pathway & electrode chambers. Need to replace reference membrane. Reference membrane replaced & issue resolved. | Cleaned the fluidic pathway & electrode chambers. Need to replace reference membrane. Reference membrane replaced & issue resolved. | SWO-129761 | No | Cleaned the fluidic pathway & electrode chambers. Need to replace reference membrane. Reference membrane replaced & issue resolved. | Cleaned the fluidic pathway & electrode chambers. Need to replace reference membrane. Reference membrane replaced & issue resolved. | Assisted Replacement | Electrode | Membrane | Reference
ABL835 FLEX | 次亜塩素酸で洗浄、測定室洗浄、比較電極次亜塩素酸で洗浄。 | B | SWO-129346 | No | Washed with hypochlorite, washed in the measuring chamber, and washed with comparative electrode hypochlorite. |  | Cleaned | Fluidic pathway | pH/BG Module | pH
ABL800 BASIC | Cleaned fluidic pathway. replaced lactate membrane. calibration done & passed. Analyser is ready & working satisfactory. | Cleaned fluidic pathway. replaced lactate membrane. calibration done & passed. Analyser is ready & working satisfactory. | SWO-129747 | No | Cleaned fluidic pathway. replaced lactate membrane. calibration done & passed. Analyser is ready & working satisfactory. | Cleaned fluidic pathway. replaced lactate membrane. calibration done & passed. Analyser is ready & working satisfactory. | Assisted Replacement | Electrode | Membrane | Lac | Sensititvity
ABL825 FLEX | Lac. Elektrode ST-03 erneuert durch XK03. Reinigungsverfahren für Lac. Elektrode bei Membranwechsel mit Med.Tech. und Anwender besprochen. Anleitung vor Ort gelassen. | Lac. Elektrode ST-03 erneuert durch XK03. Reinigungsverfahren für Lac. Elektrode bei Membranwechsel mit Med.Tech. und Anwender besprochen. Anleitung vor Ort gelassen. | SWO-129451 | No | Lac. Electrode ST-03 renewed by XK03. Cleaning process for Lac. Electrode for membrane changes with Med.Tech. and users. Instructions left on site. | Lac. Electrode ST-03 renewed by XK03. Cleaning process for Lac. Electrode for membrane changes with Med.Tech. and users. Instructions left on site. | Assisted Replacement | Electrode | Electrode | Lac | Zero Current, Contacts black
ABL837 FLEX |  | B | D14FD9F13C2F4023 | No |  |  | Assisted Replacement | Electrode | Membrane | Lac
ABL800 BASIC | cleaned the inlet gasket and sample flow path, done two point calibration and sample all results found with in the range. | cleaned the inlet gasket and sample flow path, done two point calibration and sample all results found with in the range. | SWO-129968 | No | cleaned the inlet gasket and sample flow path, done two point calibration and sample all results found with in the range. | cleaned the inlet gasket and sample flow path, done two point calibration and sample all results found with in the range. | Cleaned | Fluidic pathway | Inlet module | Inlet Gasket | Clot
ABL825 FLEX | CONFIRMED ISSUE WHEN ONSITE. CUSTOMER HAS ALREADY TRIED REMEMBRANED TO NO AVAIL. | REPLACED PCO2 SEALING RING AND ADJUSTED pH BG MODULE SETTINGS. ZERO CURRENT FOR PO2 MET. PCO2 L1 AND L3 NOW RECOVERING AT MEAN VALUES. | D14FD67E3C58F457 | No | CONFIRMED ISSUE WHEN ONSITE. CUSTOMER HAS ALREADY TRIED REMEMBRANED TO NO AVAIL. | REPLACED PCO2 SEALING RING AND ADJUSTED pH BG MODULE SETTINGS. ZERO CURRENT FOR PO2 MET. PCO2 L1 AND L3 NOW RECOVERING AT MEAN VALUES. | Defect | HW | pH/BG Module | pCO2 | Sealing ring
ABL835 FLEX |  | B | D14FB36D3C494021 | No |  |  | Setup | SW | His/Lis
AQT90 FLEX | Küvetten Transport geprüft, Punch Down geprüft, LS Sensor justiert und gereinigt, Funktionskontrolle ok | Küvetten Transport geprüft, Punch Down geprüft, LS Sensor justiert und gereinigt, Funktionskontrolle ok | SWO-129446 | No | Cuvettes transport tested, Punch Down tested, LS sensor adjusted and cleaned, function control ok | Cuvettes transport tested, Punch Down tested, LS sensor adjusted and cleaned, function control ok | Adjustment | HW | Incubator module | Sensor, Cup
ABL80 FLEX Basic version | Substituição da patilha de fecho do cartucho. Instalação do cartucho de soluções. Verificação do correcto funcionamento. | B | D150F88BB2445650 | No | Replacing the cartridge closing lever. Installing the solution cartridge. Verification of proper operation.
ABL80 FLEX CO-OX | ラッチセンサーの分解点検行うもセンサー先端が断線している模様(テスターでOL)センサー届けてもらい交換。 | B | SWO-129682 | No | The disassembly inspection of the latch sensor is performed, but the sensor tip is cut off and the sensor is delivered by the tester and the sensor is exchanged.
ABL90 | CLIENT ADVISED SENSOR BAY DOOR WOULD NOT OPEN OR CLOSE PROPERLY | OSPD FOUND LOOSE ZIFF CABLE CONNECTING TO WETSECTION PCB. RESEATED CABLE, ALL TESTED OK. | D150C19DB261B4A0 | No | CLIENT ADVISED SENSOR BAY DOOR WOULD NOT OPEN OR CLOSE PROPERLY | OSPD FOUND LOOSE ZIFF CABLE CONNECTING TO WETSECTION PCB. RESEATED CABLE, ALL TESTED OK. | Adjustment | HW | wetsection PCB | Sensor cassette door
ABL90 | Customer has been prompted to replace the solution pack and or flush the analyzer several times. Unable to pump solutions errors and failed cal solution errors are observed in the fluid transport service screen. | Performed the fluid transport procedures and re-installed the solution pack. Analyzer properly initialized, however, pCO2 sensitivity is low. Calibrations are failing as a result. After analyzer sat overnight and was allowed to calibrate and properly condition, all parameters came into acceptable calibration range. Performed the necessary calibrations and QC's. Analyzer is within acceptable operating specifications. | D1510AE4B2560B05 | No | Customer has been prompted to replace the solution pack and or flush the analyzer several times. Unable to pump solutions errors and failed cal solution errors are observed in the fluid transport service screen. | Performed the fluid transport procedures and re-installed the solution pack. Analyzer properly initialized, however, pCO2 sensitivity is low. Calibrations are failing as a result. After analyzer sat overnight and was allowed to calibrate and properly condition, all parameters came into acceptable calibration range. Performed the necessary calibrations and QC's. Analyzer is within acceptable operating specifications. | Cleaned | Fluidic pathway | Fluidic pathway | FT
ABL90 FLEX | Removed blood clots blocking sample pathway, also straightened bend in inlet probe. | Removed blood clots blocking sample pathway, also straightened bend in inlet probe. | SWO-129596 | No | Removed blood clots blocking sample pathway, also straightened bend in inlet probe. | Removed blood clots blocking sample pathway, also straightened bend in inlet probe. | Cleaned | Fluidic pathway | Fluidic pathway | clot in Hemolyzer
ABL90 FLEX PLUS | SW neu installiert, alle Daten gelöscht, Gerät wieder konfiguriert, kalibriert, mit QCs überprüft und div. Patientenmessungen durchgeführt, OK Roma.db von 29376KB auf 3856KB verkleinert | SW neu installiert, alle Daten gelöscht, Gerät wieder konfiguriert, kalibriert, mit QCs überprüft und div. Patientenmessungen durchgeführt, OK Roma.db von 29376KB auf 3856KB verkleinert | SWO-129606 | No | SW reinstalled, all data deleted, device reconfigured, calibrated, checked with QCs and various patient measurements performed, OK Roma.db reduced from 29376KB to 3856KB | SW reinstalled, all data deleted, device reconfigured, calibrated, checked with QCs and various patient measurements performed, OK Roma.db reduced from 29376KB to 3856KB | Defect | SW | SW | Hardisk full | Re install SW
ABL90 FLEX PLUS | Servicedump erstellt und versandt | Servicedump erstellt und versandt | SWO-129488 | Yes | Service dump created and shipped | Service dump created and shipped | Re classified RMED | Planned activity | Planned activity | SD request by LEH | Mail 7/9
AQT90 FLEX | 更换910-304后修复by jesson | 更换910-304后修复by jesson | SWO-129469 | No | Repair by jesson after replacement 910-304 | Repair by jesson after replacement 910-304 | Defect | HW | Storage Module | Barcode reader, Storage wheel
ABL800 FLEX | replace the consumable. guide customer for replacement & use. guide for capillary use. | replace the consumable. guide customer for replacement & use. guide for capillary use. | SWO-129915 | No | replace the consumable. guide customer for replacement & use. guide for capillary use. | replace the consumable. guide customer for replacement & use. guide for capillary use. | Assisted Replacement | Electrode | Membrane | Crea | Training, reclassify?
ABL825 FLEX |  | B | D15118F0B227ABE7 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX | Neues AU Board in GS Gerät eingebaut zur Vorbereitung SW 6.16 installiert. Wird später wieder in Kundengerät eingebaut. | Gas 1 Regulator ersetzt. Kal durchgeführt. Gas1 Fehler tritt nicht mehr auf. Drift bei PCo2 PO2 und Gas2 wird bei voller Flasche als leer angezeigt. Valve G2 schaltet nicht. Valve und Kabel ok..24V bei G1 messbar.....0,0 V bei G2 Valve messbar. In GS anderes AU Board geholt und eingebaut. AU PCB ersetzt. Key | SWO-129410 | No | New AU board installed in GS device to prepare SW 6.16. Will be reinstalled later in customer device. | Gas 1 Regulator replaced. Kal. Gas1 error no longer occurs. Drift at PCo2 PO2 and Gas2 is shown as empty at full bottle. Valve G2 does not switch. Valve and cable ok.. 24V at G1 measurable..... 0.0 V measurable at G2 Valve. In GS other AU board fetched and installed. AU PCB. Key | Defect | HW | Gas Module | Gas Regulator | Gas1
ABL800 BASIC | BG側、測定室研磨調整作業 | B | SWO-129618 | No | BG side, measurement chamber polishing adjustment work |  | Cleaned | Fluidic pathway | pH/BG Module | pCO2
ABL800 BASIC | BG測定室研磨作業 | B | SWO-129623 | No | BG measurement chamber polishing work |  | Cleaned | Fluidic pathway | pH/BG Module | pCO2
ABL825 FLEX | Im Servicemenu Oxi kompl. Test durchgeführt, Neon Lamp failed. Oximodul ausgebaut und überprüft. PCB ersetzt, SW 6.16 geladen und Oximetrie kompl. neu eingestellt, ohne Erfolg. Wieder Original Oximodul eingebaut. Neon Lampe und alle Sicherungen auf AU-Board (keine Verbindung DMS mit AU) ersetzt. 2x Oxi kompl. Test durchgeführt (stabile Werte) i.o. Kal., QC und System i.o. Funktionskontrolle i.o. | Im Servicemenu Oxi kompl. Test durchgeführt, Neon Lamp failed. Oximodul ausgebaut und überprüft. PCB ersetzt, SW 6.16 geladen und Oximetrie kompl. neu eingestellt, ohne Erfolg. Wieder Original Oximodul eingebaut. Neon Lampe und alle Sicherungen auf AU-Board (keine Verbindung DMS mit AU) ersetzt. 2x Oxi kompl. Test durchgeführt (stabile Werte) i.o. Kal., QC und System i.o. Funktionskontrolle i.o. | SWO-129496 | No | In the service menu Oxi kompl. Test performed, Neon Lamp failed. Oximodul removed and checked. PCB replaced, SW 6.16 loaded and oximetry complete. re-adjusted, to no avail. Again original Oximodul built-in. Neon lamp and all fuses on AU board (no connection DMS with AU) replaced. 2x Oxi kompl. Test performed (stable values) i.o. Kal., QC and System i.o. Functional control i.o. | In the service menu Oxi kompl. Test performed, Neon Lamp failed. Oximodul removed and checked. PCB replaced, SW 6.16 loaded and oximetry complete. re-adjusted, to no avail. Again original Oximodul built-in. Neon lamp and all fuses on AU board (no connection DMS with AU) replaced. 2x Oxi kompl. Test performed (stable values) i.o. Kal., QC and System i.o. Functional control i.o. | Defect | HW | Oxi module | Neon Lamp | Low
ABL825 FLEX | Testweise eine USB Maus an einen der hinteren USB Ports angeschlossen und die Teamleitung gebeten, im Falle eines einfrierenden Monitors zu prüfen, ob alle Funktionen mit der Maus angewählt werden können. | Testweise eine USB Maus an einen der hinteren USB Ports angeschlossen und die Teamleitung gebeten, im Falle eines einfrierenden Monitors zu prüfen, ob alle Funktionen mit der Maus angewählt werden können. | SWO-129560 | Yes | Test-coded a USB mouse connected to one of the rear USB ports and asked the team management to check in case of a freezing monitor whether all functions can be selected with the mouse. | Test-coded a USB mouse connected to one of the rear USB ports and asked the team management to check in case of a freezing monitor whether all functions can be selected with the mouse. | Re classified RMED | HW | DMS_PC | Touch screen | Planned, Enable of USB Mouse could be done by user
ABL825 FLEX | Testweise eine USB Maus an einen der hinteren USB Ports angeschlossen und die Teamleitung gebeten, im Falle eines einfrierenden Monitors zu prüfen, ob alle Funktionen mit der Maus angewählt werden können. | Testweise eine USB Maus an einen der hinteren USB Ports angeschlossen und die Teamleitung gebeten, im Falle eines einfrierenden Monitors zu prüfen, ob alle Funktionen mit der Maus angewählt werden können. | SWO-129557 | Yes | Test-coded a USB mouse connected to one of the rear USB ports and asked the team management to check in case of a freezing monitor whether all functions can be selected with the mouse. | Test-coded a USB mouse connected to one of the rear USB ports and asked the team management to check in case of a freezing monitor whether all functions can be selected with the mouse. | Re classified RMED | HW | DMS_PC | Touch screen | Planned, Enable of USB Mouse could be done by user
ABL825 FLEX | Testweise eine USB Maus an einen der hinteren USB Ports angeschlossen und die Teamleitung gebeten, im Falle eines einfrierenden Monitors zu prüfen, ob alle Funktionen mit der Maus angewählt werden können. | Testweise eine USB Maus an einen der hinteren USB Ports angeschlossen und die Teamleitung gebeten, im Falle eines einfrierenden Monitors zu prüfen, ob alle Funktionen mit der Maus angewählt werden können. | SWO-129562 | Yes | Test-coded a USB mouse connected to one of the rear USB ports and asked the team management to check in case of a freezing monitor whether all functions can be selected with the mouse. | Test-coded a USB mouse connected to one of the rear USB ports and asked the team management to check in case of a freezing monitor whether all functions can be selected with the mouse. | Re classified RMED | HW | DMS_PC | Touch screen | Planned, Enable of USB Mouse could be done by user
ABL810 FLEX | Vor dem Einbau waren die Pco2- Werte bei Level 2 und 3 häufig unterhalb der Rilibäkgrenzen und sporadisch unterhalb Herstellergrenzen. Zähler: 270 | Vor dem Einbau waren die Pco2- Werte bei Level 2 und 3 häufig unterhalb der Rilibäkgrenzen und sporadisch unterhalb Herstellergrenzen. Zähler: 270 | SWO-129673 | No | Before installation, the Pco2 values at levels 2 and 3 were often below the Rilibäk limits and sporadically below the manufacturer's limits. Counter: 270 | Before installation, the Pco2 values at levels 2 and 3 were often below the Rilibäk limits and sporadically below the manufacturer's limits. Counter: 270 | Assisted Replacement | Electrode | Electrode | pCO2 | 8x -> 8x5, modif
AQT90 FLEX | Enable Auto refresh for stabilise temperature to normal range, AU status ok. | Enable Auto refresh for stabilise temperature to normal range, AU status ok. | SWO-129937 | No | Enable Auto refresh for stabilise temperature to normal range, AU status ok. | Enable Auto refresh for stabilise temperature to normal range, AU status ok. | Power off/on | SW | Incubator module | Incubator module
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-130012 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX BASIC | 更换测试卡，试剂包正常。 | 更换测试卡，试剂包正常。 | SWO-129745 | No | Replace the test card and the reagent package is OK. | Replace the test card and the reagent package is OK.
ABL80 CO-OX | Analyzer not seeing the SP. | Reset the SP and re-installed. | D1526AF32803FF1F | No | Analyzer not seeing the SP. | Reset the SP and re-installed.
ABL80 FLEX BASIC | checked all the parameters are active. Perform the QC. All the value are within the ranges. Also perform the 2 to 3 samples and its values are accepted by physician. | checked all the parameters are active. Perform the QC. All the value are within the ranges. Also perform the 2 to 3 samples and its values are accepted by physician. | SWO-129899 | No | checked all the parameters are active. Perform the QC. All the value are within the ranges. Also perform the 2 to 3 samples and its values are accepted by physician. | checked all the parameters are active. Perform the QC. All the value are within the ranges. Also perform the 2 to 3 samples and its values are accepted by physician.
ABL90 FLEX | Fließwege gespült, soweit frei. Einlassdichtung und Einlasssonde von Anwendern bereits getauscht. Einlassanschlussdichtung erneuert. Überprüfung des Lösungspacks im Service zeigte, Analysator kann keine Lösungen fördern, vmtl. Lösungspack fehlerhaft. Rollenkranz der Förderpumpe ist in Ordnung. Neuen Lösungspack eingesetzt, Fehler behoben, alle Funktionen o. K. Lösungspack reklamiert. Kalibrationen und Qualitätskontrolle in Ordnung. | Fließwege gespült, soweit frei. Einlassdichtung und Einlasssonde von Anwendern bereits getauscht. Einlassanschlussdichtung erneuert. Überprüfung des Lösungspacks im Service zeigte, Analysator kann keine Lösungen fördern, vmtl. Lösungspack fehlerhaft. Rollenkranz der Förderpumpe ist in Ordnung. Neuen Lösungspack eingesetzt, Fehler behoben, alle Funktionen o. K. Lösungspack reklamiert. Kalibrationen und Qualitätskontrolle in Ordnung. | SWO-129879 | No | Flow paths rinsed, as far as free. inlet seal and inlet probe have already been replaced by users. Inlet connection seal renewed. Checking the solution pack in the service showed, Analyzer can not promote solutions, vmtl. solution pack faulty. Roller ring of the pump is fine. New solution pack used, bugs fixed, all functions or solution pack claimed. Calibrations and quality control in order. | Flow paths rinsed, as far as free. inlet seal and inlet probe have already been replaced by users. Inlet connection seal renewed. Checking the solution pack in the service showed, Analyzer can not promote solutions, vmtl. solution pack faulty. Roller ring of the pump is fine. New solution pack used, bugs fixed, all functions or solution pack claimed. Calibrations and quality control in order. | Assisted Replacement | SP | SP | Pump cal error
ABL90 FLEX | Replace the Sensor cassette and found at the time of initialization it shows pressure vaccum error and after that ask for replacement of inlet Gasket holder.Clean the gasket holder,flush waste line and run fluid transport in service but pressure vaccum test is not passing,shows out of range. | Clean all fludic path,clean inlet gasket holder,Replace solution pack but problem remain same.It shows leak detected and pressure vaccum error with replace ment of inlet gasket holder. | SWO-129904 | No | Replace the Sensor cassette and found at the time of initialization it shows pressure vaccum error and after that ask for replacement of inlet Gasket holder. Clean the gasket holder,flush waste line and run fluid transport in service but pressure vaccum test is not passing,shows out of range. | Clean all fludic path,clean inlet gasket holder,Replace solution pack but problem remain same. It shows leak detected and pressure vaccum error with replace ment of inlet gasket holder. | Assisted Replacement | SC | SC | Reference | 1361
ABL90 | Found that there were two bad DOA sensor cassettes and two bad solution packs. | Replaced the sensor cassettes with a sensor cassette from perfusion and a DOA solution pack from RT dept. Replaced DOA solution pack with a new solution pack and sensor cassette from perfusion. | D1522660286028FD | No | Found that there were two bad DOA sensor cassettes and two bad solution packs. | Replaced the sensor cassettes with a sensor cassette from perfusion and a DOA solution pack from RT dept. Replaced DOA solution pack with a new solution pack and sensor cassette from perfusion. | Assisted Replacement | SP | SP | Unknown
ABL90 FLEX | Replaced SP with new, pC02 sensitivity low, customer to monitor and run QCs once the calibration has passed | Replaced SP with new, pC02 sensitivity low, customer to monitor and run QCs once the calibration has passed | SWO-129887 | No | Replaced SP with new, pC02 sensitivity low, customer to monitor and run QCs once the calibration has passed | Replaced SP with new, pC02 sensitivity low, customer to monitor and run QCs once the calibration has passed | Assisted Replacement | SP | SP | Pump tubes
ABL90 FLEX | Transportwege rückwärts gespült, kalibriert und in Betrieb genommen. QC s in allen 3 Leveln bestanden. | Transportwege rückwärts gespült, kalibriert und in Betrieb genommen. QC s in allen 3 Leveln bestanden. | SWO-133082 | No | Transport routes flushed backwards, calibrated and put into operation. QC s passed in all 3 levels. | Transport routes flushed backwards, calibrated and put into operation. QC s passed in all 3 levels. | Cleaned | Fluidic pathway | Fluidic pathway | Clot
ABL90 FLEX PLUS | mitgebrachte SCi (Lot R0430) eingebaut danach Glu- und Lac-Probleme - QC Level A + B zu hoch, fällige tHb Kal durchgeführt Gerät kalibriert und überprüft Falschlieferung 1x Kal1 (944-128 Lot XK02) sowie 1x Reinigungslösung (944-128 Lot SK03) ABL8xx mitgenommenen | mitgebrachte SCi (Lot R0430) eingebaut danach Glu- und Lac-Probleme - QC Level A + B zu hoch, fällige tHb Kal durchgeführt Gerät kalibriert und überprüft Falschlieferung 1x Kal1 (944-128 Lot XK02) sowie 1x Reinigungslösung (944-128 Lot SK03) ABL8xx mitgenommenen | SWO-129813 | No | sci (Lot R0430) built-in after that glu and lac problems - QC Level A + B too high, due tHb Kal carried device calibrated and checked false delivery 1x Kal1 (944-128 Lot XK02) and 1x cleaning solution (944-128 Lot SK03) ABL8xx carried | sci (Lot R0430) built-in after that glu and lac problems - QC Level A + B too high, due tHb Kal carried device calibrated and checked false delivery 1x Kal1 (944-128 Lot XK02) and 1x cleaning solution (944-128 Lot SK03) ABL8xx carried | Assisted Replacement | SC | SC | K | Sensititvity
ABL90 FLEX PLUS |  | B | D152159728399D4D | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D15236A3283C7EAC | No |  |  | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX PLUS | Logged on to analyser and entered the activity logs. Entered the error code that had been investigated and applied the filter to show all instances since my last visit. No more occurrences of this error since I reinstalled the original power supply, performed a number of tests and left the analyser running. Customer also reported that this has not happened since. No further action required for this fault. Whilst at the analyser, I applied a live connect certificate update. Analyser running well again | Logged on to analyser and entered the activity logs. Entered the error code that had been investigated and applied the filter to show all instances since my last visit. No more occurrences of this error since I reinstalled the original power supply, performed a number of tests and left the analyser running. Customer also reported that this has not happened since. No further action required for this fault. Whilst at the analyser, I applied a live connect certificate update. Analyser running well again | SWO-130178 | Yes | Logged on to analyser and entered the activity logs. Entered the error code that had been investigated and applied the filter to show all instances since my last visit. No more occurrences of this error since I reinstalled the original power supply, performed a number of tests and left the analyser running. Customer also reported that this has not happened since. No further action required for this fault. Whilst at the analyser, I applied a live connect certificate update. Analyser running well again | Logged on to analyser and entered the activity logs. Entered the error code that had been investigated and applied the filter to show all instances since my last visit. No more occurrences of this error since I reinstalled the original power supply, performed a number of tests and left the analyser running. Customer also reported that this has not happened since. No further action required for this fault. Whilst at the analyser, I applied a live connect certificate update. Analyser running well again | Re classified RMED | Courtesy
ABL90 FLEX PLUS | Fitted new wet section PCB, however analyser would not recognise ID chip. Still unable to exit operator intervention service programs. | B | SWO-130220 | No | Fitted new wet section PCB, however analyser would not recognise ID chip. Still unable to exit operator intervention service programs. |  | Unknown | HW | Wetsection PCB | Wetsection PCB | ctBil
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-129717 | No | The replacement test card is OK. | The replacement test card is OK.
ABL825 FLEX | kompletter OxiTest -> Halogenintensität niedrig, Hemolyzer demontiert und kontrolliert, Lampeneinheit justiert, Kompletter OxiTest ok, korrekte Kal abgewartet, tHb-Kal ok, mehrere QK ok | kompletter OxiTest -> Halogenintensität niedrig, Hemolyzer demontiert und kontrolliert, Lampeneinheit justiert, Kompletter OxiTest ok, korrekte Kal abgewartet, tHb-Kal ok, mehrere QK ok | SWO-129823 | No | complete OxiTest -> halogen intensity low, hemolyzer disassembled and controlled, lamp unit adjusted, complete OxiTest ok, correct scale awaiting, tHb-Kal ok, several QK ok | complete OxiTest -> halogen intensity low, hemolyzer disassembled and controlled, lamp unit adjusted, complete OxiTest ok, correct scale awaiting, tHb-Kal ok, several QK ok | Adjustment | HW | Oxi module | Lamp unit
ABL800 BASIC | Disassembled, all modules and cables in ABL800, cleaned all modules and cables, cleaned all fluidic valves. Reassembled all modules and cables. 2 point calibration performed without error. sample measurement performed. | Disassembled, all modules and cables in ABL800, cleaned all modules and cables, cleaned all fluidic valves. Reassembled all modules and cables. 2 point calibration performed without error. sample measurement performed. | SWO-129866 | No | Disassembled, all modules and cables in ABL800, cleaned all modules and cables, cleaned all fluidic valves. Reassembled all modules and cables. 2 point calibration performed without error. sample measurement performed. | Disassembled, all modules and cables in ABL800, cleaned all modules and cables, cleaned all fluidic valves. Reassembled all modules and cables. 2 point calibration performed without error. sample measurement performed. | Cleaned | HW | Fluidic manifold | Valves | Possible revisit
ABL800 BASIC | 代替えにて正常稼働していましたが、正規部品が届いた為、交換を実施しました。 | 代替えにて正常稼働していましたが、正規部品が届いた為、交換を実施しました。 | SWO-130276 | No | It was running normally in the alternative, but because the regular parts arrived, the replacement was carried out. | It was running normally in the alternative, but because the regular parts arrived, the replacement was carried out. | Defect | HW | Oxi module | Oxi module | Lidet informativeNot very informative.
ABL835 FLEX |  | B | D1524850285368C2 | No |  |  | Setup | SW | His/Lis | His/Lis
ABL827 FLEX | MET II lower liquid sensor error 535 | Replaced measuring chamber. Also tried a new pump rotor. | D15272A028151DBD | No | MET II lower liquid sensor error 535 | Replaced measuring chamber. Also tried a new pump rotor. | Defect | HW | Met II | Measuring chamber | Most proberly the liquid sensor, or cleaning
ABL800 BASIC | Replace the Lac membrane by following proper procedure,run cleaning and then 2-point calibration and resolve the problem | Replace the Lac membrane by following proper procedure,run cleaning and then 2-point calibration and resolve the problem | SWO-129961 | No | Replace the Lac membrane by following proper procedure,run cleaning and then 2-point calibration and resolve the problem | Replace the Lac membrane by following proper procedure,run cleaning and then 2-point calibration and resolve the problem | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL80 FLEX BASIC | flush sensor cassette . reinstalled sc80. calibration pass. there may be possibility of pre-analytical issue. machine under observation. | calibrate system multiple time. qc pass. guide customer for use. | SWO-131305 | No | flush sensor cassette . reinstalled sc80. calibration pass. there may be possibility of pre-analytical issue. machine under observation. | calibrate system multiple time. qc pass. guide customer for use.
ABL90 FLEX | Clean all fludic path,clean inlet gasket holder,Replace solution pack but problem remain same.It shows leak detected and pressure vaccum error with replace ment of inlet gasket holder. | Clean all fludic path,clean inlet gasket holder,Replace solution pack but problem remain same.It shows leak detected and pressure vaccum error with replace ment of inlet gasket holder. | SWO-129962 | No | Clean all fludic path,clean inlet gasket holder,Replace solution pack but problem remain same. It shows leak detected and pressure vaccum error with replace ment of inlet gasket holder. | Clean all fludic path,clean inlet gasket holder,Replace solution pack but problem remain same. It shows leak detected and pressure vaccum error with replace ment of inlet gasket holder. | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | 更换新试剂包后解决 | 更换新试剂包后解决 | SWO-130029 | No | Resolve after replacing a new reagent kit | Resolve after replacing a new reagent kit | Assisted Replacement | SP | SP | FT | 181
AQT90 FLEX | 更换了905-953后修复(byzhenjie) | 更换了905-953后修复(byzhenjie) | SWO-129677 | No | Repair after replacement of 905-953 (byzhenjie) | Repair after replacement of 905-953 (byzhenjie) | Defect | HW | Aspiration Module | Barcode reader, Inlet
ABL825 FLEX | Felszívótű kiegyenesítése, Bementi egység cseréje, Kalibrálás, QC mérés | B | SWO-129927 | No | Suction needle straightening, input unit replacement, calibration, QC measurement |  | Defect | HW | Inlet module | Probe bend
ABL800 FLEX | Cleaned the PH electrode chamber. Run the Protein removal cycle Replaced the Lac electrode's membrane. Run the 2 point calibration, Now ABL800 is ready for measurement. | B | SWO-129931 | No | Cleaned the PH electrode chamber. Run the Protein removal cycle Replaced the Lac electrode's membrane. Run the 2 point calibration, Now ABL800 is ready for measurement. |  | Cleaned | Fluidic pathway | pH/BG Module | pH | Clot
ABL800 BASIC | ph電極にクロットが付着。先端クリーニング | B | SWO-129932 | No | Clot adheres to ph electrode. Tip cleaning |  | Cleaned | Fluidic pathway | pH/BG Module | pH | Clot on Electrode
AQT90 FLEX | Did the System Cleaning already yesterday. Cleaned the Barcode reader. Run again the same sample but was with same error code as Query. Checked the sample and advised the customer to run the same sample with plasma (after centrifuge), Customer don't have the centrifuge to separate the sample. | B | SWO-129946 | No | Did the System Cleaning already yesterday. Cleaned the Barcode reader. Run again the same sample but was with same error code as Query. Checked the sample and advised the customer to run the same sample with plasma (after centrifuge), Customer don't have the centrifuge to separate the sample. |  | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1223
ABL90 FLEX PLUS | Correctly activated Solution pack and reinstalled Solution pack and passed all tests and QC's | Correctly activated Solution pack and reinstalled Solution pack and passed all tests and QC's | SWO-130086 | No | Correctly activated Solution pack and reinstalled Solution pack and passed all tests and QC's | Correctly activated Solution pack and reinstalled Solution pack and passed all tests and QC's | User error | SP Activation | SP Activation | SP Activation
ABL80 FLEX Basic version |  | B | D15701D3002A7862 | No
ABL80 FLEX BASIC | Replaced another new Solution pack. Downloaded the Details from the machine, Raised the PER. | Replaced another new Solution pack. Downloaded the Details from the machine, Raised the PER. | SWO-131173 | No | Replaced another new Solution pack. Downloaded the Details from the machine, Raised the PER. | Replaced another new Solution pack. Downloaded the Details from the machine, Raised the PER.
ABL90 FLEX | Gerät überprüft, Sensorkassette ausgetauscht, tHb Kal durchgeführt | Gerät überprüft, Sensorkassette ausgetauscht, tHb Kal durchgeführt | SWO-130340 | No | Device checked, sensor cartridge replaced, tHb Kal performed | Device checked, sensor cartridge replaced, tHb Kal performed | Assisted Replacement | SC | SC | pCO2
ABL90 | Error 963-leak currents detected. Error occurs intermittently. | Replaced tubings and connectors. Replaced hemolyzer | D155FD068A167619 | No | Error 963-leak currents detected. Error occurs intermittently. | Replaced tubings and connectors. Replaced hemolyzer | Defect | HW | Oxi module | Hemolyzer | Leak Current
ABL90 FLEX | XL Lösungpack 2x geliefert | XL Lösungpack 2x geliefert | SWO-130161 | Yes | XL solution pack 2x delivered | XL solution pack 2x delivered | Re classified RMED | Logistical upSWO
ABL90 FLEX | erneut mehrfach kalibriert, QC abgewartet, Level1 und 2 haben Grenzen erreicht und sind ok - Level 3 nur noch bei Lac ausser Bereich, wird weiter beobachtet und gegebenenfalls meldet sich MedTech erneut | erneut mehrfach kalibriert, QC abgewartet, Level1 und 2 haben Grenzen erreicht und sind ok - Level 3 nur noch bei Lac ausser Bereich, wird weiter beobachtet und gegebenenfalls meldet sich MedTech erneut | SWO-130159 | No | recalibrated several times, QC waited, level 1 and 2 have reached limits and are ok - level 3 only with Lac except area, is further observed and if necessary MedTech reports again | recalibrated several times, QC waited, level 1 and 2 have reached limits and are ok - level 3 only with Lac except area, is further observed and if necessary MedTech reports again | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS | Vérification de la position du module d'entrée, débouchage de l'hemolyseur avec de l'eau chaude(on sent une suppression importante quand on met une pression sur la syrinque) Un petit caillot sort finalement de l'hemolyseur. Mis en route : OK | Vérification de la position du module d'entrée, débouchage de l'hemolyseur avec de l'eau chaude(on sent une suppression importante quand on met une pression sur la syrinque) Un petit caillot sort finalement de l'hemolyseur. Mis en route : OK | SWO-130518 | No | Checking the position of the input module, unblocking the hemolyseur with hot water (you feel a significant deletion when you put pressure on the syrinch) A small clot finally comes out of the hemolyseur. Started: OK | Checking the position of the input module, unblocking the hemolyseur with hot water (you feel a significant deletion when you put pressure on the syrinch) A small clot finally comes out of the hemolyseur. Started: OK | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS |  | B | D155F9598A10F7F2 | Yes |  |  | Re classified RMED | PER
ABL90 FLEX PLUS |  | B | D155E0EC8A16D30F | No |  |  | Assisted Replacement | SC | SC | Reference
ABL90 | Chequeo de equipo, sustitución de inlet gasket, sonda de aspiración, procedimiento de limpieza de circuito neumático, reinstalación de SP, ejecución de todos los programas de servicio, inicio completo, calibración y pasados QC, cumplimentación de parte de trabajo. | Sustitución inlet gasket, sonda, limpieza completa. ejecución progaramas de servicio | D15611F28A558F42 | No | Equipment check, gasket inlet replacement, suction probe, pneumatic circuit cleaning procedure, SP reinstallation, execution of all service programs, full start, calibration and QC pasts, completion of part of work. | Tued inlet gasket, probe, complete cleaning. execution progaramas of service | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX |  | B | SWO-130001 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | Se compueba que la storage wheel en la posición 7 no se alinea con el punh up. Pendiente de piezas | B | D15637B28A133A2A | No | It was consistent that the storage wheel at position 7 does not align with the punh up. Earrings |  | mail |  |  | Not Cat.
ABL800 FLEX | Cleaned the autocheck barcode reader & loaded the ampoule checked equipment is functional | Cleaned the autocheck barcode reader & loaded the ampoule checked equipment is functional | SWO-131416 | No | Cleaned the autocheck barcode reader & loaded the ampoule checked equipment is functional | Cleaned the autocheck barcode reader & loaded the ampoule checked equipment is functional | Cleaned | HW | AutoCheck Module | Barcode reader
ABL825 FLEX | Troubleshot and found that the flat membrane valve on the pH BG module was sticking. | Replaced both valves on the pH BG module and this resolved the problem. | D15604B78A274D46 | No | Troubleshot and found that the flat membrane valve on the pH BG module was sticking. | Replaced both valves on the pH BG module and this resolved the problem. | Defect | HW | pH/BG Module | Valves
ABL825 FLEX | Allgemeine Kontrolle ok. Oxy Test (Lampenwert auf 3218 gesunken Oxy PCB ersetzt und SW 6.16 geladen. Oxy Test wiederholt (Lampenwert 3601). Mehrere Kal durchgeführt ok. Kal. + QC i.o. | Allgemeine Kontrolle ok. Oxy Test (Lampenwert auf 3218 gesunken Oxy PCB ersetzt und SW 6.16 geladen. Oxy Test wiederholt (Lampenwert 3601). Mehrere Kal durchgeführt ok. Kal. + QC i.o. | SWO-130150 | No | General control ok. Oxy Test (lamp value dropped to 3218 Oxy PCB replaced and SW 6.16 charged. Oxy test (lamp value 3601). Multiple Kal performed ok. Kal. + QC i.o. | General control ok. Oxy Test (lamp value dropped to 3218 Oxy PCB replaced and SW 6.16 charged. Oxy test (lamp value 3601). Multiple Kal performed ok. Kal. + QC i.o. | Defect | HW | Oxi module | Halogene | Coluld be adjusted Changed OXI PCB
ABL825 FLEX | pCO2 und Ref. - Membran mit aktuellen Lot-Nummern bezogen, Messkammer gereinigt, Gerät kalibriert, Qc´s durchgef. | pCO2 und Ref. - Membran mit aktuellen Lot-Nummern bezogen, Messkammer gereinigt, Gerät kalibriert, Qc´s durchgef. | SWO-130130 | No | pCO2 and Ref. - Membrane covered with current lot numbers, measuring chamber cleaned, device calibrated, Qc's through. | pCO2 and Ref. - Membrane covered with current lot numbers, measuring chamber cleaned, device calibrated, Qc's through. | Assisted Replacement | Electrode | Membrane | pCO2 | Could be Ref
ABL810 FLEX | Abdeckung und Schlauchführung abgebaut, komplett gereinigt und Gerinsell am Ablauf entfernt, Schläuche montiert, Gerät zusammengesetzt, Kal und QC, Reservedichtungen für Elektroden hinterlegt | Abdeckung und Schlauchführung abgebaut, komplett gereinigt und Gerinsell am Ablauf entfernt, Schläuche montiert, Gerät zusammengesetzt, Kal und QC, Reservedichtungen für Elektroden hinterlegt | SWO-130160 | No | Cover and hose guide dismantled, completely cleaned and clot removed at the drain, hoses mounted, device assembled, Kal and QC, reserve seals for electrodes deposited | Cover and hose guide dismantled, completely cleaned and clot removed at the drain, hoses mounted, device assembled, Kal and QC, reserve seals for electrodes deposited | Unknown | HW | - | -
ABL825 FLEX |  | B | D15653668A17E0B8 | No |  |  | Defect | HW | Basic Unit | Waste tubes | Clotted
ABL800 BASIC | pH clot detection option turned off in analyzing unit set up, 2 point calibration performed successfully, sample measurements performed without errors | pH clot detection option turned off in analyzing unit set up, 2 point calibration performed successfully, sample measurements performed without errors | SWO-130465 | Yes | pH clot detection option turned off in analyzing unit set up, 2 point calibration performed successfully, sample measurements performed without errors | pH clot detection option turned off in analyzing unit set up, 2 point calibration performed successfully, sample measurements performed without errors | Re classified RMED | Installation
AQT90 FLEX |  | B | D1561B7D8A1FBF5D | No |  |  | Adjustment | HW | Cover and Chassis | Inlet Door
AQT90 FLEX | Bij nader onderzoek blijkt dat de storagewheel barcodereader opnieuw afgesteld moet worden. Na afstelling en controle, OK. | Bij nader onderzoek blijkt dat de storagewheel barcodereader opnieuw afgesteld moet worden. Na afstelling en controle, OK. | SWO-130196 | No | Further investigation shows that the Storagewheel barcode reader needs to be adjusted again. After adjustment and control, OK. | Further investigation shows that the Storagewheel barcode reader needs to be adjusted again. After adjustment and control, OK. | Adjustment | Hw | Storage Module | Barcode reader, Storage wheel
ABL800 BASIC | Replaced CO2 Membrane, Cleaned Electrode chamber, Problem remains need to replace CO2 Electrode. | B | SWO-130612 | No | Replaced CO2 Membrane, Cleaned Electrode chamber, Problem remains need to replace CO2 Electrode. |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL80 FLEX BASIC | 更换试剂包后修复 | 更换试剂包后修复 | SWO-130351 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit
ABL80 CO-OX |  | B | D15736780034B21F | No
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-133527 | No | Replacement test card repair | Replacement test card repair
ABL90 |  | B | D1572EDD0014D431 | No |  |  | Defect | SW | SW | Restart | 734
ABL90 FLEX | Kunde hat gestern SC ersetzt, danach Lac QC Fehler Level C, bei Rückruf heute morgen ist alles i.O., Kunden über aktuelle Situation der SC & LP Qualität informiert, auf regelmäßig zeitnahe Nachbestellung des ZVBM hingewiesen, CCF ausgedrückt, Rückruf von Pfl. Alexander am 13.09. | Kunde hat gestern SC ersetzt, danach Lac QC Fehler Level C, bei Rückruf heute morgen ist alles i.O., Kunden über aktuelle Situation der SC & LP Qualität informiert, auf regelmäßig zeitnahe Nachbestellung des ZVBM hingewiesen, CCF ausgedrückt, Rückruf von Pfl. Alexander am 13.09. | SWO-130358 | Yes | Customer replaced yesterday SC, then Lac QC error level C, at recall this morning everything is usually, customers informed about current situation of SC & LP quality, pointed out to regularly timely reordering of the ZVBM, CCF expressed, recall of Pfl. Alexander on 13.09. | Customer replaced yesterday SC, then Lac QC error level C, at recall this morning everything is usually, customers informed about current situation of SC & LP quality, pointed out to regularly timely reordering of the ZVBM, CCF expressed, recall of Pfl. Alexander on 13.09. | Re classified RMED | Hot line call
ABL90 | flushed - drops comming out of inlet with a lot of back pressure service showing clot in inlet section - flushed hemolyzer flushed sensor bay flushed probe flushed cassette | replaced waste line hemolyzer tube connector gasket and inlet gasket | D1574023003A58A8 | No | flushed - drops comming out of inlet with a lot of back pressure service showing clot in inlet section - flushed hemolyzer flushed sensor bay flushed probe flushed cassette | replaced waste line hemolyzer tube connector gasket and inlet gasket | Defect | HW | PM Kit | PM Kit
ABL90 |  | B | D15762970019AD29 | Yes |  |  | Re classified RMED | Installation
ABL90 |  | B | D15763E00025AD2B | Yes |  |  | Re classified RMED | Installation
ABL90 FLEX PLUS | IP Adresse Def. Gateway nach Angabe durch Frau Essers eingestellt, trotzdem keine Radiance Verbindung => vermutlich ist die Netzwerkdose nicht gepatched, Frau Essers kümmert sich darum. Sensorkassette ist angegeben mit zuerst eingesetzt 25.08.2016 (wurde eventuell testweise an diesem Datum eingebaut?), daher wird das ABlaufdatum im ABL korrekt mit 24.09.2016 angegeben. | IP Adresse Def. Gateway nach Angabe durch Frau Essers eingestellt, trotzdem keine Radiance Verbindung => vermutlich ist die Netzwerkdose nicht gepatched, Frau Essers kümmert sich darum. Sensorkassette ist angegeben mit zuerst eingesetzt 25.08.2016 (wurde eventuell testweise an diesem Datum eingebaut?), daher wird das ABlaufdatum im ABL korrekt mit 24.09.2016 angegeben. | SWO-130418 | No | IP address Def. Gateway set by Ms. Essers, still no radiance connection => presumably the network box is not patched, Ms. Essers takes care of it. Sensor cassette is specified with first inserted 25.08.2016 (was installed test-based on this date?), therefore the OBl date is correctly indicated in the ABL as 24.09.2016. | IP address Def. Gateway set by Ms. Essers, still no radiance connection => presumably the network box is not patched, Ms. Essers takes care of it. Sensor cassette is specified with first inserted 25.08.2016 (was installed test-based on this date?), therefore the OBl date is correctly indicated in the ABL as 24.09.2016. | User error | SC | SC | SC
ABL90 FLEX PLUS |  | B | D157853F00478FEC | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Gerät geprüft, LP und SC ausgetauscht, Funktionskontrolle ok | Gerät geprüft, LP und SC ausgetauscht, Funktionskontrolle ok | SWO-130656 | No | Device checked, LP and SC replaced, function control ok | Device checked, LP and SC replaced, function control ok | Assisted Replacement | SP | SP | pCO2
ABL80 FLEX BASIC | チューブキット交換、マニフォールド分解掃除。センサーカセット交換等実施。 | チューブキット交換、マニフォールド分解掃除。センサーカセット交換等実施。 | SWO-130354 | No | Tube kit replacement, manifold disassembly cleaning. Sensor cassette replacement, etc. | Tube kit replacement, manifold disassembly cleaning. Sensor cassette replacement, etc.
ABL90 FLEX | 更换测试卡，834-662，924-816，924-455各一个机器仍然不能使用，更换测试卡后正常。by jesson | 更换测试卡，834-662，924-816，924-455各一个机器仍然不能使用，更换测试卡后正常。by jesson | SWO-131471 | No | Replace the test card, 834-662, 924-816, 924-455 each machine is still not available, after replacing the test card normal. by jesson | Replace the test card, 834-662, 924-816, 924-455 each machine is still not available, after replacing the test card normal. by jesson | Assisted Replacement | SC | SC | SC
AQT90 FLEX |  | B | SWO-130361 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | Sustitucion del tube holder. Desobtruir el tubo 5. Chequeo por servicio. Realizacion de limpieza. Comprobacion con qc y muestras. | B | D157858E0014A2CF | No | Replacement of the tube holder. Unobstruct tube 5. Check for service. Cleaning. Check with qc and samples. |  | Defect | HW | Aspiration Module | Tube holder
ABL835 FLEX | Nach Erhalt AU Board neu installiert. Ersatzboard wieder ausgebaut. SW und Key eingelesen. Siehe SWO 129410 | Nach Erhalt AU Board neu installiert. Ersatzboard wieder ausgebaut. SW und Key eingelesen. Siehe SWO 129410 | SWO-130608 | No | After receiving AU board reinstalled. Replacement board removed again. SW and Key. See SWO 129410 | After receiving AU board reinstalled. Replacement board removed again. SW and Key. See SWO 129410 | Defect | HW | Gas Module | Flow regulator 1 | Revisit
ABL820 FLEX | I spoke with Teri and she said the pH well is holding fluid. She spoke with tech support and was told they would send her a seal ring to replace on the pH electrode. She said that they had not received the seal ring yet. They took the pH electrode from their old analyzer and the pH is working fine. I will send her a replacement electrode under warranty and will send her electrode back. | B | D1572BD8003A4438 | Yes | I spoke with Teri and she said the pH well is holding fluid. She spoke with tech support and was told they would send her a seal ring to replace on the pH electrode. She said that they had not received the seal ring yet. They took the pH electrode from their old analyzer and the pH is working fine. I will send her a replacement electrode under warranty and will send her electrode back. |  | Re classified RMED | Electrode | Electrode | pH | No Visit
ABL835 FLEX | Proteinentfernung durchgeführt - Messkammer gereinigt Lactat Elektrode erneuert Gerät kalibriert und überprüft | Proteinentfernung durchgeführt - Messkammer gereinigt Lactat Elektrode erneuert Gerät kalibriert und überprüft | SWO-130512 | No | Protein removal performed - measuring chamber cleaned lactate electrode renewed device calibrated and checked | Protein removal performed - measuring chamber cleaned lactate electrode renewed device calibrated and checked | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL800 BASIC | PH電極内測定チャンバー内血液詰まり除去。PH電極洗浄。比較電極メンブラン交換、インレットガスケット交換、電極モジュール用ポンプチューブ交換。 | PH電極内測定チャンバー内血液詰まり除去。PH電極洗浄。比較電極メンブラン交換、インレットガスケット交換、電極モジュール用ポンプチューブ交換。 | SWO-134843 | No | Blood clogging removal in the measurement chamber in the PH electrode. PH electrode cleaning. Comparison electrode membrane replacement, inlet gasket replacement, pump tube replacement for electrode modules. | Blood clogging removal in the measurement chamber in the PH electrode. PH electrode cleaning. Comparison electrode membrane replacement, inlet gasket replacement, pump tube replacement for electrode modules. | Defect | Electrode | Electrode | pH | Sealing ring
ABL825 FLEX | Überprüfung: Oberer Dichtring zwischen Referenz - Messkammer und Thermoblock war offenbar ab Werk nicht korrekt in Position, dadurch Blockade der Elektrode => Block demontiert und Dichtring neu positioniert. Einlassnadel mittig nach oben geknickt (Anwenderfehler!) => Nadel ersetzt. Komplettes Startverfahren mit Kal und QC Messungen => OK | Überprüfung: Oberer Dichtring zwischen Referenz - Messkammer und Thermoblock war offenbar ab Werk nicht korrekt in Position, dadurch Blockade der Elektrode => Block demontiert und Dichtring neu positioniert. Einlassnadel mittig nach oben geknickt (Anwenderfehler!) => Nadel ersetzt. Komplettes Startverfahren mit Kal und QC Messungen => OK | SWO-130511 | No | Check: Upper sealing ring between reference - measuring chamber and thermoblock was apparently not correctly positioned ex works, thus blocking the electrode => block disassembled and sealing ring repositioned. Inlet needle bent upwards in the middle (user error!) => Needle replaced. Complete start procedure with Kal and QC measurements => OK | Check: Upper sealing ring between reference - measuring chamber and thermoblock was apparently not correctly positioned ex works, thus blocking the electrode => block disassembled and sealing ring repositioned. Inlet needle bent upwards in the middle (user error!) => Needle replaced. Complete start procedure with Kal and QC measurements => OK | Defect | HW | pH/BG Module | Temoblock Seal
ABL825 FLEX | TS erneut eingefroren, Bedienung teilweise gar nicht oder nur mit starker zeitlicher Verzögerung möglich, Bedienung per Fernsupport möglich wie gewohnt. Nach Abwischen des (nur unwesentlich verschmutzten!) unteren Bildschirmrands wieder normale Funktion! | Monitor erneuert. | SWO-130499 | No | TS frozen again; Operation partially not possible at all or only with a long time delay; Remote support operation possible as usual. After wiping the (only slightly polluted!) bottom of the screen back to normal function! | Monitor renewed. | Defect | HW | DMS_PC | Touch screen | Second time Cleaned working
ABL825 FLEX | Upper Liquid Sensor im EL MET-Modul erneuert | Upper Liquid Sensor im EL MET-Modul erneuert | SWO-130558 | No | Upper Liquid Sensor renewed in EL MET module | Upper Liquid Sensor renewed in EL MET module | Defect | HW | El/Met Module | Upper LS
ABL825 FLEX | CONFIRMED ISSUE WHEN ONSITE. | PROCEED TO REPLACE pH BG LOWER LIQUID SENSOR, pH ELECTRODE, BOTH MEMBRANE VALVES ABOVE ROTOR, INLET NEEDLE, INLET GASKET, INLET LOWER SENSOR TUBING, CONNECTORS ON INLET MODULES AND TUBING BETW INLET MODULE AND pH BG MODULE. ALSO CHANGED OUT PC02 SEALING RING. | D1573581002BDCAB | No | CONFIRMED ISSUE WHEN ONSITE. | PROCEED TO REPLACE pH BG LOWER LIQUID SENSOR, pH ELECTRODE, BOTH MEMBRANE VALVES ABOVE ROTOR, INLET NEEDLE, INLET GASKET, INLET LOWER SENSOR TUBING, CONNECTORS ON INLET MODULES AND TUBING BETW INLET MODULE AND pH BG MODULE. ALSO CHANGED OUT PC02 SEALING RING. | Defect | HW | pH/BG Module | Valves
ABL80 FLEX BASIC | calibrate system multiple time. qc pass. guide customer for use. | calibrate system multiple time. qc pass. guide customer for use. | SWO-131306 | No | calibrate system multiple time. qc pass. guide customer for use. | calibrate system multiple time. qc pass. guide customer for use.
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-132200 | No | Replacement test card repair | Replacement test card repair
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-130722 | No | Replacement test card repair | Replacement test card repair
ABL80 FLEX BASIC | Solution pack need replacement. Install the new SP. After install the SP unit is working properly. | Solution pack need replacement. Install the new SP. After install the SP unit is working properly. | SWO-131275 | No | Solution pack need replacement. Install the new SP. After install the SP unit is working properly. | Solution pack need replacement. Install the new SP. After install the SP unit is working properly.
ABL80 CO-OX | Red fluid leaking from analyzer. Solution pack leaking around waste nipple. | Replaced waste connector and SP. | D15835BF7661DDCF | No | Red fluid leaking from analyzer. Solution pack leaking around waste nipple. | Replaced waste connector and SP.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-131692 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL80 FLEX | Done the zeroing calibration. All values are found to be out of range. Also corrosion was found on the connector board pins. So need to replace the sensor connector board. | Done the zeroing calibration. All values are found to be out of range. Also corrosion was found on the connector board pins. So need to replace the sensor connector board. | SWO-131336 | No | Done the zeroing calibration. All values are found to be out of range. Also corrosion was found on the connector board pins. So need to replace the sensor connector board. | Done the zeroing calibration. All values are found to be out of range. Also corrosion was found on the connector board pins. So need to replace the sensor connector board.
ABL90 |  | B | D1589CA2761DAD2C | Yes |  |  | Re classified RMED | Training
ABL90 FLEX PLUS | Done back flush x 5 on analyzer and ran calibration x3 all ok. Ran Qc's and all ok. Also advised customer of what was done. | Done back flush x 5 on analyzer and ran calibration x3 all ok. Ran Qc's and all ok. Also advised customer of what was done. | SWO-131046 | No | Done back flush x 5 on analyzer and ran calibration x3 all ok. Ran Qc's and all ok. Also advised customer of what was done. | Done back flush x 5 on analyzer and ran calibration x3 all ok. Ran Qc's and all ok. Also advised customer of what was done. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-130646 | No | The replacement test card is OK. | The replacement test card is OK.
ABL825 FLEX |  | B | D158837076384A34 | Yes |  |  | Re classified RMED | Unknown
ABL800 BASIC | Na電極交換。コンタミ除去実施。Clメンブラン以外のメンブラン定期交換。全ポンプチューブ交換。測定チャンバー、エアーライン洗浄。 | Na電極交換。コンタミ除去実施。Clメンブラン以外のメンブラン定期交換。全ポンプチューブ交換。測定チャンバー、エアーライン洗浄。 | SWO-130860 | No | Na electrode replacement. Contamination removal was carried out. Regular exchange of membranes other than Cl-Menbran. All pump tube replacement. Measuring chamber, air line cleaning. | Na electrode replacement. Contamination removal was carried out. Regular exchange of membranes other than Cl-Menbran. All pump tube replacement. Measuring chamber, air line cleaning. | Assisted Replacement | Electrode | Electrode | Na+
ABL835 FLEX | Neues Oxi PCB eingebaut SW laden Kal und QC | B | SWO-130788 | No | Neues Oxi PCB eingebaut SW laden Kal und QC |  | Defect | HW | Oxi module | Neon Lamp | Too Low
ABL825 FLEX | LS und PP überprüft, Schlauch PP BG-Modul recht platt, erneuert, intensiv gespült, Messkammer u Elektroden geprüft, Einsaugung scheint OK Kunde muss weiterhin beobachten | LS und PP überprüft, Schlauch PP BG-Modul recht platt, erneuert, intensiv gespült, Messkammer u Elektroden geprüft, Einsaugung scheint OK Kunde muss weiterhin beobachten | SWO-130727 | No | LS and PP checked, hose PP BG module quite flat, renewed, intensively rinsed, measuring chamber u electrodes tested, suction seems OK customer must continue to observe | LS and PP checked, hose PP BG module quite flat, renewed, intensively rinsed, measuring chamber u electrodes tested, suction seems OK customer must continue to observe | Defect | HW | pH/BG Module | Pump tubes
ABL825 FLEX | Monitor erneuert. | Monitor erneuert. | SWO-130753 | No | Monitor renewed. | Monitor renewed. | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX | Schoonmaken en doorspoelen van naald en needle wash unit. | Schoonmaken en doorspoelen van naald en needle wash unit. | SWO-130674 | No | Cleaning and flushing of needle and needle wash unit. | Cleaning and flushing of needle and needle wash unit. | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1223,1258,1259
ABL80 FLEX | Checked waste motor in Diagnosis and Physically, Pump not working, cleaned waste path and waste Tubing and reinstalled, Problem remains. roller wheel lock broken tried to fixed but problem not resolve, Need to replace waste motor pump and roller wheel with new one. | Checked waste motor in Diagnosis and Physically, Pump not working, cleaned waste path and waste Tubing and reinstalled, Problem remains. roller wheel lock broken tried to fixed but problem not resolve, Need to replace waste motor pump and roller wheel with new one. | SWO-131332 | No | Checked waste motor in Diagnosis and Physically, Pump not working, cleaned waste path and waste Tubing and reinstalled, Problem remains. roller wheel lock broken tried to fixed but problem not resolve, Need to replace waste motor pump and roller wheel with new one. | Checked waste motor in Diagnosis and Physically, Pump not working, cleaned waste path and waste Tubing and reinstalled, Problem remains. roller wheel lock broken tried to fixed but problem not resolve, Need to replace waste motor pump and roller wheel with new one.
ABL90 FLEX | Überprüfung: pCO2 Status 71,5 mV, Empf. 87,1 % - LP ersetzt, AQC-Fehler bleibt, Einlassdichtung Halter Einlasssonde ersetzt, Fluid Transport - i.O., Fehler bleiben, SC ersetzt (Kundenbestand), Startphase abgewartet, Kal i.O., | Überprüfung: pCO2 Status 71,5 mV, Empf. 87,1 % - LP ersetzt, AQC-Fehler bleibt, Einlassdichtung Halter Einlasssonde ersetzt, Fluid Transport - i.O., Fehler bleiben, SC ersetzt (Kundenbestand), Startphase abgewartet, Kal i.O., | SWO-130954 | No | Review: pCO2 status 71.5 mV, Empf. 87.1% - LP replaced, AQC error remains, inlet seal holder inlet probe replaced, fluid transport - i.o., error remain, SC replaced (customer stock), start phase awaited, Kal i.O., | Review: pCO2 status 71.5 mV, Empf. 87.1% - LP replaced, AQC error remains, inlet seal holder inlet probe replaced, fluid transport - i.o., error remain, SC replaced (customer stock), start phase awaited, Kal i.O., | Assisted Replacement | SC | SC | pCO2
ABL90 |  | B | D159EE56EC21EB5F | Yes |  |  | Re classified RMED | Training
ABL90 | Upon arrival, reviewed all data on unit. Unable to find correct Cap Proficiency run on Date stated of 08-24-2016. All data reviewed and unit functioning correctly on that Date and all days since. | No correction needed. Downloaded all Data logs and will start a Techline case so all data can be reviewed and response can be given to customer. | D1599103EC35E148 | No | Upon arrival, reviewed all data on unit. Unable to find correct Cap Proficiency run on Date stated of 08-24-2016. All data reviewed and unit functioning correctly on that Date and all days since. | No correction needed. Downloaded all Data logs and will start a Techline case so all data can be reviewed and response can be given to customer. | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | Replace the complete Inlet module and check for pressure vaccum now it's working fine and reinstall the SC and found that in QC PO2 is showing error of Inhomogenious sample | Replace the complete Inlet module and check for pressure vaccum now it's working fine and reinstall the SC and found that in QC PO2 is showing error of Inhomogenious sample | SWO-132854 | No | Replace the complete Inlet module and check for pressure vaccum now it's working fine and reinstall the SC and found that in QC PO2 is showing error of Inhomogenious sample | Replace the complete Inlet module and check for pressure vaccum now it's working fine and reinstall the SC and found that in QC PO2 is showing error of Inhomogenious sample | Defect | HW | Inlet module | Inlet module
ABL90 |  | B | D159B891EC21AD30 | Yes |  |  | Re classified RMED | Setup/Training
ABL90 |  | B | D159BA61EC4CAD31 | Yes |  |  | Re classified RMED |  |  |  | LEH Mail
ABL90 FLEX | Akku und Netzteil (neu: P0018N894, alt: P0020N812) getauscht. Messung von Schutzleiter zu Masse ok. Kalibration und QC's vorgenommen. | Akku und Netzteil (neu: P0018N894, alt: P0020N812) getauscht. Messung von Schutzleiter zu Masse ok. Kalibration und QC's vorgenommen. | SWO-131112 | No | Battery and power supply (new: P0018N894, old: P0020N812). Measurement of protective conductor to ground ok. calibration and QC's. | Battery and power supply (new: P0018N894, old: P0020N812). Measurement of protective conductor to ground ok. calibration and QC's. | Defect | HW | Power Supply | Power Supply
AQT90 FLEX | 重新定标正常。 | B | SWO-130313 | No | Recalibration is normal. |  | mail |  |  | Not Cat.
AQT90 FLEX | 更换902-924一个正常。 | B | SWO-130644 | No | Replace the 902-924 one normal. |  | Defect | HW | Analyzer Control Module | Barcode reader, display
ABL825 FLEX | Temperaturegulierung angepasst und über mehrere Zyklen überprüft , mehrere Kal, erneut Fehler, neue Platine getestet - ohne Erfolg, Rep abgebrochen, Material und Ersatzteile organisiert, erneut Modul ausgebaut und Temperaturfühler ersetzt, Neustart und Regulierung überprüft, i.O., mehrere Kal, QC | Zeitweise Fehlermeldung Temperaturfehler im E-Lyte Modul, Kunde wünscht überprüfung | SWO-131114 | No | Temperature angulation adjusted and checked over several cycles, multiple cal, re-made faults, new board tested - without success, rep aborted, material and spare parts organized, re-expanded module and temperature sensor replaced, restart and regulation checked, i.O., several Kal, QC | Intermittent error message temperature error in E-Lyte module, customer wishes review | Defect | HW | pH/BG Module | Termistor
ABL837 FLEX | 電極次亜洗浄 チャンバー清掃 ダミー先端結晶有の為清掃 メンブラン交換 | 電極次亜洗浄 チャンバー清掃 ダミー先端結晶有の為清掃 メンブラン交換 | SWO-132669 | No | Electrode secondary cleaning chamber cleaning dummy tip crystal having cleaning membrane replacement | Electrode secondary cleaning chamber cleaning dummy tip crystal having cleaning membrane replacement | Cleaned | Fluidic pathway | Fluidic pathway | Crea
ABL825 FLEX | Customer has changed membranes | Swapped reference electrode. Did not resolve. Replaced reference connector. Did not resolveNoticed tubing connected to modules loose. Replaced all pump tubing. Cleaned el met chamber. Reseated all electrodes | D159F6D6EC017051 | No | Customer has changed membranes | Swapped reference electrode. Did not resolve. Replaced reference connector. Did not resolveNoticed tubing connected to modules loose. Replaced all pump tubing. Cleaned el met chamber. Reseated all electrodes | Defect | HW | pH/BG Module | Pump tubes
ABL835 FLEX | Oxy-Complete-Test, Halogenlampe justiert, Lac-Elektrode membraniert, tHb-Kal durchgeführt, System und Funktion i.O. | Oxy-Complete-Test, Halogenlampe justiert, Lac-Elektrode membraniert, tHb-Kal durchgeführt, System und Funktion i.O. | SWO-130980 | No | Oxy-Complete-Test, halogen lamp adjusted, Lac-electrode membraneed, tHb-Kal performed, system and function i.O. | Oxy-Complete-Test, halogen lamp adjusted, Lac-electrode membraneed, tHb-Kal performed, system and function i.O. | Adjustment | HW | Oxi module | Lamp unit
TCM-COMBIM 54 | install of Combi M54 | install of Combi M54 | D15ACA766238AE3E | No | install of Combi M54 | install of Combi M54
ABL80 FLEX BASIC | sc80 fail to calibration.Recalibrate the system. fail to calibrate .installed new sc80 & calibrate the system. | B | SWO-131299 | No | sc80 fail to calibration. Recalibrate the system. fail to calibrate .installed new sc80 & calibrate the system.
ABL1 | Placed SC In Fridge for cooling, cleaned Dust from Sensor connector and Signal Data board and tried to Reinstalled.Problem remains. | Placed SC In Fridge for cooling, cleaned Dust from Sensor connector and Signal Data board and tried to Reinstalled.Problem remains. | SWO-131337 | No | Placed SC In Fridge for cooling, cleaned Dust from Sensor connector and Signal Data board and tried to Reinstalled.Problem remains. | Placed SC In Fridge for cooling, cleaned Dust from Sensor connector and Signal Data board and tried to Reinstalled.Problem remains.
ABL90 FLEX | Remplacement de la carte cpu et échangé la carte mémoire afin de sauvegarder la base de données. Remis en route : OK | Remplacement de la carte cpu et échangé la carte mémoire afin de sauvegarder la base de données. Remis en route : OK | SWO-131656 | No | Replacing the cpu card and swapped the memory card to back up the database. Started again: OK | Replacing the cpu card and swapped the memory card to back up the database. Started again: OK | Defect | HW | DMS_PC | CPU
ABL90 FLEX | Parameter PCO2 in der Konfiguration entsperrt, Analysator geht direkt komplett in Grün Status | Parameter PCO2 in der Konfiguration entsperrt, Analysator geht direkt komplett in Grün Status | SWO-131267 | No | Parameter PCO2 unlocked in configuration, analyzer goes directly completely into green status | Parameter PCO2 unlocked in configuration, analyzer goes directly completely into green status | User error | Setup | Parameter | Locked
ABL90 |  | B | D15B0804625581A7 | Yes |  |  | Re classified RMED | Installation | Remote support
ABL90 FLEX PLUS |  | B | D15AFB0D62098B71 | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Error found if users keep the syringe pushed on the probe to long and dose not remove when told to. The analyser goes into trouble shooting guide Inlet gasket replacement . The users do not have clearance to do this. The inlet gasket dose not need to be replaced all you need to do is to step through the guide with out removing any parts. The analyser then returns to normal operation. | Customer runs through the replacement screen and after that the analyser goes back to normal use. | SWO-131255 | No | Error found if users keep the syringe pushed on the probe to long and dose not remove when told to. The analyser goes into trouble shooting guide Inlet gasket replacement . The users do not have clearance to do this. The inlet gasket dose not need to be replaced all you need to do is to step through the guide with out removing any parts. The analyser then returns to normal operation. | Customer runs through the replacement screen and after that the analyser goes back to normal use. | User error | HW | Auto Inlet module | Sampler not removed
ABL90 FLEX PLUS | Telefonische Kontaktaufnahme mit Haus EDV. Erfragen der Verbindungsdaten. Vorortservice. Netzwerkdaten nicht in Systemeinstellung. Analysator wurde ausgeschaltet, danach keine Verbindungsdaten im Analysator. Daten wurden bei Auslieferung nicht gespeichert! Eingabe der Netzwerkdaten, freischaltung des Ping. Speicherung der Daten auf Windowsebene. Neustart des Analysators. Netzwerkverbindung hergestellt. Mehrere, nicht freigegebene Parameter werden angezeigt. Na und K über Aqure gesperrt aber nicht freigegeben. Parameter ohne Berechtigung abgeschaltet. Gesperrte Parameter freigegeben und deaktiviert. Analysator im Hinterlegten Parameterstatus. | Telefonische Kontaktaufnahme mit Haus EDV. Erfragen der Verbindungsdaten. Vorortservice. Netzwerkdaten nicht in Systemeinstellung. Analysator wurde ausgeschaltet, danach keine Verbindungsdaten im Analysator. Daten wurden bei Auslieferung nicht gespeichert! Eingabe der Netzwerkdaten, freischaltung des Ping. Speicherung der Daten auf Windowsebene. Neustart des Analysators. Netzwerkverbindung hergestellt. Mehrere, nicht freigegebene Parameter werden angezeigt. Na und K über Aqure gesperrt aber nicht freigegeben. Parameter ohne Berechtigung abgeschaltet. Gesperrte Parameter freigegeben und deaktiviert. Analysator im Hinterlegten Parameterstatus. | SWO-131263 | No | Telephone contact with Haus EDV. Inquiring about the connection data. On-site service. Network data not in system setting. Analyzer was turned off, then no connection data in the analyzer. Data was not saved on delivery! Entering the network data, activating the ping. Store data at the Windows level. Restart the analyzer. Network connection established. Multiple parameters that are not shared are displayed. Na and K blocked via Aqure but not released. Parameters without authorization. Locked parameters released and deactivated. Analyzer in the stored parameter status. | Telephone contact with Haus EDV. Inquiring about the connection data. On-site service. Network data not in system setting. Analyzer was turned off, then no connection data in the analyzer. Data was not saved on delivery! Entering the network data, activating the ping. Store data at the Windows level. Restart the analyzer. Network connection established. Multiple parameters that are not shared are displayed. Na and K blocked via Aqure but not released. Parameters without authorization. Locked parameters released and deactivated. Analyzer in the stored parameter status. | Setup | His/Lis | His/Lis | His/Lis
ABL80 CO-OX | Sustitución de la totalidad de tubos del analizador, revisión preventiva según procedimiento, limpieza y descontaminación de la salida de desperdicios chequeo y limpieza de los circuitos de fluidos así como de los circuitos electrónicos, realización del ciclo de calidad, funcionando correctamente. | Sustitución de la totalidad de tubos. | D15AE8C362503B2F | No | Replacing all the tubes of the analyzer, preventive review according to procedure, cleaning and decontamination of the waste outlet check and cleaning of fluid circuits as well as electronic circuits, realization of the cycle of quality, working properly. | Replacement of all tubes.
ABL835 FLEX | Se realiza mantenimiento preventivo según el protocolo del fabricante. Equipo operativo. Se cambia sonda de aspiracion u se hace desproteinizacion | B | D161686EB00807F1 | No | Preventive maintenance is performed according to the manufacturer's protocol. Operational team. Aspiration probe is changed or deproteinization is made |  | Defect | HW | Inlet module | Probe
ABL837 FLEX | CreaA Bメンブラン交換。溶液類交換。MetⅡモジュールの測定チャンバー及びダミー電極洗浄。エアーライン洗浄。 | CreaA Bメンブラン交換。溶液類交換。MetⅡモジュールの測定チャンバー及びダミー電極洗浄。エアーライン洗浄。 | SWO-131689 | No | CreaA B Membrane replacement. Solution exchange. Measurement chamber and dummy electrode cleaning of met module. Air line cleaning. | CreaA B Membrane replacement. Solution exchange. Measurement chamber and dummy electrode cleaning of met module. Air line cleaning. | Assisted Replacement | Electrode | Membrane | Crea | 953/QC High
ABL837 FLEX |  | B | D15B3ADA623F97CA | No |  |  | Assisted Replacement | Electrode | Electrode | pO2 | Cracked
ABL800 BASIC | Measuring Light Heat PCB交換 | Measuring Light Heat PCB交換 | SWO-131679 | No | Measuring Light Heat PCB交換 | Measuring Light Heat PCB交換 | Defect | HW | pH/BG Module | Heater PCB | 355
ABL825 FLEX | Inspected waste tubing and found waste tubing off at connector. | Reconnected tubing. | D15AC030624BDDD2 | No | Inspected waste tubing and found waste tubing off at connector. | Reconnected tubing. | Adjustment | HW | Inlet module | Tubing waste
ABL825 FLEX | Replaced reference electrode. | B | SWO-131428 | No | Replaced reference electrode. |  | Assisted Replacement | Electrode | Electrode | Reference
ABL827 FLEX |  | B | D15B1AB46209D96D | No |  |  | Defect | HW | Inlet module | Probe bend
ABL800 BASIC | Measuring light Heat PCB交換 | Measuring light Heat PCB交換 | SWO-131135 | No | Measuring light Heat PCB交換 | Measuring light Heat PCB交換 | Defect | HW | pH/BG Module | Measuring Heat PCB
ABL837 FLEX |  | B | D15ECAB3C455360B | Yes |  |  | Re classified RMED | Setup/installation
ABL827 FLEX | Eine USV zwischen Netzsteckdose und ABL geschaltet um mögliche Netzschwankungen zu puffern - ohne Erfolg - der sporadische Neustart bleibt weiterhin bestehen, Sicherungskontakte und Sicherungen auf dem AU Board mittels Glasfaserpinsel gereinigt, Power Supply getauscht, Gerät neu gestartet, nach Neustart alles Ok, mehrere Kal. durchgeführt, System- und Funktionskontrolle in Ordnung. | B | SWO-131552 | No | A UPS switched between power socket and ABL to buffer possible power fluctuations - without success - the sporadic restart remains, backup contacts and fuses on the AU board cleaned by means of fiber optic brush, power supply exchanged, device restarted, after rebooting everything Ok, multiple Kal. system and function control in order. |  | Defect | HW | DMS_PC | Power supply | Presumably AU PCB
AQT90 FLEX | cleaned the process wheel and the optical unit , then also the problem not solved. while in the measuring program the dark counts are more than 80. and standard deviation value in system check is not in range. need the replacement of optical unit. | cleaned the process wheel and the optical unit , then also the problem not solved. while in the measuring program the dark counts are more than 80. and standard deviation value in system check is not in range. need the replacement of optical unit. | SWO-131318 | No | cleaned the process wheel and the optical unit , then also the problem not solved. while in the measuring program the dark counts are more than 80. and standard deviation value in system check is not in range. need the replacement of optical unit. | cleaned the process wheel and the optical unit , then also the problem not solved. while in the measuring program the dark counts are more than 80. and standard deviation value in system check is not in range. need the replacement of optical unit. | Defect | HW | Process Module | Optical Unit | 129
ABL90 FLEX | Verbogene Einlasssonde erneuert, Sensor Kassette erneuert. Gerät mehrfach kalibriert | Einlassdichtung durch Kundeneigene ersetzt. Einlasssonde erneuert. Sensor Kassette und Lösungspack durch Kundeneigene Gene erneuert. | SWO-131628 | No | Bent inlet probe renewed, sensor cassette renewed. Device calibrated multiple times | Inlet seal replaced by customer's own. Inlet probe renewed. Sensor cassette and solution pack renewed by customer-owned genes. | Replaced | HW | Inlet module | Probe bend
ABL90 FLEX | Einlassdichtung durch Kundeneigene ersetzt. Einlasssonde erneuert. Sensor Kassette und Lösungspack durch Kundeneigene Gene erneuert. | Einlassdichtung durch Kundeneigene ersetzt. Einlasssonde erneuert. Sensor Kassette und Lösungspack durch Kundeneigene Gene erneuert. | SWO-131631 | Yes | Inlet seal replaced by customer's own. Inlet probe renewed. Sensor cassette and solution pack renewed by customer-owned genes. | Inlet seal replaced by customer's own. Inlet probe renewed. Sensor cassette and solution pack renewed by customer-owned genes. | Re classified RMED | Double UpSWO
ABL90 FLEX | Fehlersuche, Austausch Oximetrieschlauch und Schlauchsystem, tHb-Kal, Funktionskontrolle | Fehlersuche, Austausch Oximetrieschlauch und Schlauchsystem, tHb-Kal, Funktionskontrolle | SWO-131387 | No | Troubleshooting, replacement of oximetry hose and hose system, tHb-Kal, function control | Troubleshooting, replacement of oximetry hose and hose system, tHb-Kal, function control | Defect | HW | PM Kit | PM Kit
ABL815 FLEX | Schlauchwege zwischen Inlet und Abfallpumpe erneuert. Verbindungsnipple zwischen Inlet und Abfallpumpe mit einer Kanüle auf freigängigkeit überprüft und mit Wasser durchgespült. Verbogene Einlasssonde gerichtet. Einlassdichtung erneuert. Pumpen Kal und LS Justage durchgeführt. BGA durchgeführt. | Schlauchwege zwischen Inlet und Abfallpumpe erneuert. Verbindungsnipple zwischen Inlet und Abfallpumpe mit einer Kanüle auf freigängigkeit überprüft und mit Wasser durchgespült. Verbogene Einlasssonde gerichtet. Einlassdichtung erneuert. Pumpen Kal und LS Justage durchgeführt. BGA durchgeführt. | SWO-131630 | No | Hose paths between inlet and waste pump. Connection nipple between the inlet and the waste pump with a cannula checked for freeability and rinsed with water. Bent inlet probe. Inlet seal renewed. Pumps Kal and LS adjustment performed. BGA. | Hose paths between inlet and waste pump. Connection nipple between the inlet and the waste pump with a cannula checked for freeability and rinsed with water. Bent inlet probe. Inlet seal renewed. Pumps Kal and LS adjustment performed. BGA. | Defect | HW | Inlet module | Probe bend
TCM400 | Calibration sensitivity error. need to replace the membrane. | Calibration sensitivity error. need to replace the membrane. | SWO-132909 | No | Calibration sensitivity error. need to replace the membrane. | Calibration sensitivity error. need to replace the membrane.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-131956 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX |  | B | SWO-131397 | No |  |  | Defect | HW | DMS_PC | Display | LEH Mail
TCM CombiM |  | B | SWO-131384 | No
ABL90 FLEX | 更换了636-656后修复。910-298未更换 by jesson | 更换了636-656后修复。910-298未更换 by jesson | SWO-131319 | No | Repair after replacing 636-656. 910-298 Unreplaced by jesson | Repair after replacing 636-656. 910-298 Unreplaced by jesson | Defect | HW | DMS_PC | Display
AQT90 FLEX |  | B | SWO-131369 | No |  |  | Defect | HW | Cover and Chassis | PCB, Cabinet | 4548
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-131472 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX Basic version |  | B | D15ED3DAC42CB564 | No
ABL80 OSM | The SW is showing corrupted files. | Replaced CF card. | D15E7E1EC4435F29 | No | The SW is showing corrupted files. | Replaced CF card.
ABL90 FLEX | Fliesswege gespült, Einlass gereinigt, Serviceprogramme geprüft, Abfallschlauch und OxiVentilSchlauch ersetzt, Startphase, Lösungspack nochmal erneuert, mehrere Kal abgewartet, Kal & QK ok | Mit Schw. Stefanie versucht, den ABL ohne Spülvorrichtung zu spülen - ohne Erfolg, Kunde bittet um Service | SWO-131798 | No | Flow paths rinsed, inlet cleaned, service programs tested, waste hose and OxiVentilhose replaced, start phase, solution pack renewed, several kal waited, Kal & QK ok | With Schw. Stefanie tries to flush the ABL without flushing device - without success, customer asks for service | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | ABL uit en weer aangezet, hierna werkt alles weer naar behoren. Wellicht is er sprake geweest van een elektronische storing. | B | SWO-131569 | No | ABL is turned off and on again, after that everything works properly again. There may have been an electronic malfunction. |  | Defect | SW | Off/On | Oxi | 134
ABL90 | Found blockage in waste connector. Sensor cassette Lactate issues. | Install PM kit and replace waste connector init. Install software version 3.1 MR8. Replace sensor cassette from customer stock. | D15EE743C43AFF08 | No | Found blockage in waste connector. Sensor cassette Lactate issues. | Install PM kit and replace waste connector init. Install software version 3.1 MR8. Replace sensor cassette from customer stock. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Unit powers off immediately when unplugged. | Replaced battery unit with a battery from another unit that is not moved around because batteries are on a back order. Will swap out when new battery arrives. Using battery from 172N001 unit | D15EBB8BC453D834 | No | Unit powers off immediately when unplugged. | Replaced battery unit with a battery from another unit that is not moved around because batteries are on a back order. Will swap out when new battery arrives. Using battery from 172N001 unit | Defect | HW | Battery module | Battery module | EJNSC?
ABL90 FLEX | Gerät überprüft, Lösungspack erneuert, Einlaß gesäubert, Meßkassette ausgetauscht | Gerät überprüft, Lösungspack erneuert, Einlaß gesäubert, Meßkassette ausgetauscht | SWO-131813 | No | Device checked, solution pack renewed, inlet cleaned, measuring cassette replaced | Device checked, solution pack renewed, inlet cleaned, measuring cassette replaced | Assisted Replacement | SC | SC | Glu
ABL90 |  | B | D15EEA57C41E81AD | Yes |  |  | Re classified RMED | Planned activity |  |  | LEH Mail
ABL90 FLEX | Auf Control Panel Ebene den Befehl repair connection ausgeführt. Kontrolle der Datenverbindung --> iO. | Auf Control Panel Ebene den Befehl repair connection ausgeführt. Kontrolle der Datenverbindung --> iO. | SWO-131619 | No | Executed the repair connection command at the Control Panel level. Control of the data connection --> iO. | Executed the repair connection command at the Control Panel level. Control of the data connection --> iO. | Setup | SW | His/Lis | His/Lis
ABL90 FLEX PLUS | Fehlersuche, div. Programme im Service durchgef., Neustart, Kal. + QC i.O. | Fehlersuche; div. Programme im Service durchgef.; Neustart; Kal. + QC i.O. | SWO-131503 | No | Troubleshooting; various programs in the service; Restart; Kal. + QC in the first place | Troubleshooting; various programs in the service; Restart; Kal. + QC in the first place | Adjustment | HW | Pump calibration | Pump calibration
ABL90 FLEX PLUS | Service dump taken and video of screen and how this maybe happening. Tech line case will be raised. | B | SWO-131655 | No | Service dump taken and video of screen and how this maybe happening. Tech line case will be raised. |  | Unknown | HW | Auto Inlet module | Inlet Position Cal.
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-133507 | No | Repair after replacing a new test card | Repair after replacing a new test card
AQT90 FLEX | 更换了扫描器910-304后修复 by jesson | 更换了扫描器910-304后修复 by jesson | SWO-131323 | No | Repair edgy jesson after replacing scanner 910-304 | Repair edgy jesson after replacing scanner 910-304 | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX |  | B | SWO-131410 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | Sustituida aguja. Pasados controles de caliidad y comrpobado buen funcionamiento. El equipo reporta error 1029 y 1031. | B | D15EA46FC45EDB3E | No | Replaced needle. Past quality controls and well functioning. The computer reports error 1029 and 1031. |  | Defect | HW | Process Module | Optical Unit | 131
ABL837 FLEX | 測定チャンバー洗浄、除ﾀﾝﾊﾟｸ | B | SWO-131673 | No | Measuring chamber cleaning, deproteinising |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL815 FLEX | Glu Elektrode mit Cleaner behandelt, Messkammer überprüft, Fließwege überprüft, Kontakt überprüft, Membran ersetzt, Kal und QC | B | SWO-131789 | No | Glu electrode treated with cleaner, measuring chamber checked, flow paths checked, contact checked, membrane replaced, Kal and QC |  | Cleaned | Fluidic pathway | El/Met Module | Glu
ABL815 FLEX | pH Elektrode überprüft, gereinigt und neue Dichtung angebracht, Glu und Lac Elektrode gereinigt und erneut membraniert, Nullstromkalibrierung vorgenommen, mehrere Kal, QC | Intensive Reinigung der pH-Elektrode, pO2-Elektrode membraniert, korrekte Kal abgewartet, QK ok, Messungen an Blut ok | SWO-131790 | No | pH electrode checked, cleaned and new seal attached, Glu and Lac electrode cleaned and re-logged, zero current calibration performed, several Kal, QC | Intensive cleaning of the pH electrode, pO2 electrode membraneed, correct kalwaited, QK ok, measurements on blood ok | Cleaned | Fluidic pathway | pH/BG Module | pH
ABL800 BASIC | Need to replace the Measuring Chamber. Part number-902-446 | B | SWO-131644 | No | Need to replace the Measuring Chamber. Part number-902-446 |  | Defect | HW | pH/BG Module | Measuring chamber
ABL825 FLEX | Fehler ist erneut mit dem bereits ersetzten neuen Bildschirm aufgetreten. Monitor 910-320 (gebraucht) aus ehemaligem KUK Altgerätebestand eingesetzt, ausgebauten Monitor desinfiziert und zur Weitersendung an BorWe eingepackt. | Fehler ist erneut mit dem bereits ersetzten neuen Bildschirm aufgetreten. Monitor 910-320 (gebraucht) aus ehemaligem KUK Altgerätebestand eingesetzt, ausgebauten Monitor desinfiziert und zur Weitersendung an BorWe eingepackt. | SWO-131616 | No | Failed again with the new screen that has already been replaced. Monitor 910-320 (used) from former KUK waste equipment stock used, removed monitor disinfected and packed for forwarding to BorWe. | Failed again with the new screen that has already been replaced. Monitor 910-320 (used) from former KUK waste equipment stock used, removed monitor disinfected and packed for forwarding to BorWe. | Defect | HW | DMS_PC | Touch screen
ABL815 FLEX | Monitor ersetzt. | Monitor ersetzt. | SWO-131530 | No | Monitor replaced. | Monitor replaced. | Defect | HW | DMS_PC | Touch screen
ABL827 FLEX | Verdacht auf Netzschwankungen, von Med. Tech. eine USV dazwischen geschalten - brachte keine Besserung, Netzteil getauscht und Sicherungen auf AU Board kontrolliert und Kontaktflächen mittels Glasfaserpinsel gesäubert - war aber keine Verschmutzung Korrosion zu erkennen - alle Maßnahmen brachten keine Besserung. Service Dump erstellt und an TS versendet. | B | SWO-131889 | No | Suspicion of network fluctuations, from Med. Tech. Switching a UPS in between - brought no improvement, replaced power supply and fuses checked on AU board and cleaned contact surfaces by means of fiberglass brushes - but no pollution corrosion was detectable - all measures brought no improvement. Service Dump created and shipped to TS. |  | Defect | HW | DMS_PC | Power supply
AQT90 FLEX | changed the optical unit from our stock, and done the cleaning of AQT90, now the analyser working satisfactorily. | changed the optical unit from our stock, and done the cleaning of AQT90, now the analyser working satisfactorily. | SWO-132805 | No | changed the optical unit from our stock, and done the cleaning of AQT90, now the analyser working satisfactorily. | changed the optical unit from our stock, and done the cleaning of AQT90, now the analyser working satisfactorily. | Defect | HW | Process Module | Optical Unit | 129
ABL80 FLEX BASIC | Teile erneuert, Gerät nicht getestet, da aktuell nicht in Verwendung, ACHTUNG: 580-093 nicht buchbar, darum SWO nicht geschlossen! | B | SWO-131837 | No | Parts renewed, device not tested, since currently not in use, ATTENTION: 580-093 not bookable, therefore SWO not closed!
ABL90 | Verified once on site also found error 0377 (Calibration drift out of range). | Initially completed a full service flush of the system to rule out any contamination in the system. All calibrations and QC's are currently passing. Will follow up with the customer. | D16023203A13B03B | No | Verified once on site also found error 0377 (Calibration drift out of range). | Initially completed a full service flush of the system to rule out any contamination in the system. All calibrations and QC's are currently passing. Will follow up with the customer. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Startphase QC und Kalibration. | Startphase QC und Kalibration. | SWO-131720 | No | Start phase QC and calibration. | Start phase QC and calibration. | Adjustment | SW | Service | Service
ABL90 FLEX PLUS |  | B | D1604AA73A5C06F0 | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
TCM4 | 更换945-377一个正常。 | B | SWO-131951 | No | Replace the 945-377 one normal.
ABL90 FLEX |  | B | SWO-131684 | No |  |  | Defect | HW | DMS_PC | CPU
ABL800 FLEX | Checked and found PCo2 Chamber was clotted from blood particle, clean it Clean also the PH electrode chamber. Replaced the PCo2 membrane. Run the protein removal solution. Executed the 2 Point Calibration. | B | SWO-134753 | No | Checked and found PCo2 Chamber was clotted from blood particle, clean it Clean also the PH electrode chamber. Replaced the PCo2 membrane. Run the protein removal solution. Executed the 2 Point Calibration. |  | Cleaned | HW | pH/BG Module | Measuring chamber | LEH Mail
ABL825 FLEX | Upon arrival, found the analyzer presenting rinse errors. Found the analyzer leaking from the pH BG and, EI Met modules. | Checked the activity log, found error 502 Air detected when filling the inlet during pump calibration. Replaced all listed parts. Cleaned all pump rotors. replaced Na, Ca and, K membranes. Replaced the fluidic module rinse valve. replaced all tubing using the radiometer yearly service kit. | D15FF1D23A0F1D41 | No | Upon arrival, found the analyzer presenting rinse errors. Found the analyzer leaking from the pH BG and, EI Met modules. | Checked the activity log, found error 502 Air detected when filling the inlet during pump calibration. Replaced all listed parts. Cleaned all pump rotors. replaced Na, Ca and, K membranes. Replaced the fluidic module rinse valve. replaced all tubing using the radiometer yearly service kit. | Defect | HW | Basic Unit | PM Kit
ABL835 FLEX |  | B | D160125C3A278B77 | No |  |  | Setup | His/Lis | His/Lis | His/Lis
ABL825 FLEX | The customer reported that the FlexQ module is not returning to home position after sample is run on it. | Found the tray riding shaft loose and, missing the securing screw. Replaced the securing screw and, repositioned the shaft. Performed all service adjustments for the FlexQ module to confirm proper operation. | D16019CB3A352BEE | No | The customer reported that the FlexQ module is not returning to home position after sample is run on it. | Found the tray riding shaft loose and, missing the securing screw. Replaced the securing screw and, repositioned the shaft. Performed all service adjustments for the FlexQ module to confirm proper operation. | Defect | HW | FlexQ Module | FlexQ Tray | Screw missing
ABL820 FLEX | The brace on both sides of the printer were broken. | Replaced printer unit. | D15FC5103A155F37 | No | The brace on both sides of the printer were broken. | Replaced printer unit. | Defect | HW | Printer module | Printer module
AQT90 FLEX | Storage wheel checked in service program for safe possitions, fail at safe. Analyser opened and clear all cartridges from storage wheel. Analyser switched on and try to install cartridge , unable to read barcoad. analyser open and clear for cartridge barcoad scaner. Reinstall cartridge and installed. Analyserr serviced properly and handed over in a fine working conditrion. | Storage wheel checked in service program for safe possitions, fail at safe. Analyser opened and clear all cartridges from storage wheel. Analyser switched on and try to install cartridge , unable to read barcoad. analyser open and clear for cartridge barcoad scaner. Reinstall cartridge and installed. Analyserr serviced properly and handed over in a fine working conditrion. | SWO-132838 | No | Storage wheel checked in service program for safe possitions, fail at safe. Analyser opened and clear all cartridges from storage wheel. Analyser switched on and try to install cartridge , unable to read barcoad. analyser open and clear for cartridge barcoad scaner. Reinstall cartridge and installed. Analyserr serviced properly and handed over in a fine working conditrion. | Storage wheel checked in service program for safe possitions, fail at safe. Analyser opened and clear all cartridges from storage wheel. Analyser switched on and try to install cartridge , unable to read barcoad. analyser open and clear for cartridge barcoad scaner. Reinstall cartridge and installed. Analyserr serviced properly and handed over in a fine working conditrion. | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX BASIC | calibration multiple time. calibration pass. follow-up visit. | calibration multiple time. calibration pass. follow-up visit. | SWO-132905 | No | calibration multiple time. calibration pass. follow-up visit. | calibration multiple time. calibration pass. follow-up visit.
ABL80 FLEX BASIC | Cleaned Sc with distilled water and kept in fridge. Again reinstalled Sc but problem still persist. Need to replace SC. | Cleaned Sc with distilled water and kept in fridge. Again reinstalled Sc but problem still persist. Need to replace SC. | SWO-132852 | No | Cleaned Sc with distilled water and kept in fridge. Again reinstalled Sc but problem still persist. Need to replace SC. | Cleaned Sc with distilled water and kept in fridge. Again reinstalled Sc but problem still persist. Need to replace SC.
ABL80 FLEX BASIC | Removed air from solution pack pouch and reinstalled solution pack again, Problem is repeated, Need to change solution pack. | Removed air from solution pack pouch and reinstalled solution pack again, Problem is repeated, Need to change solution pack. | SWO-132853 | No | Removed air from solution pack pouch and reinstalled solution pack again, Problem is repeated, Need to change solution pack. | Removed air from solution pack pouch and reinstalled solution pack again, Problem is repeated, Need to change solution pack.
ABL80 FLEX BASIC | Replaced the SP and again perform QC but problem is still remaining. So SC need replacement. | Replaced the SP and again perform QC but problem is still remaining. So SC need replacement. | SWO-132870 | No | Replaced the SP and again perform QC but problem is still remaining. So SC need replacement. | Replaced the SP and again perform QC but problem is still remaining. So SC need replacement.
ABL90 | Multiple cassettes and solution packs tried. | On arrival analyzer was up and functioning correctly. Perform hot water back flush multiple times. Perform all service tests. Replace inlet gasket, connector gasket, and inlet probe. perform all service tests a second time. Perform back flush again. | D1614B05B02A8C33 | No | Multiple cassettes and solution packs tried. | On arrival analyzer was up and functioning correctly. Perform hot water back flush multiple times. Perform all service tests. Replace inlet gasket, connector gasket, and inlet probe. perform all service tests a second time. Perform back flush again. | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | Überprüfung Analysator - keine Defekte an Analysator und Aktivitäten-Bericht erkennbar - Kundenschreiben weitergeleitet und mit Labor besprochen ! SP LOT : XP-02 | Überprüfung Analysator - keine Defekte an Analysator und Aktivitäten-Bericht erkennbar - Kundenschreiben weitergeleitet und mit Labor besprochen ! SP LOT : XP-02 | SWO-132464 | No | Review analyzer - no defects recognizable to analyzer and activity report - customer letter forwarded and discussed with laboratory ! SP LOT : XP-02 | Review analyzer - no defects recognizable to analyzer and activity report - customer letter forwarded and discussed with laboratory ! SP LOT : XP-02 | Revisit required | SP | SP | pCO2
ABL90 FLEX | Replace SP | Replace SP | SWO-132367 | No | Replace SP | Replace SP | Assisted Replacement | SP | SP | pO2
ABL90 FLEX | Flushed analyser several times, resistance from around oximetry module Ran clot detection program several times to remove blockage | . | SWO-132366 | No | Flushed analyser several times, resistance from around oximetry module Ran clot detection program several times to remove blockage | . | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Tausch des Einlassdichtung und Spülen der Fließwege. Weiterhin Pumpenkalibrierfehler. Tausch des Lösungspacks, alle Funktionen im Service o. K. Defekten Lösungspack reklamiert. Kalibrationen und Qualitätskontrolle in Ordnung. | Tausch des Einlassdichtung und Spülen der Fließwege. Weiterhin Pumpenkalibrierfehler. Tausch des Lösungspacks, alle Funktionen im Service o. K. Defekten Lösungspack reklamiert. Kalibrationen und Qualitätskontrolle in Ordnung. | SWO-132157 | No | Replacement of the inlet seal and rinsing of the flow paths. Further pump calibration errors. Replacement of the solution pack, all functions in the service o.K. defective solution pack complained. Calibrations and quality control in order. | Replacement of the inlet seal and rinsing of the flow paths. Further pump calibration errors. Replacement of the solution pack, all functions in the service o.K. defective solution pack complained. Calibrations and quality control in order. | Assisted Replacement | SP | SP | Pump calibration
ABL90 FLEX PLUS | Failing pressure test out of limits and failing. Replaced inlet gasket and test ok. | Failing pressure test out of limits and failing. Replaced inlet gasket and test ok. | SWO-132125 | No | Failing pressure test out of limits and failing. Replaced inlet gasket and test ok. | Failing pressure test out of limits and failing. Replaced inlet gasket and test ok. | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Fin d'après-midi j'ai de nouveau changé l'appareil par l'original, échangé de pack de solution et cassette de capteur, toutes les CQ sont bien dans les gammes de controles. | Fin d'après-midi j'ai de nouveau changé l'appareil par l'original, échangé de pack de solution et cassette de capteur, toutes les CQ sont bien dans les gammes de controles. | SWO-132183 | No | Late afternoon I again changed the device by the original, exchanged solution pack and sensor cassette, all THE CQ are well within the ranges of controls. | Late afternoon I again changed the device by the original, exchanged solution pack and sensor cassette, all THE CQ are well within the ranges of controls. | Assisted Replacement | SP | SP | pO2
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-132191 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Multi
AQT90 FLEX |  | B | SWO-131972 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL800 FLEX | Crea membrane & consumable replacement . | Crea membrane & consumable replacement . | SWO-132521 | No | Crea membrane & consumable replacement . | Crea membrane & consumable replacement . | Assisted Replacement | Electrode | Membrane | Crea | New installation
ABL820 FLEX | The customer is completely unsatisfied with ABL 820 with low Pco2 controls on alls levels this issue is present for at less a year we already replaced analyser with a different with the same issue | Contact sale rep to to found out a solution | D161541CB0054CDF | No | The customer is completely unsatisfied with ABL 820 with low Pco2 controls on alls levels this issue is present for at less a year we already replaced analyser with a different with the same issue | Contact sale rep to to found out a solution | Assisted Replacement | Electrode | Membrane | pCO2 | ABL82
ABL825 FLEX | Spectrometer ausgetauscht Lichtleiter ausgetauscht Komplette Lampeneinheit ausgetauscht und justiert tHb Kal Funktionskontrolle i.o | Spectrometer ausgetauscht Lichtleiter ausgetauscht Komplette Lampeneinheit ausgetauscht und justiert tHb Kal Funktionskontrolle i.o | SWO-132013 | No | Spectrometer replaced fiber optic exchange Complete lamp unit replaced and adjusted tHb Kal function control i.o | Spectrometer replaced fiber optic exchange Complete lamp unit replaced and adjusted tHb Kal function control i.o | Defect | HW | Oxi module | Lamp unit
ABL825 FLEX | pH elektrode renset med linse papir Kørt 3 x blodprøve - OK | pH elektrode renset med linse papir Kørt 3xblodprøve - ok | SWO-132417 | No | pH electrode cleaned with lens paper driven 3 x blood test-OK | pH electrode cleaned with lens paper driven 3xblood test-ok | Cleaned | Fluidic pathway | pH/BG Module | pH | 476
ABL835 FLEX | Cl reading high or is outside package insert | Had replaced CL electrode from another unit when unit was moved and unit is still trending high. | D1618124B03DF664 | No | Cl reading high or is outside package insert | Had replaced CL electrode from another unit when unit was moved and unit is still trending high. | Assisted Replacement | Electrode | Electrode | Cl
ABL820 FLEX | reference membrane was over due for a change. analyzer locked. after reference membrane was changed analyzer was still locked. checked all corrective actions to see what was locking the analyzer. QC was locking the analyzer and removed the 10% over due lock. analyzer is still locked after running QC. found QC schedule was wiped out | replaced the QC schedule unlocked the analyzer | D1616B68B014C13B | No | reference membrane was over due for a change. analyzer locked. after reference membrane was changed analyzer was still locked. checked all corrective actions to see what was locking the analyzer. QC was locking the analyzer and removed the 10% over due lock. analyzer is still locked after running QC. found QC schedule was wiped out | replaced the QC schedule unlocked the analyzer | User error | Electrode | Membrane | Reference | Exp.
ABL800 BASIC |  | B | SWO-132569 | No |  |  | Unknown | No fault found | No fault found | No fault found
ABL805 FLEX | Frau Klubert hat im Rahmen einer Konfiguration Einträge im Pat. ID Layout entfernt - diese lassen sich nicht mehr neu anwählen, das linke Fenster ist komplett leer. Alte Konfig. Layouts neu eingespielt und Pat. Berichts Layout den Gegebenheiten entsprechend angepasst. | Frau Klubert hat im Rahmen einer Konfiguration Einträge im Pat. ID Layout entfernt - diese lassen sich nicht mehr neu anwählen, das linke Fenster ist komplett leer. Alte Konfig. Layouts neu eingespielt und Pat. Berichts Layout den Gegebenheiten entsprechend angepasst. | SWO-132181 | No | Ms. Klubert has entries in the Pat as part of a configuration. ID layout removed - these can no longer be reselected, the left window is completely empty. Old Config. layouts re-recorded and Pat. Report layout adapted to the circumstances. | Ms. Klubert has entries in the Pat as part of a configuration. ID layout removed - these can no longer be reselected, the left window is completely empty. Old Config. layouts re-recorded and Pat. Report layout adapted to the circumstances. | User error | Setup | Setup | Patient layout
ABL825 FLEX | Layouts Konfiguration eines vergleichbaren Analysators übertragen und entsprechend der gewünschten Parameter-Reihenfolge angepasst. Testausdruck, OK. | Layouts Konfiguration eines vergleichbaren Analysators übertragen und entsprechend der gewünschten Parameter-Reihenfolge angepasst. Testausdruck, OK. | SWO-132182 | No | Layouts configure a comparable analyzer and adjust it according to the desired parameter order. Test expression, OK. | Layouts configure a comparable analyzer and adjust it according to the desired parameter order. Test expression, OK. | User error | Setup | Setup | Patient layout
ABL800 BASIC | reinstalled application software6.16. adjusted halogen lamp. performed protien removal. performed two point calibration. performed thb calibration. | reinstalled application software6.16. adjusted halogen lamp. performed protien removal. performed two point calibration. performed thb calibration. | SWO-132946 | No | reinstalled application software6.16. adjusted halogen lamp. performed protien removal. performed two point calibration. performed thb calibration. | reinstalled application software6.16. adjusted halogen lamp. performed protien removal. performed two point calibration. performed thb calibration. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Key s überprüft Probleme mit dem E-Lyte Key, Demo Key für E-Lyte wieder eingeschaltet, Scanner leuchtet nicht - Gerät länger ausgeschaltet - Scanner nach Neustart wieder i.O, Tastentöne auf Wunsch von Hr. Breitmoser abgeschaltet, Kal und QC durchgeführt, Anwender werden Scanner Funktion beobachten und sich bei Problemen wieder melden | Key s überprüft Probleme mit dem E-Lyte Key, Demo Key für E-Lyte wieder eingeschaltet, Scanner leuchtet nicht - Gerät länger ausgeschaltet - Scanner nach Neustart wieder i.O, Tastentöne auf Wunsch von Hr. Breitmoser abgeschaltet, Kal und QC durchgeführt, Anwender werden Scanner Funktion beobachten und sich bei Problemen wieder melden | SWO-132161 | Yes | Key s checks problems with the E-Lyte Key, Demo Key for E-Lyte turned back on, Scanner does not light up - device is switched off longer - Scanner resumes after restart, key tones switched off at Mr. Breitmoser's request, Kal and QC performed, users are scanners Observe function and report back in case of problems | Key s checks problems with the E-Lyte Key, Demo Key for E-Lyte turned back on, Scanner does not light up - device is switched off longer - Scanner resumes after restart, key tones switched off at Mr. Breitmoser's request, Kal and QC performed, users are scanners Observe function and report back in case of problems | Re classified RMED | Setup new install
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-132540 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-134118 | No | Replacement test card repair | Replacement test card repair
ABL90 | checked fluid flow bubbles in flow from pack replaced connector gasket ande inlet gasket checked fluid flow still bubbles replaced pack | replaced pack inlet gasket connector gasket | D1626F592633911F | No | checked fluid flow bubbles in flow from pack replaced connector gasket ande inlet gasket checked fluid flow still bubbles replaced pack | replaced pack inlet gasket connector gasket | Assisted Replacement | SP | SP | FT
ABL90 FLEX | Tausch des Lösungspacks und Einsatz eines neuen aus Kopenhagen gelieferten. Los Nr. neuer Lösungspack XP-06, alter Lösungspack Lot Nr. XK-03. Kalibrationen und Qualitätskontrolle in Ordnung. | Tausch des Lösungspacks und Einsatz eines neuen aus Kopenhagen gelieferten. Los Nr. neuer Lösungspack XP-06, alter Lösungspack Lot Nr. XK-03. Kalibrationen und Qualitätskontrolle in Ordnung. | SWO-132591 | No | Replacement of the solution pack and use of a new one delivered from Copenhagen. Lot No. new solution pack XP-06, old solution pack Lot No. XK-03. Calibrations and quality control in order. | Replacement of the solution pack and use of a new one delivered from Copenhagen. Lot No. new solution pack XP-06, old solution pack Lot No. XK-03. Calibrations and quality control in order. | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX | Upper Lid with protective spring erneuert. SC Fach mehrmals im Service geöffnet und geschlossen - ok. SC eingelegt. Kalibrationen und QC´s vorgenommen. FHbF nach Rücksprache mit Frau Höpfner deaktiviert, da QC Fehler Level B vorlag. Es war kein tHb Kalibrator mehr Vorort und auch nicht im Wagenbestand. Bitte tHb Kalibrator bestellen und nach Eintreff Hr. Geldhauser 0172 1530693 anrufen | tHb Kalibration (100,7%) durchgeführt. FHbF wieder aktiviert. Kalibration und QC´s vorgenommen. | SWO-132502 | No | Upper Lid with protective spring renewed. SC compartment opened several times in the service and closed - ok. SC inserted. calibrations and QC's. FHbF deactivated after consultation with Ms. Höpfner because QC error level B was present. It was no longer a tHb calibrator in the suburb and also not in the car stock. Please order tHb Calibrator and call Mr. Geldhauser 0172 1530693 | tHb calibration (100.7%) Performed. FHbF is reactivated. calibration and QC's. | Defect | HW | Front cover | SC Door
ABL90 FLEX | tHb Kalibration (100,7%) durchgeführt. FHbF wieder aktiviert. Kalibration und QC´s vorgenommen. | tHb Kalibration (100,7%) durchgeführt. FHbF wieder aktiviert. Kalibration und QC´s vorgenommen. | SWO-132507 | No | tHb calibration (100.7%) Performed. FHbF is reactivated. calibration and QC's. | tHb calibration (100.7%) Performed. FHbF is reactivated. calibration and QC's. | Adjustment | HW | Oxi module | tHb Cal
ABL90 | pO2 QCs failing. Customer using Portex syringes. | Perform hot water back flush. Perform annual maintenance. Perform all service tests. | D162A52E26427F0F | No | pO2 QCs failing. Customer using Portex syringes. | Perform hot water back flush. Perform annual maintenance. Perform all service tests. | Cleaned | Fluidic pathway | Fluidic pathway | Portex
ABL90 FLEX | Tausch des Lösungspacks und Reklamation des alten Lösungspacks. Kalibrationen und Qualitätskontrolle in Ordnung. | Tausch des Lösungspacks und Reklamation des alten Lösungspacks. Kalibrationen und Qualitätskontrolle in Ordnung. | SWO-132590 | No | Swap the solution pack and complain the old solution pack. Calibrations and quality control in order. | Swap the solution pack and complain the old solution pack. Calibrations and quality control in order. | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX PLUS | Tausch des neuen Lösungspacks mit Lot Nr.: XP-06. Kalibrationen und Qualitätskontrolle in Ordnung. | Tausch des neuen Lösungspacks mit Lot Nr.: XP-06. Kalibrationen und Qualitätskontrolle in Ordnung. | SWO-132589 | No | Swap the new solution pack with lot no.: XP-06. Calibrations and quality control in order. | Swap the new solution pack with lot no.: XP-06. Calibrations and quality control in order. | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX PLUS |  | B | SWO-132697 | No |  |  | Assisted Replacement | SP | SP | pO2
ABL835 FLEX | CONTA DEĞİŞTİRİLEREK SORUN GİDERİLDİ. | CONTA DEĞİŞTİRİLEREK SORUN GİDERİLDİ. | SWO-132520 | Yes | THE PROBLEM WAS ADDRESSED BY CHANGING THE SEAL. | THE PROBLEM WAS ADDRESSED BY CHANGING THE SEAL. | Defect | Electrode | Electrode | pH | Sealing ring
ABL837 FLEX |  | B | D162C622263F8B7F | No |  |  | Defect | HW | External PS | UPS
ABL800 BASIC | Change membrane but error not solve. Needto change electrode | Change membrane but error not solve. Needto change electrode | SWO-132830 | No | Change membrane but error not solve. Needto change electrode | Change membrane but error not solve. Needto change electrode | Revisit required | Electrode | Membrane | pCO2
ABL800 BASIC | Found Lamp unit problem, adjusted the Halogen lamp, upgraded the software to 6.16V, restarted analyzer calibrated found Lac calibration error, replaced the Lac membrane found instrument is functional. Note: Preventive Maintenance is pending, require following membranes. pCO2, pO2, Ref, Na, Ka & Ca. | B | SWO-132332 | No | Found Lamp unit problem, adjusted the Halogen lamp, upgraded the software to 6.16V, restarted analyzer calibrated found Lac calibration error, replaced the Lac membrane found instrument is functional. Note: Preventive Maintenance is pending, require following membranes. pCO2, pO2, Ref, Na, Ka & Ca. |  | Adjustment | HW | Oxi module | Lamp unit
ABL837 FLEX |  | B | D16263A226436A2C | Yes |  |  | Re classified RMED | Clugston agree, mail sep.27
AQT90 FLEX | Barcode Scanner ersetzt. Neue MCC QC eingescannt. Div .Konfig und Fragen besprochen. | Barcode Scanner ersetzt. Neue MCC QC eingescannt. Div .Konfig und Fragen besprochen. | SWO-132505 | No | Barcode Scanner replaced. New MCC QC scanned. Div. Config and questions discussed. | Barcode Scanner replaced. New MCC QC scanned. Div. Config and questions discussed. | Defect | HW | Analyzer Control Module | Barcode reader, display
ABL80 FLEX BASIC | 更换测试修复 | 更换测试修复 | SWO-132718 | No | Replacement test repair | Replacement test repair
ABL80 FLEX BASIC | cleaned the sensor connector with alcohol and done zeroing for sensor connector reinstalled the SC80 but problem not rectified it need to replace with new connector. | cleaned the sensor connector with alcohol and done zeroing for sensor connector reinstalled the SC80 but problem not rectified it need to replace with new connector. | SWO-133381 | No | cleaned the sensor connector with alcohol and done zeroing for sensor connector reinstalled the SC80 but problem not rectified it need to replace with new connector. | cleaned the sensor connector with alcohol and done zeroing for sensor connector reinstalled the SC80 but problem not rectified it need to replace with new connector.
ABL90 FLEX | A készülék a helyszínen nem javítható, tesztelést igényel, a javítás ideéig cserekészülék került kihelyezésre | B | SWO-132648 | No | The device is not repairable on-site, it requires testing, a replacement has been carried out until the |  | Assisted Replacement | SP | SP | Pump calibration
ABL90 | Arrived onsite and verified the errors. Took apart the analyzer to determine if PM caused the problems. Found nothing wrong. Put analyzer back together and powered on. | Replaced the solution pack re-installed the sensor cassette. Sensor cassette reconditioned without errors. Calibration and built in QC came in without errors. | D163E28D9C094EAC | No | Arrived onsite and verified the errors. Took apart the analyzer to determine if PM caused the problems. Found nothing wrong. Put analyzer back together and powered on. | Replaced the solution pack re-installed the sensor cassette. Sensor cassette reconditioned without errors. Calibration and built in QC came in without errors. | Assisted Replacement | SP | SP | Chip
ABL90 FLEX | Unable to repair as new sensor cassette required. | Unable to repair as new sensor cassette required. | SWO-132903 | No | Unable to repair as new sensor cassette required. | Unable to repair as new sensor cassette required. | Revisit required | SC | SC | pO2
ABL90 FLEX | Replaced sensor bay with non stock spare | Replaced sensor bay with non stock spare | SWO-132814 | No | Replaced sensor bay with non stock spare | Replaced sensor bay with non stock spare | Defect | HW | Basic Unit | Sensor bay
ABL90 | CONFIRMED ISSUE WHEN ONSITE. PER MARK GUZMAN OF TECH SUPPORT, REPLACE POWER SUPPLY. | PROCEED TO REPLACE POWER SUPPLY, SN P0020N745, WITHOUT NEW BACKUP BATTERY. STATUS QUO, FUSE OF NEW POWER SUPPLY BLEW AND ANALYZER REMAINED ON BACKUP BATTERY. PER JAIME APONTE (SVC MGR), RETURN NEW POWER SUPPLY SN P0020N864 FOR CREDIT. | D163B3799C261677 | No | CONFIRMED ISSUE WHEN ONSITE. PER MARK GUZMAN OF TECH SUPPORT, REPLACE POWER SUPPLY. | PROCEED TO REPLACE POWER SUPPLY, SN P0020N745, WITHOUT NEW BACKUP BATTERY. STATUS QUO, FUSE OF NEW POWER SUPPLY BLEW AND ANALYZER REMAINED ON BACKUP BATTERY. PER JAIME APONTE (SVC MGR), RETURN NEW POWER SUPPLY SN P0020N864 FOR CREDIT. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX PLUS |  | B | D163B0919C3030E0 | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Motor kobling efterspændt | Motor kobling efterspændt | SWO-132990 | No | Motor coupling after tension | Motor coupling after tension | Adjustment | HW | Auto Inlet module | Motor assembly | 1391
ABL90 FLEX PLUS | Moved analyzer in to isolation room with extractor fan and still giving same error. Replaced Solution pack and error corrected pO2 value came up from 288mmHg to 297mmHg. Phoned customer on the 26 08 2016 to find out if all was ok over the weekend and she had no errors to report. | Moved analyzer in to isolation room with extractor fan and still giving same error. Replaced Solution pack and error corrected pO2 value came up from 288mmHg to 297mmHg. Phoned customer on the 26 08 2016 to find out if all was ok over the weekend and she had no errors to report. | SWO-133388 | No | Moved analyzer in to isolation room with extractor fan and still giving same error. Replaced Solution pack and error corrected pO2 value came up from 288mmHg to 297mmHg. Phoned customer on the 26 08 2016 to find out if all was ok over the weekend and she had no errors to report. | Moved analyzer in to isolation room with extractor fan and still giving same error. Replaced Solution pack and error corrected pO2 value came up from 288mmHg to 297mmHg. Phoned customer on the 26 08 2016 to find out if all was ok over the weekend and she had no errors to report. | Assisted Replacement | SP | SP | pO2
ABL90 FLEX PLUS | Replaced analyser ID PCB. | B | SWO-133297 | No | Replaced analyser ID PCB. |  | Defect | HW | Basic Unit | Analyzer PCB ID
ABL835 FLEX | Cl trending high or outside package insert after sveral weeks following a membrane change. | Checked valves on fludic and BG modules. Replaced measuring chamber, | D163DE489C600688 | No | Cl trending high or outside package insert after sveral weeks following a membrane change. | Checked valves on fludic and BG modules. Replaced measuring chamber, | Defect | HW | pH/BG Module | Cl
AQT90 FLEX | Has to replace 903-171: Needle Assebmbly. 903-329 : Needle wash unit, complete (Hct). | B | SWO-132785 | No | Has to replace 903-171: Needle Assebmbly. 903-329 : Needle wash unit, complete (Hct). |  | Defect | HW | Aspiration Module | Needle Wash Unit
AQT90 FLEX | WE NEED TO REPLACE OPTICAL UNIT | B | SWO-133054 | No | WE NEED TO REPLACE OPTICAL UNIT |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | Cleaned Internal Dust from barcode reader, Cleaned Machine Internally and restart, Working ok. | Cleaned Internal Dust from barcode reader, Cleaned Machine Internally and restart, Working ok. | SWO-132895 | No | Cleaned Internal Dust from barcode reader, Cleaned Machine Internally and restart, Working ok. | Cleaned Internal Dust from barcode reader, Cleaned Machine Internally and restart, Working ok. | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
ABL820 FLEX | Will call once Keith Clark is back from vacation | B | D165031F12552562 | No | Will call once Keith Clark is back from vacation |  | Revisit required | HW | Oxi module | Oxi module
ABL80 FLEX BASIC |  | B | SWO-132991 | No
ABL90 FLEX | Batterij vervangen. | Batterij vervangen. | SWO-133135 | No | Replace the battery. | Replace the battery. | Defect | HW | Battery module | Battery module
ABL90 FLEX | Akku ausgetauscht. Verbindungsdatem zum Radiance korrigiert. Ip Adresse 10.20.9.28 | Akku ausgetauscht. Verbindungsdatem zum Radiance korrigiert. Ip Adresse 10.20.9.28 | SWO-133246 | No | Battery replaced. Fixed the connection date to the radiance. Ip Address 10.20.9.28 | Battery replaced. Fixed the connection date to the radiance. Ip Address 10.20.9.28 | Defect | HW | Battery module | Battery module
ABL90 FLEX | Lösungspack wurde in den letzten Tagen mehrfach erneuert, Fehler kam aber immer wieder. Lösungspack und Sensorkasette erneut. Kalibriert und QC's durchgeführt. QC's in allen 3 Lvl bestanden. | Sr. Gertrude wünscht einen Service vor Ort. Bitte ein LP mitbringen! Die vorhandenen LPs wurden schon durchgetauscht. | SWO-133294 | No | Solution pack has been renewed several times in the last few days, but errors kept coming. solution pack and sensor calet again. Calibrated and QC's performed. QC's passed in all 3 Lvl. | Sr. Gertrude wants a service on site. Please bring an LP! The existing LPs have already been replaced. | Assisted Replacement | SC | SC | pO2
ABL90 FLEX | Fehlersuche, div. Test`s im Service durchgef., LP gew., Startphase abgewartet, Kal. + QC i.O. | Fehlersuche; div. Test`s im Service durchgef.; LP gew.; Startphase abgewartet; Kal. + QC i.O. | SWO-133220 | No | Troubleshooting; div. Test it in the service. LP. start-up phase awaited; Kal. + QC in the first place | Troubleshooting; div. Test it in the service. LP. start-up phase awaited; Kal. + QC in the first place | Assisted Replacement | SP | SP | FT
ABL90 FLEX PLUS |  | B | D1679ECEFE5A32D6 | No |  |  | Adjustment | HW | Auto Inlet module | Inlet drive gear
ABL825 FLEX |  | B | D167ABACFE5BFEBF | No |  |  | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Tausch des EL MET Moduls. LS Justage, Pumpenkal und Nullpunktkalibrierung. Elektrodenkontakte der Metabolite gebürstet. Kalibrationen und Qualitätskontrolle in Ordnung. Anwender beobachten weiterhin Entwicklung der Metabolite. | Fr. Brosius hat Probleme das Lac und LGluc seit 1,5 Wochen nicht stabil laufen und immer weider Driften aufweisen, sie hat schon Membranen anderer Lot benutz,Messkammer ist sauber Elektrodenspitzen wurden beim membranieren abgewischt, Sie wird jetzt noch einge Zeit abwarten und sich dann evt wieder melden | SWO-133216 | No | Replacement of the EL MET module. LS adjustment, pump scale and zero calibration. electrode contacts of the metabolites brushed. Calibrations and quality control in order. Users continue to observe the development of metabolites. | Ms. Brosius has problems that Lac and LGluc have not been running stable for 1.5 weeks and have always weider drifts, she has already used membranes of other solder, measuring chamber is clean electrode tips have been wiped off during membraneing, she will now wait a while and then contact you again as a long time | Cleaned | Fluidic pathway | El/Met Module | Glu/Lac | To be monitored
AQT90 FLEX | Checked and found that Barcode reader must have dusty. Clean the Barcode reader, Now AQT 90 was showing the TnT test. Run the System Cleaning cycle also Load the more TnT Cartridge from AQT90 FLEX 838R0307N027. Run the sample for observation | B | SWO-134755 | No | Checked and found that Barcode reader must have dusty. Clean the Barcode reader, Now AQT 90 was showing the TnT test. Run the System Cleaning cycle also Load the more TnT Cartridge from AQT90 FLEX 838R0307N027. Run the sample for observation |  | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | Cleaned Internal and Dust from Barcode reader, perform CKMB Calibration, Machine working ok. | Cleaned Internal and Dust from Barcode reader, perform CKMB Calibration, Machine working ok. | SWO-134766 | No | Cleaned Internal and Dust from Barcode reader, perform CKMB Calibration, Machine working ok. | Cleaned Internal and Dust from Barcode reader, perform CKMB Calibration, Machine working ok. | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX BASIC | cleaned waste path. installed new solution pack and sensor cassette. | cleaned waste path. installed new solution pack and sensor cassette. | SWO-133702 | No | cleaned waste path. installed new solution pack and sensor cassette. | cleaned waste path. installed new solution pack and sensor cassette.
ABL80 FLEX BASIC |  | B | SWO-133366 | No
ABL80 FLEX Basic version |  | B | D168A7D174326FF3 | No
ABL80 CO-OX |  | B | D168A86A74236FF5 | No
ABL90 FLEX | Die 600er Kassette war nach 20 Tagen aufgebraucht, eingesetzt am 07.09.16 Am gleichen Tag wurden zwei neue Kassetten bestellt, die am 09.09.16 geliefert wurden. Eine davon wurde nach Reinhardshausen geschickt, wo man sich die Ware ausgeliehen hatte. Die zweite Kassette wurde nicht gefunden. Habe Frau Bauer die 900er Kassetten empfohlen, da scheinbar mehr gemessen wird. Schwester Tina wird mit der Stationsleitung einen neuen Einweisungstermin vereinbaren, da hier dringender Schulungsbedarf besteht. Frau Bauer hat sich Kassette aus Asklepios Schlafzentrum geliehen, nachdem von HANSC vermittelt wurde. | Die 600er Kassette war nach 20 Tagen aufgebraucht, eingesetzt am 07.09.16 Am gleichen Tag wurden zwei neue Kassetten bestellt, die am 09.09.16 geliefert wurden. Eine davon wurde nach Reinhardshausen geschickt, wo man sich die Ware ausgeliehen hatte. Die zweite Kassette wurde nicht gefunden. Habe Frau Bauer die 900er Kassetten empfohlen, da scheinbar mehr gemessen wird. Schwester Tina wird mit der Stationsleitung einen neuen Einweisungstermin vereinbaren, da hier dringender Schulungsbedarf besteht. Frau Bauer hat sich Kassette aus Asklepios Schlafzentrum geliehen, nachdem von HANSC vermittelt wurde. | SWO-134028 | Yes | The 600 cassette was used up after 20 days, used on 07.09.16 On the same day two new cassettes were ordered, which were delivered on 09.09.16. One of them was sent to Reinhardshausen, where the goods had been borrowed. The second cartridge could not be found. Ms. Bauer recommended the 900 cassettes, as apparently more is measured. Sister Tina will arrange a new briefing date with the station management, as there is an urgent need for training. Ms. Bauer borrowed a cassette from Asklepio's sleep center after being brokered by HANSC. | The 600 cassette was used up after 20 days, used on 07.09.16 On the same day two new cassettes were ordered, which were delivered on 09.09.16. One of them was sent to Reinhardshausen, where the goods had been borrowed. The second cartridge could not be found. Ms. Bauer recommended the 900 cassettes, as apparently more is measured. Sister Tina will arrange a new briefing date with the station management, as there is an urgent need for training. Ms. Bauer borrowed a cassette from Asklepio's sleep center after being brokered by HANSC. | Re classified RMED | Logistical upSWO
ABL90 FLEX | Ventil und Ventilschlauch erneuert. SP90 und SC90 erneuert. | Ventil und Ventilschlauch erneuert. SP90 und SC90 erneuert. | SWO-133575 | No | valve and valve hose. SP90 and SC90. | valve and valve hose. SP90 and SC90. | Defect | HW | Oxi module | Pinch valve
ABL90 FLEX PLUS | Check out of inlet module. Check ribbon cable . Check needle and guide, washer, all ok. Perform calibration. No led to pick up inlet. | B | SWO-134098 | No | Check out of inlet module. Check ribbon cable . Check needle and guide, washer, all ok. Perform calibration. No led to pick up inlet. |  | Adjustment | HW | Auto Inlet module | Inlet Position Cal.
ABL835 FLEX | pO2 elektróda, pO2 elektróda fedél, gáz tömítő gyűrű, pH BG modul mérőkamra fölötti gumicső csere, kalibrációk, QC mérés | B | SWO-133532 | No | pO2 electrode, pO2 electrode cover, gas sealing ring, pH BG module over measuring chamber rubber tube replacement, calibration, QC measurement |  | Assisted Replacement | Electrode | Electrode | pO2
ABL825 FLEX | 廃液ｺﾈｸﾀｰ交換 | B | SWO-133597 | No | Waste connector replacement |  | Defect | HW | Inlet module | Inlet gasket
ABL827 FLEX | Einlassstutzen gereinigt. Gerät kalibriert und QC durchgeführt. | Einlassstutzen gereinigt. Gerät kalibriert und QC durchgeführt. | SWO-133580 | No | Inlet nozzle cleaned. calibrated and QC performed. | Inlet nozzle cleaned. calibrated and QC performed. | Cleaned | Fluidic pathway | Inlet module | Inlet module
ABL825 FLEX | Found a file open in background may have been causing freezing. Found bent inlet proe and a suspect waste pump rotor. | Replaced pump rotor and inlet gasket. Performed all necessary calibrations. Rebooted analyzer. | D168F8FF74573E74 | No | Found a file open in background may have been causing freezing. Found bent inlet proe and a suspect waste pump rotor. | Replaced pump rotor and inlet gasket. Performed all necessary calibrations. Rebooted analyzer. | Defect | HW | Inlet module | Probe bend
ABL815 FLEX | Membranierung: 22.09.2016, Status 306,7 pA mM, Kal-Status 27.09.2016: 192,5 pA mM Membranierung Lac am 27.09.2016, Kontrolle am 11.10.2016 | Membranierung: 22.09.2016, Status 306,7 pA mM, Kal-Status 27.09.2016: 192,5 pA mM Membranierung Lac am 27.09.2016, Kontrolle am 11.10.2016 | SWO-133382 | No | Membraneing: 22.09.2016, status 306.7 pA mM, Kal status 27.09.2016: 192.5 pA mM membraneing Lac on 27.09.2016, control on 11.10.2016 | Membraneing: 22.09.2016, status 306.7 pA mM, Kal status 27.09.2016: 192.5 pA mM membraneing Lac on 27.09.2016, control on 11.10.2016 | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
AQT90 FLEX | Reinstalled analyzer software. | Reinstalled analyzer software. | SWO-134513 | No | Reinstalled analyzer software. | Reinstalled analyzer software. | Defect | SW | Analyzer Control Module | Reinstall Main SW
ABL80 FLEX BASIC | 现场连续更换2张测试卡后修复 | 现场连续更换2张测试卡后修复 | SWO-134002 | No | Repair after 2 consecutive test cards replaced on site | Repair after 2 consecutive test cards replaced on site
ABL90 | Passes Complete pO2 in diagnostics. | Replaced suspect Sensor Interface Bd. | D16A4640EA39F73C | No | Passes Complete pO2 in diagnostics. | Replaced suspect Sensor Interface Bd. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX PLUS | Found windows error message asking to restart computer running in background, restarted analyser to close message. Message now resolved, asked customer to monitor and report back to Roger Smith | B | SWO-134307 | No | Found windows error message asking to restart computer running in background, restarted analyser to close message. Message now resolved, asked customer to monitor and report back to Roger Smith |  | Defect | SW | Windows | Windows | Slow response
ABL90 FLEX PLUS |  | B | SWO-134732 | Yes |  |  | Re classified RMED | Double entry | See | SWO-134728 | CAS-145982-X9D5
ABL90 FLEX PLUS | Inspected mechanism and inlet, all ok. Check motor coupler, all ok. The movement is in the motor not holding back the spring pressure. Stripped down inlet assembly and reassembled correctly. Now ok. | B | SWO-133893 | No | Inspected mechanism and inlet, all ok. Check motor coupler, all ok. The movement is in the motor not holding back the spring pressure. Stripped down inlet assembly and reassembled correctly. Now ok. |  | Adjustment | HW | Auto Inlet module | Inlet Position Cal.
ABL90 FLEX PLUS | Tried to calibrate with both exe and control panel. Both options timed out every time. Left them with a mouse so that they can still use analyser. | Tried to calibrate with both exe and control panel. Both options timed out every time. Left them with a mouse so that they can still use analyser. | SWO-133894 | No | Tried to calibrate with both exe and control panel. Both options timed out every time. Left them with a mouse so that they can still use analyser. | Tried to calibrate with both exe and control panel. Both options timed out every time. Left them with a mouse so that they can still use analyser. | Revisit required | HW | DMS_PC | Touch screen
TCM5 FLEX |  | B | D16A12BCEA3049DD | No
AQT90 FLEX | 更换903-364后需要重新定标，加借定标卡。 | 更换903-364后需要重新定标，加借定标卡。 | SWO-133689 | No | After replacing the 903-364, the calibration card needs to be re-ordered and the bid card is borrowed. | After replacing the 903-364, the calibration card needs to be re-ordered and the bid card is borrowed. | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | 更换903-364后需要重新定标，加借定标卡。 | SWO-133695 | No |  | After replacing the 903-364, the calibration card needs to be re-ordered and the bid card is borrowed. | Defect | HW | Process Module | Optical Unit | 131
ABL800 FLEX | Upload new software now machine is running ok. | Upload new software now machine is running ok. | SWO-134714 | No | Upload new software now machine is running ok. | Upload new software now machine is running ok. | Defect | SW | SW | SW | LEH Mail
ABL815 FLEX | Kalibrationslösungen und Spüllösung getauscht Kal- und QC QC Messungen beim Na auf Zielwert Patientenmessung durchgeführt - Na mit erwartetem Wert gemessen Kunden angeregt, die Lösungen wegen Wärmeentwicklung einer Lampe, in einen anderen Schrank zu räumen. | Kalibrationslösungen und Spüllösung getauscht Kal- und QC QC Messungen beim Na auf Zielwert Patientenmessung durchgeführt - Na mit erwartetem Wert gemessen Kunden angeregt, die Lösungen wegen Wärmeentwicklung einer Lampe, in einen anderen Schrank zu räumen. | SWO-133852 | No | Calibration solutions and rinsing solution swapped Kal and QC QC measurements performed at The Na at target patient measurement - Na with expected value measured customers stimulated to vacate the solutions due to heat development of a lamp, in another cabinet. | Calibration solutions and rinsing solution swapped Kal and QC QC measurements performed at The Na at target patient measurement - Na with expected value measured customers stimulated to vacate the solutions due to heat development of a lamp, in another cabinet. | Assisted Replacement | Consumables | Solution bottle | Solutions bottle | Na
ABL817 FLEX | alle Schlauchwege ern. - siehe JW, alle Ventile gereinigt, Lösungen ern., Messkammer gerienigt, Gerät kalibriert, QC´s durchgef., Problem besteht weiterhin - andere QC Lot vom Level 1 besorgt und eingelesen, QC´s durchgef. - Na Zielwert, tHb 0,1 g dL über dem Zielwert Die fehlerhafte QC S7385 Lot R0080 abgezogen | alle Schlauchwege ern. - siehe JW, alle Ventile gereinigt, Lösungen ern., Messkammer gerienigt, Gerät kalibriert, QC´s durchgef., Problem besteht weiterhin - andere QC Lot vom Level 1 besorgt und eingelesen, QC´s durchgef. - Na Zielwert, tHb 0,1 g dL über dem Zielwert Die fehlerhafte QC S7385 Lot R0080 abgezogen | SWO-133899 | No | all hoseways. - see JW, all valves cleaned, solutions ern., measuring chamber rubbed, device calibrated, QC's through, problem persists - other QC lot from level 1 concerned and read in, QC's through. - Na Target value, tHb 0.1 g dL above the target value The faulty QC S7385 Lot R0080 subtracted | all hoseways. - see JW, all valves cleaned, solutions ern., measuring chamber rubbed, device calibrated, QC's through, problem persists - other QC lot from level 1 concerned and read in, QC's through. - Na Target value, tHb 0.1 g dL above the target value The faulty QC S7385 Lot R0080 subtracted | User error | Consumables | QC5+ | QC5+
ABL800 BASIC | 症状再発の為、PH BG,EL Metモジュールの測定チャンバー交換。Ref電極交換。PH BG,EL Metモジュールの各バルブ整備。サンプル移動エラー及電解質 代謝ポンプ同調エラーの為、EL Metモジュールポンプチューブ交換。 | 症状再発の為、PH BG,EL Metモジュールの測定チャンバー交換。Ref電極交換。PH BG,EL Metモジュールの各バルブ整備。サンプル移動エラー及電解質 代謝ポンプ同調エラーの為、EL Metモジュールポンプチューブ交換。 | SWO-133841 | No | Replacement of measurement chamber of PH BG and EL Met modules for symptom recurrence. Ref electrode replacement. Valve maintenance of PH BG and EL Met modules. El Met module pump tube replacement due to sample movement error and electrolyte metabolic pump tuning error. | Replacement of measurement chamber of PH BG and EL Met modules for symptom recurrence. Ref electrode replacement. Valve maintenance of PH BG and EL Met modules. El Met module pump tube replacement due to sample movement error and electrolyte metabolic pump tuning error. | Defect | HW | El/Met Module | Measuring chamber | Na
ABL835 FLEX | Found cause to be broken inlet on inlet module. | Replaced inlet nipple and retubed | D16B274860121736 | No | Found cause to be broken inlet on inlet module. | Replaced inlet nipple and retubed | Defect | HW | Inlet module | Nipple kit
ABL800 BASIC | Replaced lactate membrane | Replaced lactate membrane | SWO-134785 | No | Replaced lactate membrane | Replaced lactate membrane | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Case aufgenommen - Gerät wird demnächst besucht zwecks Datenaufnahme. | B | SWO-133901 | Yes | Case recorded - device will be visited soon for data recording. |  | Re classified RMED | Phone
ABL825 FLEX | Laut Herrn Möbus ist der Fehler mit dem hängenden Monitor letzte Woche wieder aufgetreten, zurzeit arbeitet der TS einwandfrei. Möbus konnte den genauen Termin des Ausfalls nicht mehr nennen, wird weiter beobachtet! | Laut Herrn Möbus ist der Fehler mit dem hängenden Monitor letzte Woche wieder aufgetreten; zurzeit arbeitet der TS einwandfrei. Möbus konnte den genauen Termin des Ausfalls nicht mehr nennen; wird weiter beobachtet! | SWO-133874 | No | According to Mr Möbus, the error with the hanging monitor occurred again last week; currently the TS is working properly. Möbus could no longer give the exact date of the outage; will continue to be observed! | According to Mr Möbus, the error with the hanging monitor occurred again last week; currently the TS is working properly. Möbus could no longer give the exact date of the outage; will continue to be observed! | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Fehlersuche: pO2 Elektrode überprüft, wurde mit pCO2 Membrane membraniert! pO2 korrekt membraniert, Startverfahren aufgerufen. Kal und QC Messungen erfolgen durch Anwender. | Fehlersuche: pO2 Elektrode überprüft, wurde mit pCO2 Membrane membraniert! pO2 korrekt membraniert, Startverfahren aufgerufen. Kal und QC Messungen erfolgen durch Anwender. | SWO-133927 | No | Troubleshooting: pO2 electrode checked, was membraneed with pCO2 membrane! pO2 correctly membraneed, start method called. Kal and QC measurements are carried out by users. | Troubleshooting: pO2 electrode checked, was membraneed with pCO2 membrane! pO2 correctly membraneed, start method called. Kal and QC measurements are carried out by users. | User error | Electrode | Membrane | pO2 | pO2 <> pCO2 membr.
ABL815 FLEX | Nadel wurde durch A. Klubert bereits ersetzt, defekte Nadel geht an RDEU Service zur Untersuchung. Ersatznadel übergeben. | Nadel wurde durch A. Klubert bereits ersetzt; defekte Nadel geht an RDEU Service zur Untersuchung. Ersatznadel übergeben. | SWO-133892 | No | Nadel has already been replaced by A. Klubert; defective needle goes to RDEU Service for investigation. Replacement needle. | Nadel has already been replaced by A. Klubert; defective needle goes to RDEU Service for investigation. Replacement needle. | Defect | HW | Inlet module | Probe bend
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-134521 | No | Replacement test card repair | Replacement test card repair
ABL80 FLEX BASIC | reinstalled sc80 . Calibration fail. sc80 flush & reinstalled sc & sp but fail to calibrate. Replace new SC80. Calibration pass. | reinstalled sc80 . Calibration fail. sc80 flush & reinstalled sc & sp but fail to calibrate. Replace new SC80. Calibration pass. | SWO-134410 | No | reinstalled sc80 . Calibration fail. sc80 flush & reinstalled sc & sp but fail to calibrate. Replace new SC80. Calibration pass. | reinstalled sc80 . Calibration fail. sc80 flush & reinstalled sc & sp but fail to calibrate. Replace new SC80. Calibration pass.
ABL80 FLEX CO-OX | Fehlersuche : Lösungsaustritt am oberen (2.) Anschlussstutzen , Lösungspack hatte noch Flüssigkeiten für 35 Aktivitäten, Reinigung des Analysators, Neuinstallation Lösungspack, Funktionskontrolle | Fehlersuche : Lösungsaustritt am oberen (2.) Anschlussstutzen , Lösungspack hatte noch Flüssigkeiten für 35 Aktivitäten, Reinigung des Analysators, Neuinstallation Lösungspack, Funktionskontrolle | SWO-134446 | No | Troubleshooting : Solution outlet at the upper (2nd) connection connector, solution pack still had liquids for 35 activities, cleaning of the analyzer, new installation solution pack, function control | Troubleshooting : Solution outlet at the upper (2nd) connection connector, solution pack still had liquids for 35 activities, cleaning of the analyzer, new installation solution pack, function control
ABL90 FLEX | LP erneuert, Fehler bleibt bestehen. Zusätzlich ein LP als Reserve vor Ort gelassen. SC erneuert. pCO2 Werte liegen nun wieder im normalen Bereich. Lac ist allerdings nun zu hoch. Mehrfach kalibriert und QC Kontrollen durchgeführt. Lac kommt immer näher an den QC Bereich ran. Jedoch dauert dies noch ein paar Stunden. Station Bescheid gegeben, falls Lac bis morgen nicht im Bereich sein sollte wird sich die Station nochmal melden. | LP erneuert, Fehler bleibt bestehen. Zusätzlich ein LP als Reserve vor Ort gelassen. SC erneuert. pCO2 Werte liegen nun wieder im normalen Bereich. Lac ist allerdings nun zu hoch. Mehrfach kalibriert und QC Kontrollen durchgeführt. Lac kommt immer näher an den QC Bereich ran. Jedoch dauert dies noch ein paar Stunden. Station Bescheid gegeben, falls Lac bis morgen nicht im Bereich sein sollte wird sich die Station nochmal melden. | SWO-134182 | No | LP renewed, error remains. In addition, an LP as a reserve left on site. SC renewed. pCO2 values are now back in the normal range. Lac, however, is now too high. Multiple calibrated and QC checks performed. Lac is getting closer and closer to the QC area. However, this still takes a few hours. Station informed, if Lac is not in the area by tomorrow, the station will contact again. | LP renewed, error remains. In addition, an LP as a reserve left on site. SC renewed. pCO2 values are now back in the normal range. Lac, however, is now too high. Multiple calibrated and QC checks performed. Lac is getting closer and closer to the QC area. However, this still takes a few hours. Station informed, if Lac is not in the area by tomorrow, the station will contact again. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | Fehlersuche, SC + LP installiert, Startphase abgewartet, Konfig. angepaßt, Kal. + QC i.O. | Fehlersuche; SC + LP installiert; Startphase abgewartet; Konfig. angepaßt; Kal. + QC i.O. | SWO-134078 | No | Troubleshooting; SC + LP installed; start-up phase awaited; Config. adapted; Kal. + QC in the first place | Troubleshooting; SC + LP installed; start-up phase awaited; Config. adapted; Kal. + QC in the first place | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS | Replace inlet module Replace inlet pickup pcb No change to unit operational procedure. Perform complete check. Perform re install of software . DMS installation not successful. 2nd time install but the program cannot continue message on screen. Data upgrade failed - Status 2 code | B | SWO-134728 | No | Replace inlet module Replace inlet pickup pcb No change to unit operational procedure. Perform complete check. Perform re install of software . DMS installation not successful. 2nd time install but the program cannot continue message on screen. Data upgrade failed - Status 2 code |  | Defect | HW | Auto Inlet module | Auto Inlet module
TCM5 FLEX |  | B | D16B5E9F60593C94 | No
ABL800 FLEX | Replaced the Gas 1 in ABL800. Checked the Activity logs and meet with Mr. Deepak (MICU), Discussed about if the operator training is required. Run the samples, Now ABL800 is working satisfactory. | B | SWO-134454 | No | Replaced the Gas 1 in ABL800. Checked the Activity logs and meet with Mr. Deepak (MICU), Discussed about if the operator training is required. Run the samples, Now ABL800 is working satisfactory. |  | Assisted Replacement | Consumables | Gas cylinder | Gas 1
ABL825 FLEX | Remplacé la tête de l'imprimante thermique e.o garantie Calibration e.o contrôle de qualité e.o | Remplacé la tête de l'imprimante thermique e.o garantie Calibration e.o contrôle de qualité e.o | SWO-134096 | No | Replaced the head of the thermal printer e.o guaranteed Calibration e.o quality control e.o | Replaced the head of the thermal printer e.o guaranteed Calibration e.o quality control e.o | Defect | HW | Printer module | Printer Head
ABL815 FLEX | Gerät geprüft, Schlauchsystem gespült, Pumpenschlauch im Elyte Modul ersetzt, Einlassdichtung (aus Kundenbestand) ausgetauscht, Funktionskontrolle ok | Gerät geprüft, Schlauchsystem gespült, Pumpenschlauch im Elyte Modul ersetzt, Einlassdichtung (aus Kundenbestand) ausgetauscht, Funktionskontrolle ok | SWO-134465 | No | Device tested, hose system flushed, pump hose replaced in Elyte module, inlet seal (from customer stock) replaced, function control ok | Device tested, hose system flushed, pump hose replaced in Elyte module, inlet seal (from customer stock) replaced, function control ok | Defect | HW | El/Met Module | Pump tubes
ABL800 BASIC | Barcode was not clear. so entered bar code of cylinder manually. Then restart system. Unit is working ok. | Barcode was not clear. so entered bar code of cylinder manually. Then restart system. Unit is working ok. | SWO-134336 | No | Barcode was not clear. so entered bar code of cylinder manually. Then restart system. Unit is working ok. | Barcode was not clear. so entered bar code of cylinder manually. Then restart system. Unit is working ok. | Defect | Consumables | Gas cylinder | Barcode Quality
ABL825 FLEX | Konfiguration geprüft, Datenquelle Abfrage angepasst, Funktionskontrolle ok | Konfiguration geprüft, Datenquelle Abfrage angepasst, Funktionskontrolle ok | SWO-134466 | No | Configuration checked, data source query customized, function control ok | Configuration checked, data source query customized, function control ok | Setup | His/Lis | His/Lis | His/Lis
ABL800 BASIC | Cleaned Electrodes chamber and replace cal1, cal2 solutions. | Cleaned Electrodes chamber and replace cal1, cal2 solutions. | SWO-134802 | No | Cleaned Electrodes chamber and replace cal1, cal2 solutions. | Cleaned Electrodes chamber and replace cal1, cal2 solutions. | Assisted Replacement | Electrode | Membrane | Reference | Why not REFERENCE?
ABL80 FLEX CO-OX | neuen Hemolysator eingebaut, Gerät kalibriert, | neuen Hemolysator eingebaut, Gerät kalibriert, | SWO-134614 | No | new hemolysator installed, device calibrated, | new hemolysator installed, device calibrated,
ABL80 FLEX BASIC | sc80 & sp80 fail to calibrate. reinstalled sc80 & sp80. the sensor connector faulty . need to check the Sensor connector board. | sc80 & sp80 fail to calibrate. reinstalled sc80 & sp80. the sensor connector faulty . need to check the Sensor connector board. | SWO-134737 | No | sc80 & sp80 fail to calibrate. reinstalled sc80 & sp80. the sensor connector faulty . need to check the Sensor connector board. | sc80 & sp80 fail to calibrate. reinstalled sc80 & sp80. the sensor connector faulty . need to check the Sensor connector board.
ABL80 FLEX CO-OX | SC und LP erneuert, da der Status nicht eindeutig war, Gerät kalibriert, QC´s durchgef. Der Kunde hat immer wieder verschiedene Probleme - anschließend muss das LP oder die SC erneuert werden. Mit dem Kunden die möglichen Fehlerursachen besprochen - LP und SC - Lagertemperatur - Das Gerät sollte nicht trocken laufen, davon kann das Gerät und das Zubehör einen Schaden nehmen Daten-Log. erstellt | Der Kunde bittet um einen Service und Rückruf. | SWO-134673 | No | SC and LP renewed because the status was not unique, device calibrated, QC's through. The customer always has different problems - after that the LP or sc has to be renewed. Discussed with the customer the possible causes of errors - LP and SC - storage temperature - The device should not run dry, of which the device and the accessories can take a damage data log. Created | The customer asks for a service and callback.
ABL80 CO-OX | QC low out of limits on S7730 lot 146 QC. | Checked zero settings. OK. No supplies available. Prepared credit forms for 2 sensors and 1 solution pack for overnight delivery. | D16C9055D615816F | No | QC low out of limits on S7730 lot 146 QC. | Checked zero settings. OK. No supplies available. Prepared credit forms for 2 sensors and 1 solution pack for overnight delivery.
ABL80 CO-OX | pO2 QC low out of limits on level 1 S7730 lot 146 QC on two new analyzers. | Checked zero settings. OK. Replacing solution pack no help. Replaced sensor. Used customers stock. | D16C856CD612816B | No | pO2 QC low out of limits on level 1 S7730 lot 146 QC on two new analyzers. | Checked zero settings. OK. Replacing solution pack no help. Replaced sensor. Used customers stock.
ABL90 FLEX | Nach dem Neustart Uhrzeit I.O. | Nach dem Neustart Uhrzeit I.O. | SWO-134586 | No | After restarting time I.O. | After restarting time I.O. | Setup | SW | Time setting | Time setting
ABL90 FLEX | Gerät nochmal überprüft. Alles ok. SD gespeichert. | Gerät nochmal überprüft. Alles ok. SD gespeichert. | SWO-134554 | Yes | device checked again. All ok. SD. | device checked again. All ok. SD. | Re classified RMED | PER data collect
ABL90 FLEX PLUS | Tausch der CF Karte. Einspielen des Betriebssystems und der Applikations Software. Einspielung erfolgreich. Konfiguration des Analysators und Überprüfung, Funktion in Ordnung. Konfiguration der Netzwerkverbindung, Analysator verbunden, Daten werden übertragen. tHb Kalibration mit eigenem Kalibrator durchgeführt. Kalibrationen und Qualitätskontrolle in Ordnung. | Tausch der CF Karte. Einspielen des Betriebssystems und der Applikations Software. Einspielung erfolgreich. Konfiguration des Analysators und Überprüfung, Funktion in Ordnung. Konfiguration der Netzwerkverbindung, Analysator verbunden, Daten werden übertragen. tHb Kalibration mit eigenem Kalibrator durchgeführt. Kalibrationen und Qualitätskontrolle in Ordnung. | SWO-134692 | No | Exchange of the CF card. Inserting the operating system and application software. Recording successful. Configuration of the analyzer and verification, function in order. Configuration of the network connection, analyzer connected, data is transmitted. tHb calibration with its own calibrator. Calibrations and quality control in order. | Exchange of the CF card. Inserting the operating system and application software. Recording successful. Configuration of the analyzer and verification, function in order. Configuration of the network connection, analyzer connected, data is transmitted. tHb calibration with its own calibrator. Calibrations and quality control in order. | Defect | HW | DMS_PC | CF card
ABL90 FLEX PLUS | Replaced touch screen. | Replaced touch screen. | SWO-134821 | No | Replaced touch screen. | Replaced touch screen. | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX | 更换903-158修复(by andrew) | 更换903-158修复(by andrew) | SWO-133739 | No | Replacement 903-158 repair (by andrew) | Replacement 903-158 repair (by andrew) | Defect | HW | Basic Unit | Solution Pack PCB Id
ABL835 FLEX | Vloeistofkanaal doorgespoten met demiwater en nieuwe tubing gemonteerd. Liquid sensoren gekalibreerd nu zonder cal fout. LS-oxi 123. | Vloeistofkanaal doorgespoten met demiwater en nieuwe tubing gemonteerd. Liquid sensoren gekalibreerd nu zonder cal fout. LS-oxi 123. | SWO-134627 | No | Liquid channel sprayed with demi water and new tubing mounted. Liquid sensors calibrated now without CAL error. LS-Oxi 123. | Liquid channel sprayed with demi water and new tubing mounted. Liquid sensors calibrated now without CAL error. LS-Oxi 123. | Defect | HW | Oxi module | Tube
ABL800 FLEX | Verified the available UPS technically. Connect the ABL800 on UPS. Meet with Biomedical and discussed about the Radiometer samplers to avoid the clotting to minimize the breakdown. | B | SWO-134458 | No | Verified the available UPS technically. Connect the ABL800 on UPS. Meet with Biomedical and discussed about the Radiometer samplers to avoid the clotting to minimize the breakdown. |  | Setup | Ext HW | UPS | UPS | LEH Mail
ABL810 FLEX | Demokeys deaktiviert und Parameterkeys neu installiert initialisiert. Kalibriert und QC's durchgeführt. | B | SWO-134629 | No | Demokeys disabled and parameter keys reinstalled initialized. Calibrated and QC's performed. |  | Setup | Demo key exp | Demo Key Exp | Demo key expired
ABL825 FLEX | Austausch Druckkopf Testausdruck i.O. | Austausch Druckkopf Testausdruck i.O. | SWO-134657 | No | Replacement print head test expression i.O. | Replacement print head test expression i.O. | Defect | HW | Printer module | Printer Head
ABL825 FLEX | Einlassnadel erneuert, Gerät kalibriert, QC durchgef. | Einlassnadel erneuert, Gerät kalibriert, QC durchgef. | SWO-134661 | No | Inlet needle renewed, device calibrated, QC through. | Inlet needle renewed, device calibrated, QC through. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Datenextrakt gezogen zwecks Weiterleitung an RDEU Service. | Datenextrakt gezogen zwecks Weiterleitung an RDEU Service. | SWO-134610 | No | Data extract drawn for forwarding to RDEU Service. | Data extract drawn for forwarding to RDEU Service. | Unknown | SW | Auto Check Module | Repeated QC
ABL815 FLEX | Nullstromabgleich im Servicemenü durchgeführt, 2 Pkt. Kal, OK. | Nullstromabgleich im Servicemenü durchgeführt, 2 Pkt. Kal, OK. | SWO-134641 | No | Zero current leveling performed in the service menu, 2 pts. Kal, OK. | Zero current leveling performed in the service menu, 2 pts. Kal, OK. | Adjustment | HW | pH/BG Module | Zero current | pO2
ABL815 FLEX | Dateien kontrolliert, Nasssektion sichtkontrolliert, keine Fehler feststellbar. Aufruf einer Kal 1 => Driftwert =0 | Dateien kontrolliert, Nasssektion sichtkontrolliert, keine Fehler feststellbar. Aufruf einer Kal 1 => Driftwert =0 | SWO-134667 | No | Files controlled, wet section visually controlled, no errors detected. Calling a Kal 1 => Drift value =0 | Files controlled, wet section visually controlled, no errors detected. Calling a Kal 1 => Drift value =0 | No fault found | No fault found | No fault found | No fault found
AQT90 FLEX | defekte Kassette entfernt ( die Führung der Kassette ist gebrochen), alle Kassetten entfernt und auf defekt kontrolliert aber die restlichen sind IO, Startphase Systemüberprüfung und Systemreinigung ohne Fehler durchgelaufen System wieder einsatzbereit | defekte Kassette entfernt ( die Führung der Kassette ist gebrochen), alle Kassetten entfernt und auf defekt kontrolliert aber die restlichen sind IO, Startphase Systemüberprüfung und Systemreinigung ohne Fehler durchgelaufen System wieder einsatzbereit | SWO-134852 | No | defective cassette removed (the cartridge's guide is broken), all cassettes removed and checked for defective but the rest are IO, start phase system check and system cleaning without fault sit gone system ready for use again | defective cassette removed (the cartridge's guide is broken), all cassettes removed and checked for defective but the rest are IO, start phase system check and system cleaning without fault sit gone system ready for use again | Defect | Consumables | Test Kit | Cartridge | Broken
ABL80 FLEX BASIC |  | B | SWO-134837 | No
ABL80 CO-OX | Results suppressed on about 50% of samples. Error 292 and 581. Solution pack appears to not be sealing. | Replaced pack, checked internal tubing and pinch valve. | D1702178380739EB | No | Results suppressed on about 50% of samples. Error 292 and 581. Solution pack appears to not be sealing. | Replaced pack, checked internal tubing and pinch valve.
ABL90 FLEX | A tesztelt műszer visszaszállítása és üzembehelyezése, Kalibrálás, QC mérések | B | SWO-135059 | No | Redelivery and commissioning of the tested instrument, calibration, QC measurements |  | Defect | HW | Inlet module | Connector gasket
ABL825 FLEX |  | B | D17082D33861DD10 | Yes |  |  | Re classified RMED | Remote support | Demo key exp
ABL800 BASIC | cleaned the pressure transducer but it is not working.Cleaned the touch screen and calibrated the touch screen,kept instrument for under observation presently working. | cheeked UPS power supply to instrument ,Instrument internal connections found in good condition.re installed the Operating system, application software but problem remaining same . maybe it need to replace with analyzer control PCB and CPU. | SWO-135243 | No | cleaned the pressure transducer but it is not working. Cleaned the touch screen and calibrated the touch screen,kept instrument for under observation presently working. | cheeked UPS power supply to instrument ,Instrument internal connections found in good condition.re installed the Operating system, application software but problem remaining same . maybe it need to replace with analyzer control PCB and CPU. | Cleaned | HW | DMS_PC | Touch screen
ABL800 BASIC | Calibrated the touch screen | B | SWO-142711 | No | Calibrated the touch screen |  | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | found that po2 is failing calibration for sensitivity. | clean the tip of the electrode which corrected the issue. | D1721D69AE352C2A | No | found that po2 is failing calibration for sensitivity. | clean the tip of the electrode which corrected the issue. | Cleaned | Electrode | Electrode | pO2
ABL825 FLEX | weak electrodes | Replaced PH and REF electrodes | D170913A385A6D22 | No | weak electrodes | Replaced PH and REF electrodes | Assisted Replacement | Electrode | Electrode | pH
ABL80 CO-OX |  | B | D175468310178405 | No
ABL90 | Problem reoccuring after replacement of Sensor Interface Bd. | Replaced suspect Wet Section Bd. | D1718FF7AE06442C | No | Problem reoccuring after replacement of Sensor Interface Bd. | Replaced suspect Wet Section Bd. | Defect | HW | Basic Unit | Wet sec PCB
ABL90 |  | B | D171A230AE46D8D0 | Yes |  |  | Re classified RMED | Setup
ABL90 | NONE NECESSARY | REPLACED POWER SUPPLY | D171AC36AE541A73 | No | NONE NECESSARY | REPLACED POWER SUPPLY | Defect | HW | Power Supply | Power Supply
ABL90 | REPLACED POWER SUPPLY AND SYSTEM SHOWS NO CHANGE | ORDERED NEW POWER SUPPLY | D171AAECAE511A6F | No | REPLACED POWER SUPPLY AND SYSTEM SHOWS NO CHANGE | ORDERED NEW POWER SUPPLY | Defect | HW | Power Supply | Power Supply
ABL90 | TECHLINE RECOMENDS POWER SUPPLY REPLACEMENT | REPLACED POWER SUPPLY | D171A89EAE1F1A6B | No | TECHLINE RECOMENDS POWER SUPPLY REPLACEMENT | REPLACED POWER SUPPLY | Defect | HW | Power Supply | Power Supply
ABL90 FLEX PLUS | Checked unit, back flush | Checked unit, back flush | SWO-135588 | No | Checked unit, back flush | Checked unit, back flush | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS |  | B | D171871EAE4BC12C | No |  |  | Defect | HW | Cassette retainer | Cassette retainer | Clot
TCM4 |  | B | D1718DC0AE5971A7 | No
ABL835 FLEX | İNLET KASKET VE NİPPLE DEĞİŞTİRİLEREK SORUN GİDERİLDİ. | İNLET KASKET VE NİPPLE DEĞİŞTİRİLEREK SORUN GİDERİLDİ. | SWO-135928 | Yes | THE PROBLEM WAS ADDRESSED BY CHANGING THE INLET HELMET AND The NIPPLE. | THE PROBLEM WAS ADDRESSED BY CHANGING THE INLET HELMET AND The NIPPLE.
ABL837 FLEX | Cleaned Hemolyzer with protein remover ran service THb cal and test passed ok. But still failing QC 3. New Hemolyzer required part on order . | Advised customer engineer required. | SWO-135194 | No | Cleaned Hemolyzer with protein remover ran service THb cal and test passed ok. But still failing QC 3. New Hemolyzer required part on order . | Advised customer engineer required. | Revisit required | HW | Oxi module | Hemolyzer
ABL825 FLEX | analyzer will not print activity logs, logged into radiance but hospital printer is not working. worked with hospital IT to resolve the printer error. | showed customer how to print the logs from the radiance computer | D171ABEBAE3BA00A | Yes | analyzer will not print activity logs, logged into radiance but hospital printer is not working. worked with hospital IT to resolve the printer error. | showed customer how to print the logs from the radiance computer | Re classified RMED | His/Lis | Setup | Ext printer
ABL820 FLEX | Oxiymetrieprobleme bei QCs und Messungen. Oximetriemodu (R0081N005) ausgebaut und neues Modul (2068N009) eingebaut. Halfen Kamp eingestellt. Complete Oxymetrietest durchgeführt --> i.O. Mehrere Kalibrierungen durchgeführ. Service tHb Kal, normale tHb Kal ( Fcuv:0,9904) und Kalverifizierung durchgeführt. --> i.O. QCs für alle Level durchgeführt und von Labor überprüfen lassen. --> i.O. ABL wieder Messbereit. | Oxiymetrieprobleme bei QCs und Messungen. Oximetriemodu (R0081N005) ausgebaut und neues Modul (2068N009) eingebaut. Halfen Kamp eingestellt. Complete Oxymetrietest durchgeführt --> i.O. Mehrere Kalibrierungen durchgeführ. Service tHb Kal, normale tHb Kal ( Fcuv:0,9904) und Kalverifizierung durchgeführt. --> i.O. QCs für alle Level durchgeführt und von Labor überprüfen lassen. --> i.O. ABL wieder Messbereit. | SWO-135338 | No | Oxiymetry problems in QCs and measurements. Oximetrymodu (R0081N005) and new module (2068N009) installed. Halfen Kamp set. Complete oxymetry test performed --> i.O. Performs multiple calibrations. Service tHb Kal, normal tHb Kal ( Fcuv:0,9904) and calverification performed. --> i.O. QCs for all levels performed and checked by laboratory. --> i.O. ABL ready to measure again. | Oxiymetry problems in QCs and measurements. Oximetrymodu (R0081N005) and new module (2068N009) installed. Halfen Kamp set. Complete oxymetry test performed --> i.O. Performs multiple calibrations. Service tHb Kal, normal tHb Kal ( Fcuv:0,9904) and calverification performed. --> i.O. QCs for all levels performed and checked by laboratory. --> i.O. ABL ready to measure again. | Defect | HW | Oxi module | Oxi module
ABL810 FLEX | Nach Wartungsdienst besteht Fehler weiterhin - Erneuerung pO2 Elektrode - Funktionsüberprüfung etc. ohne Beanstandung | Nach Wartungsdienst besteht Fehler weiterhin - Erneuerung pO2 Elektrode - Funktionsüberprüfung etc. ohne Beanstandung | SWO-135285 | No | After maintenance service, errors persist - renewal pO2 electrode - function check etc. without complaint | After maintenance service, errors persist - renewal pO2 electrode - function check etc. without complaint | Assisted Replacement | Electrode | Electrode | pO2
ABL800 BASIC | cheeked UPS power supply to instrument ,Instrument internal connections found in good condition.re installed the Operating system, application software but problem remaining same . maybe it need to replace with analyzer control PCB and CPU. | cheeked UPS power supply to instrument ,Instrument internal connections found in good condition.re installed the Operating system, application software but problem remaining same . maybe it need to replace with analyzer control PCB and CPU. | SWO-135261 | No | cheeked UPS power supply to instrument ,Instrument internal connections found in good condition.re installed the Operating system, application software but problem remaining same . maybe it need to replace with analyzer control PCB and CPU. | cheeked UPS power supply to instrument ,Instrument internal connections found in good condition.re installed the Operating system, application software but problem remaining same . maybe it need to replace with analyzer control PCB and CPU. | Revisit required | HW | DMS_PC | Touch screen
ABL800 BASIC | Check the analyzer found touch screen was not working, tried to calibrate the touch screen, but calibration was not accepting, tried replacing standby old display unit but it was not working, need to replace new display unit and CPU unit. | Check the analyzer found touch screen was not working, tried to calibrate the touch screen, but calibration was not accepting, tried replacing standby old display unit but it was not working, need to replace new display unit and CPU unit. | SWO-135893 | No | Check the analyzer found touch screen was not working, tried to calibrate the touch screen, but calibration was not accepting, tried replacing standby old display unit but it was not working, need to replace new display unit and CPU unit. | Check the analyzer found touch screen was not working, tried to calibrate the touch screen, but calibration was not accepting, tried replacing standby old display unit but it was not working, need to replace new display unit and CPU unit. | Revisit required | HW | DMS_PC | Touch screen
ABL80 FLEX CO-OX | Geräteaufstellung, Einlaufverfahren durchgeführt. | Geräteaufstellung, Einlaufverfahren durchgeführt. | SWO-135674 | No | device installation, inlet procedure carried out. | device installation, inlet procedure carried out.
ABL80 CO-OX | Checked the data logs and found the error message. | Replaced the hemolyzer. | D172B94C2405DAAA | No | Checked the data logs and found the error message. | Replaced the hemolyzer.
ABL80 FLEX BASIC | sc80 reinstalled. sc80 fail to calibrate. check the logs. Replace the sc80 solution pack. | sc80 reinstalled. sc80 fail to calibrate. check the logs. Replace the sc80 solution pack. | SWO-135631 | No | sc80 reinstalled. sc80 fail to calibrate. check the logs. Replace the sc80 solution pack. | sc80 reinstalled. sc80 fail to calibrate. check the logs. Replace the sc80 solution pack.
ABL80 OSM | checked screen function | completed a screen calibration | D172CDAA24313356 | No | checked screen function | completed a screen calibration
ABL80 CO-OX | Paula tried to put in 105 for the patient reference ranges. It converted back to 104. I tried it on the 80 I was working on and also tech support and Jim Gibbons were able to reproduce the issue. If you put any odd number past 103 for Glu reference range it will convert to the even number. The off number will not stay. | Passed on to tech support and opened a techline case. | D172D68F2459CBCF | No | Paula tried to put in 105 for the patient reference ranges. It converted back to 104. I tried it on the 80 I was working on and also tech support and Jim Gibbons were able to reproduce the issue. If you put any odd number past 103 for Glu reference range it will convert to the even number. The off number will not stay. | Passed on to tech support and opened a techline case.
ABL90 FLEX | Vor Ort manuelle QC Level 3 ausgelöst. Danach war der Analysator einsatzbereit. Mehrfach kal und QC wiederholt. Die Fehlermeldung trat nicht mehr auf. Die Flüssigkeiten werden sehr gut transportiert ( Visuell Flow Check). Po2 complete Test durchgeführt im Service Menü. Keine Fehlermeldung. Die Senorbay nachgestellt damit die Pins weiter unten sind. | Vor Ort manuelle QC Level 3 ausgelöst. Danach war der Analysator einsatzbereit. Mehrfach kal und QC wiederholt. Die Fehlermeldung trat nicht mehr auf. Die Flüssigkeiten werden sehr gut transportiert ( Visuell Flow Check). Po2 complete Test durchgeführt im Service Menü. Keine Fehlermeldung. Die Senorbay nachgestellt damit die Pins weiter unten sind. | SWO-135540 | No | On-site manual QC Level 3 triggered. After that, the analyzer was ready for use. Repeated several times kal and QC. The error message no longer occurred. The liquids are transported very well (Visual Flow Check). Po2 complete test performed in the service menu. No error message. The Senorbay is recreated so that the pins are further down. | On-site manual QC Level 3 triggered. After that, the analyzer was ready for use. Repeated several times kal and QC. The error message no longer occurred. The liquids are transported very well (Visual Flow Check). Po2 complete test performed in the service menu. No error message. The Senorbay is recreated so that the pins are further down. | No fault found | No fault found | No fault found | pO2
ABL90 FLEX | keine, Gerät nagelneu, Lösungspack nicht verfügbar, Gerät fehlerhaft verlassen | keine, Gerät nagelneu, Lösungspack nicht verfügbar, Gerät fehlerhaft verlassen | SWO-135625 | No | none, device brand new, solution pack not available, device left incorrectly | none, device brand new, solution pack not available, device left incorrectly | Revisit required | SP | SP | SP
ABL90 FLEX PLUS |  | B | D172C1E32445EFBD | No |  |  | No fault found | FT | FT | Bend probe
ABL90 FLEX PLUS | Straightened bent probe. Unblocked sample pathway. | Straightened bent probe. Unblocked sample pathway. | SWO-135629 | No | Straightened bent probe. Unblocked sample pathway. | Straightened bent probe. Unblocked sample pathway. | Replaced | Probe bend | Probe bend | Probe bend
ABL90 FLEX PLUS |  | B | D172EA382439485F | No |  |  | Defect | HW | DMS_PC | CPU
ABL800 FLEX | Taken the Data Backup Cleaned the fuses on Analyzer control PCB Start the ABL800, showing Leakage error Cleaned the Gasket and Needle and checked the complete fluidic path for blockage. Read all the Solution bottles barcode Replaced the Po2 membrane Do the 2 point calibration. | Taken the Data Backup Cleaned the fuses on Analyzer control PCB Start the ABL800, showing Leakage error Cleaned the Gasket and Needle and checked the complete fluidic path for blockage. Read all the Solution bottles barcode Replaced the Po2 membrane Do the 2 point calibration. | SWO-135362 | No | Taken the Data Backup Cleaned the fuses on Analyzer control PCB Start the ABL800, showing Leakage error Cleaned the Gasket and Needle and checked the complete fluidic path for blockage. Read all the Solution bottles barcode Replaced the Po2 membrane Do the 2 point calibration. | Taken the Data Backup Cleaned the fuses on Analyzer control PCB Start the ABL800, showing Leakage error Cleaned the Gasket and Needle and checked the complete fluidic path for blockage. Read all the Solution bottles barcode Replaced the Po2 membrane Do the 2 point calibration. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL837 FLEX | Replaced faulty Hemolyzer and setup and test ok Cal,QC ok passed ok | Advised customer engineer required. | SWO-135492 | No | Replaced faulty Hemolyzer and setup and test ok Cal,QC ok passed ok | Advised customer engineer required. | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX |  | B | D1741CF09A2E1523 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL827 FLEX | Found some blockage in the waste pump tubes Found some problems with the WDC reports | Replaced waste pump tubes and flushed the waste lines Replaced the inlet gasket | D172CDB8242557F5 | No | Found some blockage in the waste pump tubes Found some problems with the WDC reports | Replaced waste pump tubes and flushed the waste lines Replaced the inlet gasket | Defect | HW | Basic Unit | Waste tubes
ABL825 FLEX | warranty replacement looks like manufacturing flaw | replaced k electrode | D1730CE124013E1F | No | warranty replacement looks like manufacturing flaw | replaced k electrode | Assisted Replacement | Electrode | Electrode | K+
ABL825 FLEX | Bereits mit Medizintechnik (Herr Berlinghoff) mittels Kreuztausch getestet - Erneuerung pCO2 Elektrode etc. - Funktionsüberprüfung etc. ohne Beanstandung ! | Bereits mit Medizintechnik (Herr Berlinghoff) mittels Kreuztausch getestet - Erneuerung pCO2 Elektrode etc. - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-135461 | No | Already tested with medical technology (Mr. Berlinghoff) by means of cross-exchange - renewal pCO2 electrode etc. - functional check etc. without complaint ! | Already tested with medical technology (Mr. Berlinghoff) by means of cross-exchange - renewal pCO2 electrode etc. - functional check etc. without complaint ! | Assisted Replacement | Electrode | Electrode | pCO2
ABL835 FLEX | Allgemeine Kontrolle Nadel zu weit vorne, Inlet-Achse hat spiel Motor inkl. Achse ersetzt i.o | B | SWO-135586 | No | General control needle too far forward, Inlet axis has game motor incl. axis replaced i.o |  | Defect | HW | Inlet module | Inlet module
AQT90 FLEX | Pisa.ini dosyası içerisinde com port aktivasyonu için yapılan ayarda pisa dosyası bozulmuş. Hata düzeltilerek sorun giderildi. | Pisa.ini dosyası içerisinde com port aktivasyonu için yapılan ayarda pisa dosyası bozulmuş. Hata düzeltilerek sorun giderildi. | SWO-135721 | Yes | The pisa file is corrupted in the pisa.ini file in the setting for com port activation. Fixed the bug and fixed the problem. | The pisa file is corrupted in the pisa.ini file in the setting for com port activation. Fixed the bug and fixed the problem.
AQT90 FLEX | cleaned process wheel and optical unit. | cleaned the process wheel and optical unit now it is working satisfactorily. | SWO-136206 | No | cleaned process wheel and optical unit. | cleaned the process wheel and optical unit now it is working satisfactorily. | Cleaned | HW | Process Module | Process wheel
ABL90 FLEX | Szervízteszt, Átmosás, Bementi tűtartó csere, Kalibrálás, QC mérés | B | SWO-135742 | No | Service test, relaying, input pin replacement, calibration, QC measurement |  | Defect | HW | Inlet module | Inlet gasket
ABL90 FLEX | Zunächst per Remote im Servicemenü div Tests durchgeführt, dann vor Ort einlassdichtung sowie anschlussdichtung ersetzt, div Reinigungsarbeiten durchgeführt, Neustart, kalibrationen sowie QC Messungen | Zunächst per Remote im Servicemenü div Tests durchgeführt, dann vor Ort einlassdichtung sowie anschlussdichtung ersetzt, div Reinigungsarbeiten durchgeführt, Neustart, kalibrationen sowie QC Messungen | SWO-135992 | No | First performed remotely in the service menu div tests, then on-site inlet seal and connection seal replaced, div cleaning work performed, restart, calibrations and QC measurements | First performed remotely in the service menu div tests, then on-site inlet seal and connection seal replaced, div cleaning work performed, restart, calibrations and QC measurements | Defect | HW | Inlet module | Inlet gasket
ABL90 | Performed inspection, reviewed data logs, all QC in range, calibration acceptable. | Completed PM, updated software to 3.1MR8, customers installed fresh SC and SP of different lot#s. Reran linearity, all parameters now within specifications. | D174080D9A2DAA98 | No | Performed inspection, reviewed data logs, all QC in range, calibration acceptable. | Completed PM, updated software to 3.1MR8, customers installed fresh SC and SP of different lot#s. Reran linearity, all parameters now within specifications. | Assisted Replacement | SP | SP | pH
ABL90 FLEX | Replace power supply with non stock spare | Replace power supply with non stock spare | SWO-136137 | No | Replace power supply with non stock spare | Replace power supply with non stock spare | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | Replace power supply with non stock spare | Replace power supply with non stock spare | SWO-135852 | No | Replace power supply with non stock spare | Replace power supply with non stock spare | Defect | HW | Power Supply | Power Supply
AQT90 FLEX |  | B | SWO-135723 | No |  |  | Defect | HW | Process Module | BNP Sensor
ABL825 FLEX | 代替ユニットセット、引上げ修理 | B | SWO-135705 | No | Alternative unit set, raise repair |  | Defect | HW | Oxi module | Oxi module | Neon lamp
ABL800 BASIC | Try to fix probe but still same error continue. We have to change inlet probe. | Try to fix probe but still same error continue. We have to change inlet probe. | SWO-135684 | No | Try to fix probe but still same error continue. We have to change inlet probe. | Try to fix probe but still same error continue. We have to change inlet probe. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Lac und pCO2 membraniert mehrere Kalibrationen QC Level 1 und 2 beide Parameter nahe Mittelwert | Lac und pCO2 membraniert mehrere Kalibrationen QC Level 1 und 2 beide Parameter nahe Mittelwert | SWO-135726 | No | Lac and pCO2 membranes several calibrations QC Level 1 and 2 both parameters near mean | Lac and pCO2 membranes several calibrations QC Level 1 and 2 both parameters near mean | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Had to make Halogen adjustments blindly due to 6.16 software.Unable to adjust out Halogen values.Could not duplicate Assymmetry error. | Replaced Lamp Assy.Made necessary adjustments. | D1831E5E22477D52 | No | Had to make Halogen adjustments blindly due to 6.16 software. Unable to adjust out Halogen values. Could not duplicate Assymmetry error. | Replaced Lamp Assy.Made necessary adjustments. | Defect | HW | Oxi module | Lamp unit
ABL810 FLEX | El Met Modul (902-824R2339N001) inclusive 5 Dummyelektroden leihweise eingebaut. AC 5+ Lösungen für eingescannt. Kalibrationen und QC's vorgenommen. QC Level 3 wurde nach Einbau um 1,2mmhg höher gemessen. Änderung inAQURE von Herrn Kozelj vorgenommen. | El Met Modul (902-824R2339N001) inclusive 5 Dummyelektroden leihweise eingebaut. AC 5+ Lösungen für eingescannt. Kalibrationen und QC's vorgenommen. QC Level 3 wurde nach Einbau um 1,2mmhg höher gemessen. Änderung inAQURE von Herrn Kozelj vorgenommen. | SWO-135884 | No | El Met module (902-824R2339N001) including 5 dummy electrodes on loan built-in. AC 5+ solutions for scanned. calibrations and QC's. QC Level 3 was measured 1.2mmhg higher after installation. Amendment to AQURE by Mr Kozelj. | El Met module (902-824R2339N001) including 5 dummy electrodes on loan built-in. AC 5+ solutions for scanned. calibrations and QC's. QC Level 3 was measured 1.2mmhg higher after installation. Amendment to AQURE by Mr Kozelj. | Assisted Replacement | Electrode | Electrode | pCO2 | ABL8X
AQT90 FLEX |  | B | SWO-135706 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | フィルター清掃、P5中古と交換 | B | SWO-142988 | No | Filter cleaning, replacing with P5 used |  | Defect | HW | Aspiration Module | Pump - P5
ABL80 FLEX BASIC | Calibration pass. Qc pass. machine working pass.Routine visit. | Calibration pass. Qc pass. machine working pass.Routine visit. | SWO-136177 | No | Calibration pass. Qc pass. machine working pass. Routine visit. | Calibration pass. Qc pass. machine working pass. Routine visit.
ABL90 FLEX | Probenweg mit Spritze gespült, Einlassdichtung gewechselt, Installierung LP (gleiches), Kal-Fehler Glu und Lac, Wechsel SC vom Kunden, Kal, QC Inlet Gasket with Holder geliefert für Reparaturzwecke | Probenweg mit Spritze gespült, Einlassdichtung gewechselt, Installierung LP (gleiches), Kal-Fehler Glu und Lac, Wechsel SC vom Kunden, Kal, QC Inlet Gasket with Holder geliefert für Reparaturzwecke | SWO-137226 | No | Sample path with syringe rinsed, inlet seal changed, installation LP (same), Kal-Fault Glu and Lac, Change SC from customer, Kal, QC Inlet Gasket with Holder supplied for repair purposes | Sample path with syringe rinsed, inlet seal changed, installation LP (same), Kal-Fault Glu and Lac, Change SC from customer, Kal, QC Inlet Gasket with Holder supplied for repair purposes | Assisted Replacement | SC | SC | FT
ABL90 FLEX | Kein Material vor Ort, Praxis ab Mittag geschlossen, Reparatur am Montag, den 10.10.2016 mit Material aus der GS10 | SP90 ersetzt, SC90 ersetzt, Startphase abgewartet, Kal, tHb-Kal & QK ok | SWO-137296 | No | No material on site, practice closed from noon, repair on Monday, 10.10.2016 with material from the GS10 | SP90 replaced, SC90 replaced, start phase awaited, Kal, tHb-Kal & QK ok | Revisit required | SP | SP | FT
ABL90 FLEX | Arrived on site and attended analyser to run a service tip to allow the analyser to be connected via a comport for the data connection. The analyser currently is set up as standard to allow only an external barcode reader to be connected and to rectify this, an alteration to the Ini file must be made. Connected keyboard and sub stick and took a copy backup of the current Ini file. Made the alteration to the current file on the analyser to allow the connection and then took a copy of this also. Committed the changes to the c drive and rebooted the analyser to allow the changes to take place. Once analyser had been rebooted, tests were done to check that the connection would now work. No results being seen as there was no send button on the screen to enable transmission. Further investigation showed that the com port that the analyser was set to was com port 2 when it should have been com port 1. Altered this and then sent a result through. Analysis showed that the results were being sent from the analyser but not being posted in the millennium system. Analyser sending data as it should. Customer needs to work out the posting of the information on their system. | Arrived on site and attended analyser to run a service tip to allow the analyser to be connected via a comport for the data connection. The analyser currently is set up as standard to allow only an external barcode reader to be connected and to rectify this, an alteration to the Ini file must be made. Connected keyboard and sub stick and took a copy backup of the current Ini file. Made the alteration to the current file on the analyser to allow the connection and then took a copy of this also. Committed the changes to the c drive and rebooted the analyser to allow the changes to take place. Once analyser had been rebooted, tests were done to check that the connection would now work. No results being seen as there was no send button on the screen to enable transmission. Further investigation showed that the com port that the analyser was set to was com port 2 when it should have been com port 1. Altered this and then sent a result through. Analysis showed that the results were being sent from the analyser but not being posted in the millennium system. Analyser sending data as it should. Customer needs to work out the posting of the information on their system. | SWO-136162 | No | Arrived on site and attended analyser to run a service tip to allow the analyser to be connected via a comport for the data connection. The analyser currently is set up as standard to allow only an external barcode reader to be connected and to rectify this, an alteration to the Ini file must be made. Connected keyboard and sub stick and took a copy backup of the current Ini file. Made the alteration to the current file on the analyser to allow the connection and then took a copy of this also. Committed the changes to the c drive and rebooted the analyser to allow the changes to take place. Once analyser had been rebooted, tests were done to check that the connection would now work. No results being seen as there was no send button on the screen to enable transmission. Further investigation showed that the com port that the analyser was set to was com port 2 when it should have been com port 1. Altered this and then sent a result through. Analysis showed that the results were being sent from the analyser but not being posted in the millennium system. Analyser sending data as it should. Customer needs to work out the posting of the information on their system. | Arrived on site and attended analyser to run a service tip to allow the analyser to be connected via a comport for the data connection. The analyser currently is set up as standard to allow only an external barcode reader to be connected and to rectify this, an alteration to the Ini file must be made. Connected keyboard and sub stick and took a copy backup of the current Ini file. Made the alteration to the current file on the analyser to allow the connection and then took a copy of this also. Committed the changes to the c drive and rebooted the analyser to allow the changes to take place. Once analyser had been rebooted, tests were done to check that the connection would now work. No results being seen as there was no send button on the screen to enable transmission. Further investigation showed that the com port that the analyser was set to was com port 2 when it should have been com port 1. Altered this and then sent a result through. Analysis showed that the results were being sent from the analyser but not being posted in the millennium system. Analyser sending data as it should. Customer needs to work out the posting of the information on their system. | Setup | His/Lis | His/Lis | His/Lis
ABL90 | Pressed the service soft reset button, powered off on, no change. Checked connections, ok. Put CPU in adjacent analyzer, ok. Display unit needs replaced. Ordered parts. | B | D175782D1045C356 | No | Pressed the service soft reset button, powered off on, no change. Checked connections, ok. Put CPU in adjacent analyzer, ok. Display unit needs replaced. Ordered parts. |  | Revisit required | HW | DMS_PC | Display
ABL90 FLEX PLUS | Motor inlet vervangen, tevens de inlet coupling vervangen. | Motor inlet vervangen, tevens de inlet coupling vervangen. | SWO-136042 | No | Replace the inlet coupling. | Replace the inlet coupling. | Defect | HW | Auto Inlet module | Auto Inlet module | 119
AQT90 FLEX | 更换了扫描器902-924、905-953和液路传感器902-918后解决(by zhenjie) | 更换了扫描器902-924、905-953和液路传感器902-918后解决(by zhenjie) | SWO-135663 | No | Replaced scanner 902-924, 905-953 and liquid road sensor 902-918 post-fix (by zhenjie) | Replaced scanner 902-924, 905-953 and liquid road sensor 902-918 post-fix (by zhenjie) | Defect | HW | Aspiration Module | Barcode reader, Inlet
ABL800 FLEX | replace the crea membrane & consumable. calibration pass. machine working fine. | replace the crea membrane & consumable. calibration pass. machine working fine. | SWO-136178 | No | replace the crea membrane & consumable. calibration pass. machine working fine. | replace the crea membrane & consumable. calibration pass. machine working fine. | Assisted Replacement | Electrode | Membrane | Crea
ABL825 FLEX | REVISIÓN DE HISTORICOS, SUSTITUCIÓN DE LA MEMBRANA DE LACTATO. SUSTUCIÓN DE LA ENTRADA Y LIMPIEZA TUBOS DE FLUIDICA. DESPROTEINIZACIÓN, LIMPIEZA Y LUBRICACIÓN DE LA VALVULA OXI. CALIBACIONES Y CONTROLES | REVISIÓN DE HISTORICOS, SUSTITUCIÓN DE LA MEMBRANA DE LACTATO. SUSTUCIÓN DE LA ENTRADA Y LIMPIEZA TUBOS DE FLUIDICA. DESPROTEINIZACIÓN, LIMPIEZA Y LUBRICACIÓN DE LA VALVULA OXI | D1791178721A7172 | No | HISTORICAL REVIEW, SUBSTITUTE OF THE LACTATO MEMBRANE. SUBSTITUTE OF INPUT AND CLEANING FLUIDICA TUBES. DEPROTEINIZATION, CLEANING AND LUBRICATION OF THE OXI VALVE. CALIBANS AND CONTROLS | HISTORICAL REVIEW, SUBSTITUTE OF THE LACTATO MEMBRANE. SUBSTITUTE OF INPUT AND CLEANING FLUIDICA TUBES. DEPROTEINIZATION, CLEANING AND LUBRICATION OF THE OXI VALVE | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL820 FLEX | found that user stated there was a clot that had gotten caught in the system and now PH PCO2 and Po2 will not calibrate. They are getting error leak current detected. | tried cleaning all electrodes and well. will have to order boards. | D1782055FC248230 | No | found that user stated there was a clot that had gotten caught in the system and now PH PCO2 and Po2 will not calibrate. They are getting error leak current detected. | tried cleaning all electrodes and well. will have to order boards. | Revisit required | HW | Fluidic manifold | Valves | email
ABL800 BASIC | Try to fix printer assembly but still it is not working.Need to change new printer assembly | Try to fix printer assembly but still it is not working.Need to change new printer assembly | SWO-135969 | No | Try to fix printer assembly but still it is not working. Need to change new printer assembly | Try to fix printer assembly but still it is not working. Need to change new printer assembly | Revisit required | HW | Printer module | Printer Head
ABL825 FLEX | Schläuche neu verlegt, Fehler mit Testmessung überprüft, Fehler behoben. | Schläuche neu verlegt, Fehler mit Testmessung überprüft, Fehler behoben. | SWO-136171 | No | Hoses re-laid, errors checked with test measurement, bug fixed. | Hoses re-laid, errors checked with test measurement, bug fixed. | Defect | HW | Fluidic manifold | Tubes | Assembly
ABL827 FLEX |  | B | D175436F100BFD8C | No |  |  | Defect | HW | pH/BG Module | Valve membrane
ABL800 BASIC | Cleaning tubing with hypo chloride solution. Then given two point calibration. checked with QC. Instrument is working fine. | Cleaning tubing with hypo chloride solution. Then given two point calibration. checked with QC. Instrument is working fine. | SWO-136197 | No | Cleaning tubing with hypo chloride solution. Then given two point calibration. checked with QC. Instrument is working fine. | Cleaning tubing with hypo chloride solution. Then given two point calibration. checked with QC. Instrument is working fine. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX |  | B | D1757AC1103B2D81 | No |  |  | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken)
AQT90 FLEX | Membran Pumpe aus Technikerbestand erneuert, neues Ersatzteil muss bestellt werden, Gerät neu gestartet Systemüberprüfung, Systemreinigung und QCs durchgeführt | Membran Pumpe aus Technikerbestand erneuert, neues Ersatzteil muss bestellt werden, Gerät neu gestartet Systemüberprüfung, Systemreinigung und QCs durchgeführt | SWO-136120 | No | Membrane pump from technician stock renewed, new spare part must be ordered, device restarted system inspection, system cleaning and QCs carried out | Membrane pump from technician stock renewed, new spare part must be ordered, device restarted system inspection, system cleaning and QCs carried out | Defect | HW | Process Module | Membrane Pump | 168
TCM TOSCA | Sensoren fungerer som den skal og sendt retur til Allan drews på medicoteknisk afd. | B | SWO-136083 | No | The sensor works as it should and returned to Allan Drews on the medical devices AfD.
ABL80 FLEX BASIC |  | B | SWO-136234 | No
AQT90 FLEX | cleaned the process wheel and optical unit now it is working satisfactorily. | cleaned the process wheel and optical unit now it is working satisfactorily. | SWO-136204 | No | cleaned the process wheel and optical unit now it is working satisfactorily. | cleaned the process wheel and optical unit now it is working satisfactorily. | Cleaned | HW | Process Module | Process wheel
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-137208 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX |  | B | SWO-136325 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-136346 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
AQT90 FLEX |  | B | SWO-136344 | No |  |  | Defect | HW | Cover and Chassis | Inlet Door
ABL80 FLEX BASIC | 更换试剂包修复 | 更换试剂包修复 | SWO-137610 | No | Replacement reagent kit repair | Replacement reagent kit repair
ABL80 FLEX BASIC | solution pack reset. calibrate system multiple time. replace new pack . | solution pack reset. calibrate system multiple time. replace new pack . | SWO-138900 | No | solution pack reset. calibrate system multiple time. replace new pack . | solution pack reset. calibrate system multiple time. replace new pack .
ABL90 FLEX | Erneuerung Einlass- und Anschlussdichtung - Allgemeine Überprüfung und Spülung Analysator ohne Fehlerbeseitigung - Neuer Sensor demontiert - Undichtigkeit an rechter Metall-Kapillare festgestellt - Gummirest beseitigt und Neuinitialisierung Sensor - Analysator ohne Beanstandung ! | Erneuerung Einlass- und Anschlussdichtung - Allgemeine Überprüfung und Spülung Analysator ohne Fehlerbeseitigung - Neuer Sensor demontiert - Undichtigkeit an rechter Metall-Kapillare festgestellt - Gummirest beseitigt und Neuinitialisierung Sensor - Analysator ohne Beanstandung ! | SWO-137494 | No | Renewal Inlet and connection seal - General inspection and flushing analyzer without troubleshooting - New sensor dismantled - Leakage detected on right metal capillary - Rubber residue removed and reinitialization sensor - Analyzer without Complaint! | Renewal Inlet and connection seal - General inspection and flushing analyzer without troubleshooting - New sensor dismantled - Leakage detected on right metal capillary - Rubber residue removed and reinitialization sensor - Analyzer without Complaint! | Defect | HW | Inlet module | Connector gasket | Rubber
ABL90 | Unable to pump solution errors Also External QC failing Po2 too high. | I suspect an air leak Installed Pm kit...Problem solved brought system to ready after installing a new SP that was due for replacement. | D17926DD721804DC | No | Unable to pump solution errors Also External QC failing Po2 too high. | I suspect an air leak Installed Pm kit... Problem solved brought system to ready after installing a new SP that was due for replacement. | Defect | HW | PM Kit | PM Kit
ABL90 FLEX | WiFi-Verbindung gestartet - keine Authentifizierung, ABL aus- u. eingeschaltet - gleiche Meldung, IT versucht, zu erreichen - kein Personal erreichbar. Abbruch der Arbeiten u. Info an die Station | FSE-Einsatz erforderlich | SWO-137639 | No | WiFi connection started - no authentication, ABL off and turned on - same message, IT trying to reach - no staff reachable. Cancellation of work and info to the station | FSE deployment required | Setup | SW | Wifi | Wifi
ABL90 FLEX | SP90 ersetzt, SC90 ersetzt, Startphase abgewartet, Kal, tHb-Kal & QK ok | SP90 ersetzt, SC90 ersetzt, Startphase abgewartet, Kal, tHb-Kal & QK ok | SWO-137299 | No | SP90 replaced, SC90 replaced, start phase awaited, Kal, tHb-Kal & QK ok | SP90 replaced, SC90 replaced, start phase awaited, Kal, tHb-Kal & QK ok | Assisted Replacement | SP | SP | FT
ABL90 FLEX PLUS | Roma.ini-Eintrag Barcode-Parameter erweitert (22=0, 23=0) - Funktionsüberprüfung etc. | Roma.ini-Eintrag Barcode-Parameter erweitert (22=0, 23=0) - Funktionsüberprüfung etc. | SWO-137537 | No | Roma.ini entry Barcode parameters extended (22=0, 23=0) - function check etc. | Roma.ini entry Barcode parameters extended (22=0, 23=0) - function check etc. | Setup | SW | Roma.ini | Barcode setup
ABL90 FLEX PLUS | Spring type shaft coupling not attached to motor shaft correctly. Tightened grub screws on coupling, now ok. | Spring type shaft coupling not attached to motor shaft correctly. Tightened grub screws on coupling, now ok. | SWO-137521 | No | Spring type shaft coupling not attached to motor shaft correctly. Tightened grub screws on coupling, now ok. | Spring type shaft coupling not attached to motor shaft correctly. Tightened grub screws on coupling, now ok. | Adjustment | HW | Auto Inlet module | Inlet drive gear | Service tip….
ABL90 FLEX |  | B | SWO-137367 | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX |  | B | SWO-137373 | Yes |  |  | Re classified RMED | Same day upSWO
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-137474 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | K | 378
ABL90 FLEX | 更换924-816一个正常。---jack.chu | 更换924-816一个正常。---jack.chu | SWO-137320 | No | Replace the 924-816 one normal. ---jack.chu | Replace the 924-816 one normal. ---jack.chu | Defect | HW | Inlet module | Inlet Gasket
ABL90 | SUSTITUCIÓN DE GASKET Y SONDA REVISIÓN PREVENTIVA SEGÚN PROCEDIMIENTO SUSTITUCIÓN DE SC PASADO QC | SUSTITUCIÓN DE GASKET Y SONDA REVISIÓN PREVENTIVA SEGÚN PROCEDIMIENTO SUSTITUCIÓN DE SC | D1792E007229316C | No | GASKET SUBSTITUTE AND SONDA PRE-EMPTIVE REVIEW AS SC PASSED QC SUBSTITUTE | GASKET AND SONDA SUBSTITUTE PRE-EMPTIVE REVIEW AS SC SUBSTITUTE | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-137942 | No | Repair after replacing a new test card | Repair after replacing a new test card | Assisted Replacement | SC | SC | Glu
AQT90 FLEX |  | B | SWO-137360 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
AQT90 FLEX | 更换903-364，905-908后修复。---jack.chu | 更换903-364，905-908后修复。---jack.chu | SWO-136242 | No | Replace 903-364, 905-908 after repair. ---jack.chu | Replace 903-364, 905-908 after repair. ---jack.chu | Defect | HW | Process Module | Optical Unit | 131
ABL800 FLEX | Reloaded Application software | Reloaded Application software | SWO-142740 | No | Reloaded Application software | Reloaded Application software | Defect | SW | Roma.ini | Roma.ini
ABL827 FLEX |  | B | D17962017231E828 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Na. Elektrode und Messkammer geprüft und gereinigt, Funktionskontrolle ok | B | SWO-137523 | No | well. Electrode and measuring chamber tested and cleaned, function control ok |  | Cleaned | Electrode | Electrode | Na+
ABL835 FLEX | Im Serviceprogramm LS und Pumpenkalibrierung durchgeführt, kein Erfolg. Alle LS überprüft und Inlet Liquid Sensor ersetzt. LS justage und Pumpenkalibrierung durchgeführt, alles i.o. Kal., QC und System i.o. Funktionskontrolle i.o. | Im Serviceprogramm LS und Pumpenkalibrierung durchgeführt, kein Erfolg. Alle LS überprüft und Inlet Liquid Sensor ersetzt. LS justage und Pumpenkalibrierung durchgeführt, alles i.o. Kal., QC und System i.o. Funktionskontrolle i.o. | SWO-137483 | No | Performed in the service program LS and pump calibration, no success. All LS checked and Inlet Liquid Sensor replaced. LS adjustment and pump calibration performed, all i.o. Cal., QC and system i.o. function control i.o. | Performed in the service program LS and pump calibration, no success. All LS checked and Inlet Liquid Sensor replaced. LS adjustment and pump calibration performed, all i.o. Cal., QC and system i.o. function control i.o. | Defect | HW | Inlet module | LS
ABL825 FLEX | Pumpenschlauch wieder korrekt angebracht, Kal & QC durchgeführt, Ursache mit Anwender besprochen. | Pumpenschlauch wieder korrekt angebracht, Kal & QC durchgeführt, Ursache mit Anwender besprochen. | SWO-137631 | No | Pump hose reattached correctly, Kal & QC performed, cause discussed with user. | Pump hose reattached correctly, Kal & QC performed, cause discussed with user. | Adjustment | HW | pH/BG Module | Tubes | Black
AQT90 FLEX | Perform System Cleaning, Cleaned Internal Dust and Reference Cup. | Perform System Cleaning, Cleaned Internal Dust and Reference Cup. | SWO-140549 | No | Perform System Cleaning, Cleaned Internal Dust and Reference Cup. | Perform System Cleaning, Cleaned Internal Dust and Reference Cup. | Cleaned | HW | Process Module | Cup wash unit | Why not service tip
AQT90 FLEX | div. Sensoren gereinigt, Lichtschranke der Optischen Einheit ern., Gerät neu gestartet, QC´s durchgef. | Frau F. wünscht einen Vorortservice. | SWO-137484 | No | div. Sensors cleaned, light barrier of the optical unit ern., device restarted, QC's through. | Ms. F. wants a suburban service. | Revisit required | Revisit required | Revisit required | Revisit required | Lars Email
ABL80 FLEX BASIC | Analyzer teruggeplaatst bij klant , na opstart een calibratie en controle uitgevoerd : OK Tevens een test uitgevoerd ivm de koppeling met Glims, de resultaten komen wel in Aqure maar niet in Glims, dus dit moet nog even nagecheckt worden op de Aqure server. | Analyzer teruggeplaatst bij klant , na opstart een calibratie en controle uitgevoerd : OK Tevens een test uitgevoerd ivm de koppeling met Glims, de resultaten komen wel in Aqure maar niet in Glims, dus dit moet nog even nagecheckt worden op de Aqure server. | SWO-138118 | No | Analyzer retrieved from customer, after start-up a calibration and control performed: OK also performed a test due to the coupling with Glims, the results are in Aqure but not in Glims, so this has to be checked on the Aqure server. | Analyzer retrieved from customer, after start-up a calibration and control performed: OK also performed a test due to the coupling with Glims, the results are in Aqure but not in Glims, so this has to be checked on the Aqure server.
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-138320 | No | Repair after replacing test card | Repair after replacing test card
ABL80 FLEX CO-OX | Schlauchsystem erneuert, Fließwege überprüft, i.O., sO2 und Glu Fehler, tHb kalibriert, i.O.,Gerät einkalibriert | Schlauchsystem erneuert, Fließwege überprüft, i.O., sO2 und Glu Fehler, tHb kalibriert, i.O.,Gerät einkalibriert | SWO-137720 | No | Hose system renewed, flow paths checked, i.O., sO2 and Glu error, tHb calibrated, i.O.,device calibrated | Hose system renewed, flow paths checked, i.O., sO2 and Glu error, tHb calibrated, i.O.,device calibrated
ABL90 FLEX | Lösungspack geliefert und ersetzt da Chipdaten nicht gelesen werden konnten Kalibration, QC Kontrolle. Ein Sensor sowie ein lösungspack zusätzlich geliefert. | Lösungspack geliefert und ersetzt da Chipdaten nicht gelesen werden konnten Kalibration, QC Kontrolle. Ein Sensor sowie ein lösungspack zusätzlich geliefert. | SWO-137816 | Yes | Solution pack delivered and replaced since chip data could not be read calibration, QC control. A sensor as well as a solution pack are also supplied. | Solution pack delivered and replaced since chip data could not be read calibration, QC control. A sensor as well as a solution pack are also supplied. | Re classified RMED | Logistical upSWO
AQT90 FLEX | Verificação dos registos de dados. Realização de 5 limpezas, verificação da unidade de lavagem, verificação do circuito de liquido da unidade de lavagem da agulha até bomba 4. Substituição da bomba 4. Verificação do correcto funcionamento com a execução de 5 amostras. | B | D17A9F41E815FF56 | No | Verification of data records. Conducting 5 cleanings, checking the washing unit, checking the liquid circuit of the needle wash unit until pump 4. Replacing the pump 4. Verification of the correct operation with the execution of 5 samples. |  | Cleaned | HW | Aspiration Module | HCT Sensor | 1218
ABL837 FLEX | Membran skiftet hvorefter føjsomheden steg | Membran skiftet hvorefter føjsomheden steg | SWO-137678 | No | The membrane switch after which the | The membrane switch after which the | Assisted Replacement | Electrode | Membrane | Lac | 379
ABL810 FLEX | EL MET-Modul montiert (mit 5 Dummy Elektroden) - Einstellungen und Funktionsüberprüfung (QC Level 3 pCO2 zielwertig) - Analysator transportfähig gemacht und in Auto verladen ! | EL MET-Modul montiert (mit 5 Dummy Elektroden) - Einstellungen und Funktionsüberprüfung (QC Level 3 pCO2 zielwertig) - Analysator transportfähig gemacht und in Auto verladen ! | SWO-137861 | No | EL MET module mounted (with 5 dummy electrodes) - settings and function verification (QC Level 3 pCO2 target) - Analyzer made transportable and loaded in car ! | EL MET module mounted (with 5 dummy electrodes) - settings and function verification (QC Level 3 pCO2 target) - Analyzer made transportable and loaded in car ! | Assisted Replacement | Electrode | Electrode | pCO2 | El/met added
ABL815 FLEX | Inletmodul demontiert und gereinigt, Innenraum des Analysators gereinigt, Funktionskontrolle ok | Inletmodul demontiert und gereinigt, Innenraum des Analysators gereinigt, Funktionskontrolle ok | SWO-138047 | No | Inlet module dismantled and cleaned, interior of the analyzer cleaned, function control ok | Inlet module dismantled and cleaned, interior of the analyzer cleaned, function control ok | Defect | HW | Inlet module | Inlet cal. | Leaking
ABL825 FLEX | Flex-Q Modul erneuert und abgeglichen, Gerät kalibriert, QC´s durchgef. ausgebautes Modul 902-753R1226N002 | B | SWO-137689 | No | Flex-Q module renewed and matched, device calibrated, QC's through. extended module 902-753R1226N002 |  | Defect | HW | FlexQ Module | FlexQ Module | 882
ABL815 FLEX | Monitor der alten Generation eingebaut. Auf Kundenwunsch versucht die tastentöne Stumm zu schalten, allerdings ist dies mit dem neuen Mainboard nicht möglich. | Monitor der alten Generation eingebaut. Auf Kundenwunsch versucht die tastentöne Stumm zu schalten, allerdings ist dies mit dem neuen Mainboard nicht möglich. | SWO-137763 | No | Monitor of the old generation built-in. At the customer's request, the touch tones are tried to mute, but this is not possible with the new motherboard. | Monitor of the old generation built-in. At the customer's request, the touch tones are tried to mute, but this is not possible with the new motherboard. | Defect | HW | DMS_PC | Touch screen
ABL835 FLEX | Untere BGA Modul LS ersetzt. Gerät kalibriert und QC durchgef I.O. | Untere BGA Modul LS ersetzt. Gerät kalibriert und QC durchgef I.O. | SWO-138205 | No | Replaced lower BGA module LS. calibrated and QC through I.O. | Replaced lower BGA module LS. calibrated and QC through I.O. | Defect | HW | pH/BG Module | LS, lower
ABL815 FLEX | pH und Ref. Messkammer geöffnet und gereinigt, pH Dichtring erneuert, Ref. Elektrode hatte zwei schwarze Dichtringe durch Membranierung - einen Dichtring entfernt, LS Justage und Pumpen Kal. durchgeführt, kein Austritt von Flüssigkeiten aus pH Messkammer, System- und Funktionskontrolle in Ordnung. | pH und Ref. Messkammer geöffnet und gereinigt, pH Dichtring erneuert, Ref. Elektrode hatte zwei schwarze Dichtringe durch Membranierung - einen Dichtring entfernt, LS Justage und Pumpen Kal. durchgeführt, kein Austritt von Flüssigkeiten aus pH Messkammer, System- und Funktionskontrolle in Ordnung. | SWO-137791 | No | pH and Ref. Measuring chamber opened and cleaned, pH sealing ring renewed, ref. Electrode had two black sealing rings by membraneing - a sealing ring removed, LS adjustment and pumps Kal. no leakage of liquids from pH measuring chamber, system and function control in order. | pH and Ref. Measuring chamber opened and cleaned, pH sealing ring renewed, ref. Electrode had two black sealing rings by membraneing - a sealing ring removed, LS adjustment and pumps Kal. no leakage of liquids from pH measuring chamber, system and function control in order. | User error | Electrode | Electrode | Reference | Sealing ring
AQT90 FLEX | Fehlersuche, Wheelobjects getestet - StorageWheel relativ hoher Schlupf aber noch im Bereich, IncubatorWheel positioniert icht immer korrekt - Sensoren gereinigt - leichter Metallabrieb auf den 5 Positionssensorenfestgestellt, Fluidtransport getestet, LS in Ordnung, Hct Zelle lässt sich nicht kalibrieren - läuft über während der Kalibrierung der Zellkonstante- Software stürzt hier auch wiedeholbar ab sodass Gerät neu gestartet werden muß - ausgebaut und ersetzt, mit neuer Hct Einheit identisches Ergebniss- immer am Ende der Kal läuft Zelle über und Softwareabsturz - abgebrochen, Gerät wird zurückgeschickt | Fehlersuche, Wheelobjects getestet - StorageWheel relativ hoher Schlupf aber noch im Bereich, IncubatorWheel positioniert icht immer korrekt - Sensoren gereinigt - leichter Metallabrieb auf den 5 Positionssensorenfestgestellt, Fluidtransport getestet, LS in Ordnung, Hct Zelle lässt sich nicht kalibrieren - läuft über während der Kalibrierung der Zellkonstante- Software stürzt hier auch wiedeholbar ab sodass Gerät neu gestartet werden muß - ausgebaut und ersetzt, mit neuer Hct Einheit identisches Ergebniss- immer am Ende der Kal läuft Zelle über und Softwareabsturz - abgebrochen, Gerät wird zurückgeschickt | SWO-138022 | Yes | Troubleshooting, Wheelobjects tested - StorageWheel relatively high slip but still in the range, IncubatorWheel always positions I'm correctly - sensors cleaned - light metal abrasion detected on the 5 position sensors, fluid transport tested, LS ok, Hct Cell cannot be calibrated - runs over during the calibration of the cell constant software crashes here also how recoverable so that device must be restarted - removed and replaced, with new Hct unit identical result - always at the end of the cal runs cell over and Software crash - canceled, device is returned | Troubleshooting, Wheelobjects tested - StorageWheel relatively high slip but still in the range, IncubatorWheel always positions I'm correctly - sensors cleaned - light metal abrasion detected on the 5 position sensors, fluid transport tested, LS ok, Hct Cell cannot be calibrated - runs over during the calibration of the cell constant software crashes here also how recoverable so that device must be restarted - removed and replaced, with new Hct unit identical result - always at the end of the cal runs cell over and Software crash - canceled, device is returned | Re classified RMED
ABL90 FLEX | Mit IT die Registry-Einträge angepasst, Planmäßige Password-Änderung ausgeschaltet, Registry korrigiert und gesichert, ABL neu gestartet, Kal u. QC i.O. Daten nachgesendet | Mit IT die Registry-Einträge angepasst, Planmäßige Password-Änderung ausgeschaltet, Registry korrigiert und gesichert, ABL neu gestartet, Kal u. QC i.O. Daten nachgesendet | SWO-138131 | No | With IT the registry entries adjusted, scheduled password change switched off, registry corrected and backed up, ABL restarted, Kal & QC i.O. data sent | With IT the registry entries adjusted, scheduled password change switched off, registry corrected and backed up, ABL restarted, Kal & QC i.O. data sent | Setup | His/Lis | His/Lis | His/Lis
ABL90 | Customer claims when running regular blood gas that results are displayed under Cal Verification layout. | Deleted Reloaded Syringe mode setup. | D17BB3795E4618A6 | No | Customer claims when running regular blood gas that results are displayed under Cal Verification layout. | Deleted Reloaded Syringe mode setup. | Setup | SW | Patient layout | Patient layout
ABL90 FLEX | Fehlersuche, keine Verunreinigung zu erkennen, Scanner ersetzt, Barcode lesbar, Kal und QC | Fehlersuche, keine Verunreinigung zu erkennen, Scanner ersetzt, Barcode lesbar, Kal und QC | SWO-138185 | No | Troubleshooting, no contamination to detect, scanner replaced, barcode readable, Kal and QC | Troubleshooting, no contamination to detect, scanner replaced, barcode readable, Kal and QC | Defect | HW | DMS_PC | Barcode Reader
ABL90 FLEX | ABL90 ZNA Kommunikationseinstellungen angepasst: - IP Einstellungen überprüft - Aqure Konfiguration überprüft - LIS KIS Lommunikationseinstellungen entfernt ABL90 AWR Kommunikationseinstellungen angepasst: - IP Einstellungen überprüft - Aqure Konfiguration überprüft - LIS KIS Kommunikationseinstellungen angepasst ABL90 ITS Kommunikationseinstellungen angepasst: - IP Einstellungen überprüft - Aqure Konfiguration überprüft - fehlerhafte LIS KIS Konfiguration berichtigt und angepasst (High- und Low Level Protokoll nicht korrekt) - Patientendaten ab 7.10.16 12:00 Uhr nachgesendet - aktuelle Patientenmessung wurde übertragen | ABL90 ZNA Kommunikationseinstellungen angepasst: - IP Einstellungen überprüft - Aqure Konfiguration überprüft - LIS KIS Lommunikationseinstellungen entfernt ABL90 AWR Kommunikationseinstellungen angepasst: - IP Einstellungen überprüft - Aqure Konfiguration überprüft - LIS KIS Kommunikationseinstellungen angepasst ABL90 ITS Kommunikationseinstellungen angepasst: - IP Einstellungen überprüft - Aqure Konfiguration überprüft - fehlerhafte LIS KIS Konfiguration berichtigt und angepasst (High- und Low Level Protokoll nicht korrekt) - Patientendaten ab 7.10.16 12:00 Uhr nachgesendet - aktuelle Patientenmessung wurde übertragen | SWO-138162 | No | ABL90 ZNA Communication Settings Adjusted: - IP Settings Checked - Aqure Configuration Checked - LIS KIS Lomcommunication Settings Removed ABL90 AWR Communication Settings Adjusted: - IP Settings Checked - Aqure Configuration Checked - LIS KIS communication settings adapted ABL90 ITS communication settings adjusted: - IP settings checked - Aqure configuration checked - faulty LIS KIS configuration corrected and adjusted (high and low level protocol not correct) - Patient data sent from 7.10.16 12:00 - current patient measurement was transmitted | ABL90 ZNA Communication Settings Adjusted: - IP Settings Checked - Aqure Configuration Checked - LIS KIS Lomcommunication Settings Removed ABL90 AWR Communication Settings Adjusted: - IP Settings Checked - Aqure Configuration Checked - LIS KIS communication settings adapted ABL90 ITS communication settings adjusted: - IP settings checked - Aqure configuration checked - faulty LIS KIS configuration corrected and adjusted (high and low level protocol not correct) - Patient data sent from 7.10.16 12:00 - current patient measurement was transmitted | Setup | His/Lis | His/Lis | His/Lis
ABL90 FLEX PLUS | Probleme treten ausschließlich bei einem Baby auf, Kal und tHb Kal überprüft i.O., QC überprüft, Serviceprogramme durchgeführt, oHne Beanstandung, QC | Probleme treten ausschließlich bei einem Baby auf, Kal und tHb Kal überprüft i.O., QC überprüft, Serviceprogramme durchgeführt, oHne Beanstandung, QC | SWO-138199 | No | Problems occur exclusively with a baby, Kal and tHb Kal checked i.O., QC checked, service programs carried out, OHne complaint, QC | Problems occur exclusively with a baby, Kal and tHb Kal checked i.O., QC checked, service programs carried out, OHne complaint, QC | No fault found | HW | Oxi module | Spectrum mismatch
ABL90 FLEX PLUS |  | B | D17B994C5E5FAEAD | No |  |  | Defect | HW | Auto Inlet module | Motor assembly
ABL90 FLEX | 更换了910-285和910-324后修复---jack.chu | 更换了910-285和910-324后修复---jack.chu | SWO-137683 | No | Repair edited after replacing 910-285 and 910-324---jack.chu | Repair edited after replacing 910-285 and 910-324---jack.chu | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX | 更换了910-285和910-324后修复---jack.chu | 更换了910-285和910-324后修复---jack.chu | SWO-137701 | Yes | Repair edited after replacing 910-285 and 910-324---jack.chu | Repair edited after replacing 910-285 and 910-324---jack.chu | Re classified RMED | Double entry
AQT90 FLEX |  | B | SWO-137972 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364后修复---jack.chu | 更换903-364后修复---jack.chu | SWO-137685 | No | Repair ---jack.chu after replacing 903-364 | Repair ---jack.chu after replacing 903-364 | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | Felszívótű kiegyenesítése, Fecskendő szennyes elszívócső tisztítás, AC modul ellenőrzése, Szivárgás tesztek, Kalibrálás, QC mérés | B | SWO-138074 | No | Suction needle straightening, syringe with dirty exhaust nozzle cleaning, AC module check, leak tests, calibration, QC measurement |  | Defect | HW | Inlet module | Probe bend
ABL815 FLEX | 更换902-832，902-821各一个正常。---jack.chu | B | SWO-138302 | No | Replace 902-832, 902-821 each normal. ---jack.chu |  | Defect | HW | pH/BG Module | LS
ABL815 FLEX | Inletmodul geprüft, Austauschmodul installiert, Ca. Elektrode neu membraniert, Funktionskontrolle ok | Anwenderin startet das Gerät neu; stellt Ästimierung ein; beobachtet Schlupfmeldung und meldet sich ggf. erneut | SWO-139431 | No | Inlet module tested, replacement module installed, approx. electrode re-membraneed, function control ok | user restarts the device; ceases aesthetics; observes slip message and, if necessary, reports again | Defect | HW | Inlet module | Leaking | Preheater
ABL825 FLEX | Analyzer lockup has happened three times in 60 days. Had Lauren send picture of screen. Picture shows the windows background screen and no program running. | Followed procedure for installing windows and analyzer software. Reinstalled all system software. | D17B676C5E572044 | No | Analyzer lockup has happened three times in 60 days. Had Lauren send picture of screen. Picture shows the windows background screen and no program running. | Followed procedure for installing windows and analyzer software. Reinstalled all system software. | Defect | SW | SW | Windows
ABL800 BASIC | Clean all the chamber,zero current calibrate in service program ,run protein removal and 2-point calibration and resolve the problem | Clean all the chamber,zero current calibrate in service program ,run protein removal and 2-point calibration and resolve the problem | SWO-138917 | No | Clean all the chamber,zero current calibrate in service program ,run protein removal and 2-point calibration and resolve the problem | Clean all the chamber,zero current calibrate in service program ,run protein removal and 2-point calibration and resolve the problem | Adjustment | HW | pH/BG Module | Zero current
ABL80 FLEX BASIC | Solution pack need replacement. Replace the solution pack with new one. After replace the solution unit is working properly. | Solution pack need replacement. Replace the solution pack with new one. After replace the solution unit is working properly. | SWO-138838 | No | Solution pack need replacement. Replace the solution pack with new one. After replace the solution unit is working properly. | Solution pack need replacement. Replace the solution pack with new one. After replace the solution unit is working properly.
ABL80 FLEX BASIC | blood clotted at waste manifold. Cleaned the waste manifold,now instrument working fine. | blood clotted at waste manifold. Cleaned the waste manifold,now instrument working fine. | SWO-138966 | No | blood clotted at waste manifold. Cleaned the waste manifold,now instrument working fine. | blood clotted at waste manifold. Cleaned the waste manifold,now instrument working fine.
ABL90 FLEX | Gerät aus- und eingeschaltet, Bildschirmschoner überprüft - i.O., SC ersetzt (Kundenbestand), Kal u. QC i.O. tHb-Kal | Gerät aus- und eingeschaltet, Bildschirmschoner überprüft - i.O., SC ersetzt (Kundenbestand), Kal u. QC i.O. tHb-Kal | SWO-138634 | No | Device turned off and on, screen saver checked - i.O., SC replaces (customer stock), Kal & QC i.O. tHb-Kal | Device turned off and on, screen saver checked - i.O., SC replaces (customer stock), Kal & QC i.O. tHb-Kal | Defect | SW | Screen saver | Screen saver
ABL90 FLEX | Frau Fiedler fordert nochmal Service an da OSM den Fehler nach Update (OSM) die fehlende Übertragung an Radiometer adressiert. Anfahrt zum Kunden nach Telefonaten. Während Anfahrt erfolgte Rückmeldung von Frau Fiedler dass sich OSM nochmals die Sache angeschaut hat und den Fehler Systemseitig behoben hat. Einsatz nicht mehr nötig. Abbruch der Anfahrt und Rückreise zum eigentlichen Service. | Frau Fiedler fordert nochmal Service an da OSM den Fehler nach Update (OSM) die fehlende Übertragung an Radiometer adressiert. Anfahrt zum Kunden nach Telefonaten. Während Anfahrt erfolgte Rückmeldung von Frau Fiedler dass sich OSM nochmals die Sache angeschaut hat und den Fehler Systemseitig behoben hat. Einsatz nicht mehr nötig. Abbruch der Anfahrt und Rückreise zum eigentlichen Service. | SWO-138548 | Yes | Ms. Fiedler again requests service because OSM addresses the error after update (OSM) the missing transmission to radiometer. Directions to the customer after phone calls. While approached, Ms. Fiedler reported that OSM had looked at the matter again and fixed the error on the system side. No longer necessary use. Cancellation of the journey and return journey to the actual service. | Ms. Fiedler again requests service because OSM addresses the error after update (OSM) the missing transmission to radiometer. Directions to the customer after phone calls. While approached, Ms. Fiedler reported that OSM had looked at the matter again and fixed the error on the system side. No longer necessary use. Cancellation of the journey and return journey to the actual service. | Re classified RMED | Phone support
ABL90 FLEX PLUS | Installed UPS post repair from supplier. | Replaced repaired UPS with a new unit. | SWO-139086 | Yes | Installed UPS post repair from supplier. | Replaced repaired UPS with a new unit. | Re classified RMED | UPS failing
ABL90 FLEX PLUS |  | B | SWO-138561 | No |  |  | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX PLUS | Ran fluid transport tests, failing pressure test vacuum value too high. Inlet probe bent and gasket cleaned. Still reading too high after straightening and cleaning. Reseated | It has been suggested that a new probe be fitted if available. If this does not fix the analyser, an service visit will be required. Local Engineer to contact site first thing on the 13th October to check on the analyser status and arrange a service call if required. Please contact one of the following in the order they appear, 01162586840 mark jones 01162586487 roger smith 01162584657 Stephen Gardener | SWO-138558 | No | Ran fluid transport tests, failing pressure test vacuum value too high. Inlet probe bent and gasket cleaned. Still reading too high after straightening and cleaning. Reseated | It has been suggested that a new probe be fitted if available. If this does not fix the analyser, an service visit will be required. Local Engineer to contact site first thing on the 13th October to check on the analyser status and arrange a service call if required. Please contact one of the following in the order they appear, 01162586840 mark jones 01162586487 roger smith 01162584657 Stephen Gardener | Defect | HW | Auto Inlet module | Probe bend
ABL835 FLEX | OXIﾓｼﾞｭｰﾙ引上げの為、代替モジュール設置・動作確認 | OXIﾓｼﾞｭｰﾙ引上げの為、代替モジュール設置・動作確認 | SWO-138985 | No | Installation and operation confirmation of alternative modules for OXI module pulling | Installation and operation confirmation of alternative modules for OXI module pulling | Defect | HW | Oxi module | Oxi module | 581
ABL825 FLEX | CPU Unit DMS910 ersetzt. Kal., QC und System i.o. Funktionskontrolle i.o. | CPU Unit DMS910 ersetzt. Kal., QC und System i.o. Funktionskontrolle i.o. | SWO-138499 | No | CPU Unit DMS910 ersetzt. Kal., QC und System i.o. Funktionskontrolle i.o. | CPU Unit DMS910 ersetzt. Kal., QC und System i.o. Funktionskontrolle i.o. | Defect | HW | DMS_PC | CPU
ABL800 BASIC | Meet with Dr. Meenakshi regarding the query of lyte parameter. Checked the roller and adjusted the supporting guide by opening the PC module. Now Printer is working fine. Note- Requested to kindly handle the printer with care and use always the 112mm size paper. | B | SWO-138450 | No | Meet with Dr. Meenakshi regarding the query of lyte parameter. Checked the roller and adjusted the supporting guide by opening the PC module. Now Printer is working fine. Note- Requested to kindly handle the printer with care and use always the 112mm size paper. |  | Adjustment | HW | DMS_PC | Printer Head
ABL825 FLEX | COMPLETE OXI TEST FAILED ON NEON LAMP POSITION | REPLACED SPECTROMETER AND COMPLETED OXI TEST WITH PASSING RESULTS. ALL OXI PARASMETERS DEACTIVATION AS PER CUSTOMER REQUEST | D17CEDCDD44E9DF6 | No | COMPLETE OXI TEST FAILED ON NEON LAMP POSITION | REPLACED SPECTROMETER AND COMPLETED OXI TEST WITH PASSING RESULTS. ALL OXI PARASMETERS DEACTIVATION AS PER CUSTOMER REQUEST | Defect | HW | Oxi module | Spectrometer
ABL827 FLEX | Geräte Archiv überprüft - unter Aktivitäten ist mehrmals an einem Tag aufgeführt, dass der ABL AU nicht mit DMS verbunden meldet und dann immer wieder einen kurzen Neustart bringt mit Temp. Meldungen in allen Modulen, ein Tausch einer gebrauchten AU Platine brachte keine Besseung, da diese wohl defekt war, Gerät wieder in den Urzustand versetzt und neu gestartet, Kunde darüber informiert, dass ein Leihgerät organisiert wird, Kunde ist mit dem Problem sehr unzufrieden, Gerät im Service überprüft, System- und Funktionskontrolle in Ordnung. | Geräte Archiv überprüft - unter Aktivitäten ist mehrmals an einem Tag aufgeführt, dass der ABL AU nicht mit DMS verbunden meldet und dann immer wieder einen kurzen Neustart bringt mit Temp. Meldungen in allen Modulen, ein Tausch einer gebrauchten AU Platine brachte keine Besseung, da diese wohl defekt war, Gerät wieder in den Urzustand versetzt und neu gestartet, Kunde darüber informiert, dass ein Leihgerät organisiert wird, Kunde ist mit dem Problem sehr unzufrieden, Gerät im Service überprüft, System- und Funktionskontrolle in Ordnung. | SWO-138680 | No | Devices archive checked - under Activities is listed several times in a day that the ABL AU does not report to DMS and then repeatedly brings a short restart with Temp. Messages in all modules, a replacement of a used AU board did not bring any seam, as this was probably defective, device returned to its original state and restarted, informs customer that a rental device is being organized, customer is very dissatisfied with the problem , device checked in service, system and function control in order. | Devices archive checked - under Activities is listed several times in a day that the ABL AU does not report to DMS and then repeatedly brings a short restart with Temp. Messages in all modules, a replacement of a used AU board did not bring any seam, as this was probably defective, device returned to its original state and restarted, informs customer that a rental device is being organized, customer is very dissatisfied with the problem , device checked in service, system and function control in order. | Revisit required | HW | Basic Unit | Analyzer control PCB
AQT90 FLEX | 代替入れ替え | B | SWO-138616 | No | Alternate replacement |  | Defect | HW | Aspiration Module | Tubes
TCM/Tosca |  | B | D17CD326D418DD42 | No
ABL80 FLEX CO-OX | Mit Frau Funken abgeklärt, welche Daten fehlen und wie diese erstellt werden können. Daten erstellt und versendet | SC und LP erneuert, da der Status nicht eindeutig war, Gerät kalibriert, QC´s durchgef. Der Kunde hat immer wieder verschiedene Probleme - anschließend muss das LP oder die SC erneuert werden. Mit dem Kunden die möglichen Fehlerursachen besprochen - LP und SC - Lagertemperatur - Das Gerät sollte nicht trocken laufen, davon kann das Gerät und das Zubehör einen Schaden nehmen Daten-Log. erstellt | SWO-139011 | No | With Mrs Funken, it is clear what data is missing and how it can be created. Data created and sent | SC and LP renewed because the status was not unique, device calibrated, QC's through. The customer always has different problems - after that the LP or sc has to be renewed. Discussed with the customer the possible causes of errors - LP and SC - storage temperature - The device should not run dry, of which the device and the accessories can take a damage data log. Created
ABL90 FLEX | Um den Kunden schnell helfen zu können, wurde aus einem Kondi Analysator die gesamte Wetsection ausgebaut und der Chipdatensatz zum Speichern und auslesen der Daten und des LP quer getauscht. Die defekte Weetsection wird in der Werkstatt auf Fehler untersucht. Und es werden dann dementsprechend die nötigen Ersatzteile bestellt. Analysator wieder hochgefahren. Und Verbrauchsartikel eingesetzt. | Um den Kunden schnell helfen zu können, wurde aus einem Kondi Analysator die gesamte Wetsection ausgebaut und der Chipdatensatz zum Speichern und auslesen der Daten und des LP quer getauscht. Die defekte Weetsection wird in der Werkstatt auf Fehler untersucht. Und es werden dann dementsprechend die nötigen Ersatzteile bestellt. Analysator wieder hochgefahren. Und Verbrauchsartikel eingesetzt. | SWO-138823 | No | In order to be able to help the customers quickly, the entire wetsection was removed from a Kondi analyzer and the chip data record for storing and reading out the data and the LP was exchanged across the board. The defective weetsection is examined in the workshop for errors. And the necessary spare parts are then ordered accordingly. Analyzer restarted. And consumables used. | In order to be able to help the customers quickly, the entire wetsection was removed from a Kondi analyzer and the chip data record for storing and reading out the data and the LP was exchanged across the board. The defective weetsection is examined in the workshop for errors. And the necessary spare parts are then ordered accordingly. Analyzer restarted. And consumables used. | Defect | HW | Wet section | Wet section | Is this SP pO2 ??
ABL90 FLEX | Replaced faulty Sensor Cassette. | Replaced faulty Sensor Cassette. | SWO-138922 | No | Replaced faulty Sensor Cassette. | Replaced faulty Sensor Cassette. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | Performed manual flush procedure, ok Enter service and perform each fluid transport calibration, all had failed but passed manually | Performed manual flush procedure, ok Enter service and perform each fluid transport calibration, all had failed but passed manually | SWO-138845 | No | Performed manual flush procedure, ok Enter service and perform each fluid transport calibration, all had failed but passed manually | Performed manual flush procedure, ok Enter service and perform each fluid transport calibration, all had failed but passed manually | User error | Sample result | Sampler handling | Sampler handling
ABL90 |  | B | D17E40554A030121 | No |  |  | Defect | HW | Oxi module | Pinch valve
ABL90 |  | NEW, REPLACEMENT STILL NOT WORKING WITH BACKUP BATTERY TO TOGGLE BETWEEN POWER SOURCE WHEN MAIN POWER CORD IS UNPLUGGED. WILL REORDER BOTH BATTERY PACK AND POWER SUPPLY AND RETURN TO REPLACE. | D17E15624A52DD85 | No |  | NEW, REPLACEMENT STILL NOT WORKING WITH BACKUP BATTERY TO TOGGLE BETWEEN POWER SOURCE WHEN MAIN POWER CORD IS UNPLUGGED. WILL REORDER BOTH BATTERY PACK AND POWER SUPPLY AND RETURN TO REPLACE. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | SC erneuert | SC erneuert | SWO-138812 | No | SC renewed | SC renewed | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS |  | B | D17E53164A2B1E13 | No |  |  | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS |  | B | D19F60BD461D4AA5 | No
ABL90 FLEX PLUS | Replaced repaired UPS with a new unit. | Replaced repaired UPS with a new unit. | SWO-139089 | Yes | Replaced repaired UPS with a new unit. | Replaced repaired UPS with a new unit. | Re classified RMED | UPS failing
ABL90 FLEX PLUS | Rinsed inlet, which was very dirty. Ran visual flow check, all looked good. Ran test again. | B | SWO-138853 | No | Rinsed inlet, which was very dirty. Ran visual flow check, all looked good. Ran test again. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | 更换备用件903-364修复机器---jack.chu | 更换备用件903-364修复机器---jack.chu | SWO-138879 | No | Replace spare parts 903-364 repair machine ---jack.chu | Replace spare parts 903-364 repair machine ---jack.chu | Defect | HW | Process Module | Optical Unit | 131
ABL800 FLEX | Cleaned electrode chambers, 2 point calibration performed successfully | Cleaned electrode chambers, 2 point calibration performed successfully | SWO-138882 | No | Cleaned electrode chambers, 2 point calibration performed successfully | Cleaned electrode chambers, 2 point calibration performed successfully | Cleaned | Fluidic pathway | pH/BG Module | pCO2
ABL825 FLEX |  | B | D17E4BF34A43439B | No |  |  | Setup | SW | Roma.ini | Barcode reader
ABL820 FLEX | found that the inlet nipple is broken causing fluid to hit the inlet flap detector. | have board on order to correct the issue. | D17EA3FE4A29CC6B | No | found that the inlet nipple is broken causing fluid to hit the inlet flap detector. | have board on order to correct the issue. | Revisit required | HW | Inlet module | Inlet Flap Detector PCB
ABL835 FLEX | Customer has concerns regarding levle 1 and 2 pco2 qc | tested operation of gas 2 vlave, tried differnt styles of ph oring and installed a different lot membranes, no concernable change. Contacted ASM to discuss issue. | D17E403F4A2ECE0D | No | Customer has concerns regarding levle 1 and 2 pco2 qc | tested operation of gas 2 vlave, tried differnt styles of ph oring and installed a different lot membranes, no concernable change. Contacted ASM to discuss issue. | Assisted Replacement | Electrode | Membrane | pCO2
ABL835 FLEX | Values were 2 sd below mean on level 1 and 2 qc | Replaced Ph oring with black sealing ring, ran cals and qc | D17E3FA44A4ECE0A | No | Values were 2 sd below mean on level 1 and 2 qc | Replaced Ph oring with black sealing ring, ran cals and qc | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL837 FLEX |  | B | D17E12D14A2E2D61 | No |  |  | Setup | His/Lis | His/Lis | Aqure
ABL827 FLEX | Tausch der Analyzer Control PCB, Gerät neu gestartet, Satelliten Software neu aufgespielt, Gerät startet Software Pflege, Roma.db gesichert, Kira deinstalliert, Software V6.16 neu installiert, Option Keys installiert, Update Type durchgeführt, mehrmals kalibriert, System- und Funktionskontrolle in Ordnung. | Tausch der Analyzer Control PCB, Gerät neu gestartet, Satelliten Software neu aufgespielt, Gerät startet Software Pflege, Roma.db gesichert, Kira deinstalliert, Software V6.16 neu installiert, Option Keys installiert, Update Type durchgeführt, mehrmals kalibriert, System- und Funktionskontrolle in Ordnung. | SWO-139030 | No | Exchange of Analyzer Control PCB, device restarted, satellite software replayed, device starts software maintenance, Roma.db backed up, Kira uninstalled, software V6.16 reinstalled, option keys installed, update type performed, multiple calibrated, system and functional control in order. | Exchange of Analyzer Control PCB, device restarted, satellite software replayed, device starts software maintenance, Roma.db backed up, Kira uninstalled, software V6.16 reinstalled, option keys installed, update type performed, multiple calibrated, system and functional control in order. | Defect | HW | Basic Unit | Analyzer control PCB
ABL825 FLEX | Schlauch neu aufgesteckt, System gereinigt. Glucose membraniert. | Schlauch neu aufgesteckt, System gereinigt. Glucose membraniert. | SWO-138798 | No | Hose replugged, system cleaned. glucose membranes. | Hose replugged, system cleaned. glucose membranes. | Defect | HW | Inlet module | Tubes
AQT90 FLEX | Backup gemaakt van data, tevens servicedump gemaakt. Data Collection Guide ingevuld. Halfjaarlijks onderhoud uitgevoerd, en checks gedaan. | Backup gemaakt van data, tevens servicedump gemaakt. Data Collection Guide ingevuld. Halfjaarlijks onderhoud uitgevoerd, en checks gedaan. | SWO-138786 | No | Backup made of data, also service dump created. Data Collection Guide. Half-yearly maintenance carried out, and checks done. | Backup made of data, also service dump created. Data Collection Guide. Half-yearly maintenance carried out, and checks done. | Cleaned | HW | Aspiration Module | Aspiration Module
AQT90 FLEX |  | B | D17E38594A2BAE45 | No |  |  | Defect | Consumables | Test Kit | Cartridge | 31
AQURE | Enable Centralised User Manager on these analayser | Enable Centralised User Manager on these analayser | SWO-138851 | Yes | Enable Centralised User Manager on these analayser | Enable Centralised User Manager on these analayser
ABL80 FLEX BASIC | Check the power supply LED, it is not working. So power supply is not working properly, it need replacement. | Check the power supply LED, it is not working. So power supply is not working properly, it need replacement. | SWO-140214 | No | Check the power supply LED, it is not working. So power supply is not working properly, it need replacement. | Check the power supply LED, it is not working. So power supply is not working properly, it need replacement.
AQT90 FLEX | 更换备用件903-364修复---jack.chu | 更换备用件903-364修复---jack.chu | SWO-138880 | No | Replace spare parts 903-364 repair ---jack.chu | Replace spare parts 903-364 repair ---jack.chu | Defect | HW | Process Module | Optical Unit | 131
ABL800 FLEX | Performed complete oximetry test & 2 point calibration & QC done values within the range | Performed complete oximetry test & 2 point calibration & QC done values within the range | SWO-142760 | No | Performed complete oximetry test & 2 point calibration & QC done values within the range | Performed complete oximetry test & 2 point calibration & QC done values within the range | Adjustment | HW | Oxi module | Oxi module | complete oxi test
ABL800 BASIC | we need to replace new oxi module. | B | SWO-139007 | No | we need to replace new oxi module. |  | Defect | HW | Oxi module | Leaking
AQT90 FLEX |  | B | SWO-138953 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL80 FLEX BASIC | Found pCO2, pO2 calibration error, perfomred the 2p calibration found same error, suggested to replace with new SC. | Found pCO2, pO2 calibration error, perfomred the 2p calibration found same error, suggested to replace with new SC. | SWO-140239 | No | Found pCO2, pO2 calibration error, perfomred the 2p calibration found same error, suggested to replace with new SC. | Found pCO2, pO2 calibration error, perfomred the 2p calibration found same error, suggested to replace with new SC.
ABL80 FLEX Basic version |  | B | D1820DF4AC5FA6A2 | No
ABL80 FLEX CO-OX | Leihgerät aufgebaut, Verbindungsdaten eingegeben, ASTM.ini Version (19) geändert, QC Daten manuell nachgesendet | B | SWO-139172 | No | Loaner built up, connection data entered, ASTM.ini version (19) changed, QC data manually sent
ABL80 FLEX CO-OX | Gerät überprüft, Hemolyser Schlauchkit erneuert, tHb Kal,o.k. | Hr. Bültel wird das LP80 erneuern, wünscht aber einen vor Ort Service und braucht Hilfe bei der Reklamation des Lösungspacks. | SWO-139171 | No | Device checked, Hemolyser hose kit renewed, tHb Kal,o.k. | Mr. Bültel will renew the LP80, but wants an on-site service and needs help with the complaint of the solution pack.
ABL90 FLEX | Uhrzeit nach R-Vorgabe eingestellt | Uhrzeit nach R-Vorgabe eingestellt | SWO-140501 | No | Time set according to R-default | Time set according to R-default | Setup | SW | Time setting | Time setting
ABL90 FLEX | Uhrzeit nach R-Vorgabe eingestellt. | Uhrzeit nach R-Vorgabe eingestellt. | SWO-140496 | No | Time set according to R specification. | Time set according to R specification. | Setup | SW | Time setting | Time setting
ABL90 FLEX | zu 1) Nochmals drei mal mit Kaffeesahne gemessen, danach tHb Kalibration in Ordnung. zu 2) LP nochmals erneuert. Kalibrationen und QC's vorgenommen. | zu 1) Nochmals drei mal mit Kaffeesahne gemessen, danach tHb Kalibration in Ordnung. zu 2) LP nochmals erneuert. Kalibrationen und QC's vorgenommen. | SWO-139365 | No | 1) Measured again three times with coffee cream, then tHb calibration fine. to 2) LP renewed again. calibrations and QC's. | 1) Measured again three times with coffee cream, then tHb calibration fine. to 2) LP renewed again. calibrations and QC's. | Assisted Replacement | SP | SP | Glu/Lac
ABL90 FLEX | Netzwerkeinstellungen überprüft, Verbindung überprüft, Datensendung mit ASTMNet überprüft-Daten werden übertragen, Info Labor u. EDV | B | SWO-139088 | No | Network settings checked, connection checked, data sent with ASTMNet verified data are transferred, Info Lab and COMPUTER |  | Setup | His/Lis | His/Lis | His/Lis
ABL90 FLEX | Power supply ersetzt, startet nicht Ohne Akkupack probiert, startet nicht Durch Leihgerät RSCH ersetzt i.o Konfiguration geladen i.o | B | SWO-139272 | No | Power supply replaces, does not start without battery pack tried, does not start by loanr RSCH replaces i.o configuration loaded i.o |  | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | Einweisung durchgeführt. Fehleranalyse- Behebung, Einlassmodul , SC , LP und Spülverfahren erklärt. | Einweisung durchgeführt. Fehleranalyse- Behebung, Einlassmodul , SC , LP und Spülverfahren erklärt. | SWO-139641 | Yes | instruction carried out. Error analysis fix, inlet module, SC , LP and rinse procedure explained. | instruction carried out. Error analysis fix, inlet module, SC , LP and rinse procedure explained. | Re classified RMED | Training
ABL90 | Viewed logs - error 1343 - solution pack and cassette had been changed. Back flush also done. | Installed a maintenance kit. | D182EF11223AC138 | No | Viewed logs - error 1343 - solution pack and cassette had been changed. Back flush also done. | Installed a maintenance kit. | Defect | HW | PM Kit | PM Kit | 1343
ABL90 FLEX PLUS | Jetzige Einstellungen Zeitserver wird genutzt. RMedDstManagerService auf gestoppt. Zeitzone: GMT +1 Automatische Umstellung Sommerzeit gesetzt. Daten auf CF-Karte gesichert. IP-Adresse auf 10.1.51.30 geändert. Korrekte Zeit wird angezeigt. | Jetzige Einstellungen Zeitserver wird genutzt. RMedDstManagerService auf gestoppt. Zeitzone: GMT +1 Automatische Umstellung Sommerzeit gesetzt. Daten auf CF-Karte gesichert. IP-Adresse auf 10.1.51.30 geändert. Korrekte Zeit wird angezeigt. | SWO-139161 | No | Current settings time server is used. RMedDstManagerService stopped. Time zone: GMT +1 Automatic changeover daylight saving time set. Data backed up to CF card. IP address changed to 10.1.51.30. Correct time is displayed. | Current settings time server is used. RMedDstManagerService stopped. Time zone: GMT +1 Automatic changeover daylight saving time set. Data backed up to CF card. IP address changed to 10.1.51.30. Correct time is displayed. | Setup | SW | Time setting | Time setting
ABL90 FLEX PLUS | SC 90 geprüft und ausgetauscht (aus Kundenbestand), Funktionskontrolle ok | SC 90 geprüft und ausgetauscht (aus Kundenbestand), Funktionskontrolle ok | SWO-139423 | No | SC 90 checked and replaced (from customer inventory), function control ok | SC 90 checked and replaced (from customer inventory), function control ok | Assisted Replacement | SC | SC | Lac
ABL90 FLEX PLUS |  | B | D18215AAAC2BD8FD | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Uhrzeit Einstellung überarbeitet, Hacken bei autom. Sommer Winterzeit gesetzt, Einstellungen gesichert und durch Neustadt überprüft, Datenbank Fehler durch Neustart beseitigt. | Uhrzeit Einstellung überarbeitet, Hacken bei autom. Sommer Winterzeit gesetzt, Einstellungen gesichert und durch Neustadt überprüft, Datenbank Fehler durch Neustart beseitigt. | SWO-139207 | No | Time setting revised, hacking at autom. Summer winter time set, settings backed up and checked by Neustadt, database error fixed by restart. | Time setting revised, hacking at autom. Summer winter time set, settings backed up and checked by Neustadt, database error fixed by restart. | Setup | SW | Time setting | Time setting
ABL90 FLEX | 更换新试剂包后解决 | 更换新试剂包后解决 | SWO-139110 | No | Resolve after replacing a new reagent kit | Resolve after replacing a new reagent kit | Assisted Replacement | SP | SP | pO2
ABL800 BASIC | Found printer head mechanical problem, replaced the new printer head our HYD stock, restarted instrument found instrument is functional. | Found printer head mechanical problem, replaced the new printer head our HYD stock, restarted instrument found instrument is functional. | SWO-139854 | No | Found printer head mechanical problem, replaced the new printer head our HYD stock, restarted instrument found instrument is functional. | Found printer head mechanical problem, replaced the new printer head our HYD stock, restarted instrument found instrument is functional. | Defect | HW | Printer module | Printer Head
ABL825 FLEX | REVISIÓN DE HISTORICOS, REAJUSTE DE LA LAMPARA HALÓGENA. TEST COMPLETO DE OXI, CAL THB CALIBRACIONES Y CONTROLES | REVISIÓN DE HISTORICOS, REAJUSTE DE LA LAMPARA HALÓGENA. TEST COMPLETO DE OXI, CAL THB | D1847FCB98249231 | No | HISTORICAL REVIEW, ADJUSTMENT OF THE HALGENAL LAMPARA. FULL TEST OF OXI, CAL THB CALIBRATIONS AND CONTROLS | HISTORICAL REVIEW, ADJUSTMENT OF THE HALGENAL LAMPARA. FULL OXY TEST, CAL THB | Adjustment | HW | Oxi module | Lamp unit | Halogen
ABL825 FLEX | AC-Modul demontiert, Reinigung, Montage, Startphase, pO2-Elektrode membraniert, korrekte Kal abgewartet, QK ok | AC-Modul demontiert, Reinigung, Montage, Startphase, pO2-Elektrode membraniert, korrekte Kal abgewartet, QK ok | SWO-139269 | No | AC module dismantled, cleaning, assembly, start-up phase, pO2 electrode membraneed, correct scale awaiting, QK ok | AC module dismantled, cleaning, assembly, start-up phase, pO2 electrode membraneed, correct scale awaiting, QK ok | Adjustment | HW | Auto Check Module | Breaker arm
ABL825 FLEX | ooked in data logs and found numerous times QC level 1 and level 2 are out high. S7735 LOT 657, reading 7.087, customer range is 7.089 - 7.101, insert range is 7.078 - to 7.118. Looked over analyzer, found missing black oring on PO2 electrode and found that blue oring on PH was on incorrectly. Had problem getting lactate to come in for calibration. Tried 5 2pt cals still not in. remembraned lactate electrode from customers stock. | Installed PO2 electrode Oring and correctly installed PH oring all was from customer stock. remembraned lactate electrode from customers stock. | D182112CAC0EB2C4 | No | ooked in data logs and found numerous times QC level 1 and level 2 are out high. S7735 LOT 657, reading 7.087, customer range is 7.089 - 7.101, insert range is 7.078 - to 7.118. Looked over analyzer, found missing black oring on PO2 electrode and found that blue oring on PH was on incorrectly. Had problem getting lactate to come in for calibration. Tried 5 2pt cals still not in. remembraned lactate electrode from customers stock. | Installed PO2 electrode Oring and correctly installed PH oring all was from customer stock. remembraned lactate electrode from customers stock. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL825 FLEX | looked in data logs and found numerous times QC level 1 and level 2 are out high. S7735 LOT 657, reading 7.103, customer range is 7.089 - 7.101, insert range is 7.078 - to 7.118. Looked over analyzer, found bent inlet probe, found missing black oring on PO2 electrode and found that blue oring on PH was on incorrectly. | Replaced inlet probe, PO2 electrode and correctly installed PH oring all was from customer stock. | D181F690AC606486 | No | looked in data logs and found numerous times QC level 1 and level 2 are out high. S7735 LOT 657, reading 7.103, customer range is 7.089 - 7.101, insert range is 7.078 - to 7.118. Looked over analyzer, found bent inlet probe, found missing black oring on PO2 electrode and found that blue oring on PH was on incorrectly. | Replaced inlet probe, PO2 electrode and correctly installed PH oring all was from customer stock. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL810 FLEX | Analysator aufgestellt - Inbetriebnahme - Module nicht aktiv - eventuell muss sich Analysator erst klimatisieren - Ersatzgerät wieder installiert ! | AU-Board aus Ersatzgerät zu Testzwecken montiert (keine Fehlerbeseitigung) - Software neu installiert (keine Fehlerbeseitigung) - Power Distibution PCB bestellt ! | SWO-139200 | No | Analyzer set up - commissioning - modules not active - maybe ananalyzer has to have to air-conditioning first - replacement device reinstalled ! | AU-Board assembled from replacement device for testing purposes (no troubleshooting) - Software reinstalled (no troubleshooting) - Power Distibution PCB ordered ! | Assisted Replacement | Electrode | Electrode | pCO2 | El/met added
ABL810 FLEX | AU-Board aus Ersatzgerät zu Testzwecken montiert (keine Fehlerbeseitigung) - Software neu installiert (keine Fehlerbeseitigung) - Power Distibution PCB bestellt ! | AU-Board aus Ersatzgerät zu Testzwecken montiert (keine Fehlerbeseitigung) - Software neu installiert (keine Fehlerbeseitigung) - Power Distibution PCB bestellt ! | SWO-139244 | Yes | AU-Board assembled from replacement device for testing purposes (no troubleshooting) - Software reinstalled (no troubleshooting) - Power Distibution PCB ordered ! | AU-Board assembled from replacement device for testing purposes (no troubleshooting) - Software reinstalled (no troubleshooting) - Power Distibution PCB ordered ! | Re classified RMED | Double entry
ABL810 FLEX | Analysator heruntergefahren - Montage EL MET Modul - Einstellungen und Funktionsüberprüfung etc. | Analysator heruntergefahren - Montage EL MET Modul - Einstellungen und Funktionsüberprüfung etc. | SWO-139202 | No | Analyzer shut down - Assembly EL MET module - Settings and function check etc. | Analyzer shut down - Assembly EL MET module - Settings and function check etc. | Assisted Replacement | Electrode | Electrode | pCO2 | El/met added
ABL815 FLEX | Oxi Modul demontiert, Leihmodul installiert, Ersatzteile bestellt, Funktionskontrolle ok | Oxi Modul demontiert, Leihmodul installiert, Ersatzteile bestellt, Funktionskontrolle ok | SWO-139422 | No | Oxi module dismantled, rental module installed, spare parts ordered, function control ok | Oxi module dismantled, rental module installed, spare parts ordered, function control ok | Revisit required | HW | Oxi module | Oxi module
AQT90 FLEX | Cleaned Reference Cup, Internal Dust and Perform System Cleaning. | Cleaned Reference Cup, Internal Dust and Perform System Cleaning. | SWO-140550 | No | Cleaned Reference Cup, Internal Dust and Perform System Cleaning. | Cleaned Reference Cup, Internal Dust and Perform System Cleaning. | Cleaned | HW | Process Module | Cup wash unit | Why not service tip
AQT90 FLEX | 引上げ | 分解組立て実施 インキュベーションホイールと底面に干渉箇所あり 使用しているネジが長かったため 正規のネジに交換してOK 測定テスト、CALテストOK | SWO-141639 | No | Pulling | Disassembly and assembly execution There is an interference point in the incubation wheel and the bottom, and because the screw used was long, i replaced it with a regular screw and changed it to a regular screw, and ok measurement test, CAL test OK | Lars email |  |  | Not Cat.
AQT90 FLEX | Gerät überprüft, Membranpumpe erneuert, Reinigung durchgeführt | Gerät überprüft, Membranpumpe erneuert, Reinigung durchgeführt | SWO-139467 | No | Device checked, diaphragm pump renewed, cleaning carried out | Device checked, diaphragm pump renewed, cleaning carried out | Defect | HW | Process Module | Membrane Pump | 168
ABL80 FLEX CO-OX | Fehlersuche Überprüfung im Service Programm abgesprungen Schlauch wieder korrekt angebracht Installation Neues LP Testmessung i.O. | Fehlersuche Überprüfung im Service Programm abgesprungen Schlauch wieder korrekt angebracht Installation Neues LP Testmessung i.O. | SWO-139443 | No | Troubleshooting check in the service program bounced off hose again correctly attached Installation New LP test measurement in the following. | Troubleshooting check in the service program bounced off hose again correctly attached Installation New LP test measurement in the following.
ABL80 FLEX BASIC | 更换了2个试剂包后修复 | 更换了2个试剂包后修复 | SWO-139741 | No | Repaired after replacing 2 reagent kits | Repaired after replacing 2 reagent kits
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-139466 | No | Replacement test card repair | Replacement test card repair
ABL80 CO-OX | Fluid Blockage flow error. | Removed SP and powered down analyzer. Powered on and re-installed SP | D183C595223F0FC2 | No | Fluid Blockage flow error. | Removed SP and powered down analyzer. Powered on and re-installed SP
ABL90 FLEX | Flüssigkeitswege überprüft und gespült, Nadel erneuert. Flüssigkeitstransport Test ok. Sensorkasetten Temperatur kann nicht erreicht werden. Sensorkasette gewechselt und Reklamiert Alle Tests, Kalibrierungen und QC wurden bestanden. | Kd. benötigt einen Servicetechniker zur Beseitigung der Verstopfung. | SWO-139999 | No | Liquid pathways checked and rinsed, needle renewed. Liquid transport test ok. Sensor-caste temperature cannot be reached. Sensor caette changed and complaint All tests, calibrations and QC were passed. | Kd. needs a service technician to remove the blockage. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 |  | B | D183598C2246DDC0 | No |  |  | Defect | HW | Power Supply | Power Supply
ABL90 | Unit has been much better since blue tooth devise was replced but still disconnects intermittently. Other wirless units do not. | Checked network settings. Checked allow echo request in ICMP properties, | D183220022530319 | No | Unit has been much better since blue tooth devise was replced but still disconnects intermittently. Other wirless units do not. | Checked network settings. Checked allow echo request in ICMP properties, | Setup | His/Lis | His/Lis | His/Lis
ABL90 FLEX PLUS | Radiometer Mitarbeiterin S. Bertram vor Ort hat Analysator gespült und mehrere Messungen vollzogen - Analysator ohne Beanstandung | Radiometer Mitarbeiterin S. Bertram vor Ort hat Analysator gespült und mehrere Messungen vollzogen - Analysator ohne Beanstandung | SWO-139506 | No | Radiometer employee S. Bertram on site has flushed analyzer and made several measurements - analyzer without complaint | Radiometer employee S. Bertram on site has flushed analyzer and made several measurements - analyzer without complaint | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | Machine taken to workshop. Stripped and tried to unblock the hemolyzer but was unable to unblock the hemolyzer. | Machine taken to workshop. Stripped and tried to unblock the hemolyzer but was unable to unblock the hemolyzer. | SWO-140857 | No | Machine taken to workshop. Stripped and tried to unblock the hemolyzer but was unable to unblock the hemolyzer. | Machine taken to workshop. Stripped and tried to unblock the hemolyzer but was unable to unblock the hemolyzer. | Defect | HW | Oxi module | Hemolyzer | Clot
AQT90 FLEX |  | B | SWO-139316 | No |  |  | Defect | HW | Cover and Chassis | Inlet Door | 4548
ABL835 FLEX | PH BG CHAMBER'IN ÇATLAK OLDUĞU TESPİT EDİLDİ.CHAMBER GEÇİCİ YAPIŞTIRILARAK SORUN GİDERİLDİ. | PH BG CHAMBER'IN ÇATLAK OLDUĞU TESPİT EDİLDİ.CHAMBER GEÇİCİ YAPIŞTIRILARAK SORUN GİDERİLDİ. | SWO-140903 | Yes | PH BG CHAMBER HAS BEEN IDENTIFIED AS CRACKED.CHAMBER HAS BEEN PASSED THE PROBLEM BY MAKING A TRANSITION. | PH BG CHAMBER HAS BEEN IDENTIFIED AS CRACKED.CHAMBER HAS BEEN PASSED THE PROBLEM BY MAKING A TRANSITION.
ABL815 FLEX | Oxi LS Fehler beseitigt | Fehleranalyse mit Anwender Termin für 18.10.16 vereinbart | SWO-139332 | No | Oxi LS error fixed | Error analysis agreed with user appointment for 18.10.16 | Defect | HW | Oxi module | Hemolyzer | Noise
ABL825 FLEX | Comprovação dos registos de cal. e QC, medidas e erros do sistema. Substituição do detector de posição da sonda. Inicialização. Calibração e verificação com QC. | B | D18352D622018CB2 | No | Verification of the records of lime and QC, measures and errors of the system. Replacing the probe position detector. Startup. Calibration and verification with QC. |  | Defect | HW | Inlet module | Friction detector unit
ABL810 FLEX | Nach Fehlersuche Auto Check aus Reparatur-Ersatzgerät montiert ( + Gegencheck) - Analysator ohne Beanstandung - Bemerkung: Neuteil-Bestellung erfolgt | Nach Fehlersuche Auto Check aus Reparatur-Ersatzgerät montiert ( + Gegencheck) - Analysator ohne Beanstandung - Bemerkung: Neuteil-Bestellung erfolgt | SWO-139502 | No | After troubleshooting Auto Check assembled from repair replacement device ( + countercheck) - Analyzer without complaint - Remark: New parts order is made | After troubleshooting Auto Check assembled from repair replacement device ( + countercheck) - Analyzer without complaint - Remark: New parts order is made | Revisit required | HW | Auto Check Module | AutoCheck Module
ABL810 FLEX | Erneuerung Power Distribution PCB brachte keine Fehlerbeseitigung - Fehlersuche - Auto Check Modul defekt - Autocheck aus Ersatzgerät montiert - Analysator ohne Beanstandung - Siehe Case Autocheck Modul defekt | Erneuerung Power Distribution PCB brachte keine Fehlerbeseitigung - Fehlersuche - Auto Check Modul defekt - Autocheck aus Ersatzgerät montiert - Analysator ohne Beanstandung - Siehe Case Autocheck Modul defekt | SWO-139501 | No | Renewal Power Distribution PCB brought no trouble-making - Troubleshooting - Auto Check module defective - Autocheck assembled from replacement device - Analyzer without complaint - See Case Autocheck module defective | Renewal Power Distribution PCB brought no trouble-making - Troubleshooting - Auto Check module defective - Autocheck assembled from replacement device - Analyzer without complaint - See Case Autocheck module defective | Assisted Replacement | Electrode | Electrode | pCO2 | El/met added
ABL810 FLEX | Montage EL MET Modul - Update Type - Funktionsüberprüfung etc. ohne Beanstandung! | Montage EL MET Modul - Update Type - Funktionsüberprüfung etc. ohne Beanstandung! | SWO-139622 | No | Mounting EL MET Module - Update Type - Function check etc. without complaint! | Mounting EL MET Module - Update Type - Function check etc. without complaint! | Assisted Replacement | Electrode | Electrode | pCO2 | El/met added
ABL810 FLEX | Montage EL MET-Modul - Update Type (815) - Funktionsüberprüfung etc. ohne Beanstandung | Montage EL MET-Modul - Update Type (815) - Funktionsüberprüfung etc. ohne Beanstandung | SWO-139625 | No | Installation EL MET-Module - Update Type (815) - Functional check etc. without complaint | Installation EL MET-Module - Update Type (815) - Functional check etc. without complaint | Assisted Replacement | Electrode | Electrode | pCO2 | El/met added
ABL810 FLEX | Montage EL MET-Modul - Update Type (815) - Funktionsüberprüfung etc. ohne Beanstandung ! | Montage EL MET-Modul - Update Type (815) - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-139623 | No | Installation EL MET-Module - Update Type (815) - Functional check etc. without complaint ! | Installation EL MET-Module - Update Type (815) - Functional check etc. without complaint ! | Assisted Replacement | Electrode | Electrode | pCO2 | El/met added
ABL827 FLEX |  | B | D186C76184600AD8 | No |  |  | Defect | HW | Inlet module | Probe bend
ABL827 FLEX | Gerät geprüft, Messkammer BGa Modul und pCo2 membran ersetzt. Alle Ventile geprüft I.O. Mehrere Kalibrationen und QC level 2 und 3 durchgef. I.O. | Gerät geprüft, Messkammer BGa Modul und pCo2 membran ersetzt. Alle Ventile geprüft I.O. Mehrere Kalibrationen und QC level 2 und 3 durchgef. I.O. | SWO-139437 | No | device tested, measuring chamber BGa module and pCo2 membrane replaced. All valves tested I.O. Several calibrations and QC level 2 and 3. Ok. | device tested, measuring chamber BGa module and pCo2 membrane replaced. All valves tested I.O. Several calibrations and QC level 2 and 3. Ok. | Defect | HW | pH/BG Module | Measuring chamber | 8X7
ABL800 BASIC | cleaned the gasket by hypo solution and also flushed the tubing, now working satisfactorily. | cleaned the gasket by hypo solution and also flushed the tubing, now working satisfactorily. | SWO-140548 | No | cleaned the gasket by hypo solution and also flushed the tubing, now working satisfactorily. | cleaned the gasket by hypo solution and also flushed the tubing, now working satisfactorily. | Cleaned | Fluidic pathway | Inlet module | Inlet Gasket
ABL800 BASIC | changed Lac membrane. Then given calibration. Unit is working fine, | changed Lac membrane. Then given calibration. Unit is working fine, | SWO-140539 | No | changed Lac membrane. Then given calibration. Unit is working fine, | changed Lac membrane. Then given calibration. Unit is working fine, | Assisted Replacement | Electrode | Membrane | Lac
ABL800 BASIC | Cleaned electrode chambers, performed 2 point calibration, samples measurement performed | Cleaned electrode chambers, performed 2 point calibration, samples measurement performed | SWO-140522 | No | Cleaned electrode chambers, performed 2 point calibration, samples measurement performed | Cleaned electrode chambers, performed 2 point calibration, samples measurement performed | Cleaned | HW | pH/BG Module | Measuring chamber
AQT90 FLEX | Check in Service program. Storage wheel not OK. Service for Storage wheel and periphery. Again problem reported. Check in Service program, Cartridge lock not OK. Service for cartridge lock. Analyzer OK and ready for sample. Run sample but after 2 min with a noise analyzer become Non operational. Check for service program and how was sample performed and found, Cartridge lock PCB may be problem. Need to replace Cartridge lock PCB. | Check in Service program. Storage wheel not OK. Service for Storage wheel and periphery. Again problem reported. Check in Service program, Cartridge lock not OK. Service for cartridge lock. Analyzer OK and ready for sample. Run sample but after 2 min with a noise analyzer become Non operational. Check for service program and how was sample performed and found, Cartridge lock PCB may be problem. Need to replace Cartridge lock PCB. | SWO-139508 | No | Check in Service program. Storage wheel not OK. Service for Storage wheel and periphery. Again problem reported. Check in Service program, Cartridge lock not OK. Service for cartridge lock. Analyzer OK and ready for sample. Run sample but after 2 min with a noise analyzer become Non operational. Check for service program and how was sample performed and found, Cartridge lock PCB may be problem. Need to replace Cartridge lock PCB. | Check in Service program. Storage wheel not OK. Service for Storage wheel and periphery. Again problem reported. Check in Service program, Cartridge lock not OK. Service for cartridge lock. Analyzer OK and ready for sample. Run sample but after 2 min with a noise analyzer become Non operational. Check for service program and how was sample performed and found, Cartridge lock PCB may be problem. Need to replace Cartridge lock PCB. | Revisit required | HW | Storage Module | PCB, Cartridge Lock
AQT90 FLEX | Safe Positions ausgeführt. Danach war AQT wieder Messbereit. Folgende Maßnahmen vorbeugend durchgeführt. Keine Restlichen Cups Processwheel vorhanden. Photo Interuptor des Processwheel gereinigt. Schrauben auf Processwheel nachgezogen. Processwheel im Service überprüft (mehrmals Positionen geändert) - ok. Systemüberprüfung - ok. Zwei Systemreinigungen durchgeführt - ok. Jeweils eine LQC pro Parameter durchgeführt. Beim Auswerfen der Cups aus dem Processwheel trat Fehler erneut auf. Pufferkassette manuell aus AQT entfernt. Cup entfernt. Safe Positions ok. Neue Pufferkassette eingesetzt. Nochmals eine Reinigung und eine LQC pro Parametern durchgeführt. Systemüberprüfung - ok. | Safe Positions ausgeführt. Danach war AQT wieder Messbereit. Folgende Maßnahmen vorbeugend durchgeführt. Keine Restlichen Cups Processwheel vorhanden. Photo Interuptor des Processwheel gereinigt. Schrauben auf Processwheel nachgezogen. Processwheel im Service überprüft (mehrmals Positionen geändert) - ok. Systemüberprüfung - ok. Zwei Systemreinigungen durchgeführt - ok. Jeweils eine LQC pro Parameter durchgeführt. Beim Auswerfen der Cups aus dem Processwheel trat Fehler erneut auf. Pufferkassette manuell aus AQT entfernt. Cup entfernt. Safe Positions ok. Neue Pufferkassette eingesetzt. Nochmals eine Reinigung und eine LQC pro Parametern durchgeführt. Systemüberprüfung - ok. | SWO-139580 | No | Safe Positions. After that, AQT was ready to measure again. The following preventive measures are carried out. No remaining Cups Processwheel available. Photo Interuptor of the Processwheel cleaned. screws on Processwheel. Processwheel checked in service (changed positions several times) - ok. System check - ok. Two system cleanings performed - ok. One LQC per parameter. Errors occurred again when ejecting the cups from the processwheel. Manually removed from AQT. Cup removed. Safe Positions ok. New buffer cartridge inserted. Again a cleaning and one LQC per parameter performed. System check - ok. | Safe Positions. After that, AQT was ready to measure again. The following preventive measures are carried out. No remaining Cups Processwheel available. Photo Interuptor of the Processwheel cleaned. screws on Processwheel. Processwheel checked in service (changed positions several times) - ok. System check - ok. Two system cleanings performed - ok. One LQC per parameter. Errors occurred again when ejecting the cups from the processwheel. Manually removed from AQT. Cup removed. Safe Positions ok. New buffer cartridge inserted. Again a cleaning and one LQC per parameter performed. System check - ok. | Cleaned | HW | Process Module | Sensor
ABL90 | Customer had replaced solution packL evel 3 reading approx. 80 where the range is 3 -28. Ran level 3 of IVK-R5 having approx the same range. Level 3 read 120 and gave an inhomogeneous sample error. Possible silicone contamination. | Flushed sensor cassette and ran several blood samples | D184A7DC98547641 | No | Customer had replaced solution packL evel 3 reading approx. 80 where the range is 3 -28. Ran level 3 of IVK-R5 having approx the same range. Level 3 read 120 and gave an inhomogeneous sample error. Possible silicone contamination. | Flushed sensor cassette and ran several blood samples | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL800 BASIC | Replace the auto check scanner and check ,Now it's working properly.Replace the scheduled membranes and check all tubings and check for any blockage and the complete preventive maintenance. | Replace the auto check scanner and check ,Now it's working properly.Replace the scheduled membranes and check all tubings and check for any blockage and the complete preventive maintenance. | SWO-140592 | No | Replace the auto check scanner and check ,Now it's working properly. Replace the scheduled membranes and check all tubings and check for any blockage and the complete preventive maintenance. | Replace the auto check scanner and check ,Now it's working properly. Replace the scheduled membranes and check all tubings and check for any blockage and the complete preventive maintenance. | Defect | HW | Auto Check Module | Barcode Reader
ABL810 FLEX | Neue pCO2 Membran eingesetzt, Messkammern u. Kontakte gereinigt. Cl- Dummy Dichtung erneuert. Kal. i.O. QCs in 3 Lvl. fast mittelwertig. | Neue pCO2 Membran eingesetzt, Messkammern u. Kontakte gereinigt. Cl- Dummy Dichtung erneuert. Kal. i.O. QCs in 3 Lvl. fast mittelwertig. | SWO-140016 | No | New pCO2 membrane inserted, measuring chambers and contacts cleaned. Cl- Dummy seal renewed. Kal. i.O. QCs in 3 Lvl. almost average. | New pCO2 membrane inserted, measuring chambers and contacts cleaned. Cl- Dummy seal renewed. Kal. i.O. QCs in 3 Lvl. almost average. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Einstellungen erklärt und dem Bediener gezeigt. | Einstellungen erklärt und dem Bediener gezeigt. | SWO-139742 | No | settings and shown to the operator. | settings and shown to the operator. | Setup | SW | Demo key exp | Demo key exp | 6
ABL800 BASIC | Replace the membrane,clean all chambers,check all pump tubings,clean gasket and done complete service.Install the permanent option key. | Replace the membrane,clean all chambers,check all pump tubings,clean gasket and done complete service.Install the permanent option key. | SWO-140593 | No | Replace the membrane,clean all chambers,check all pump tubings,clean gasket and done complete service. Install the permanent option key. | Replace the membrane,clean all chambers,check all pump tubings,clean gasket and done complete service. Install the permanent option key. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL837 FLEX | Error 316 Neon Number too low and error 258 neon lamp intensity too low.The Oxi Complete test failed neon position and neon lamp intensity. The OXI Pcb did not resolved the issue. After the OXI PCB was installed the Spectrometer A D converted failed but neon number too low error disappeared. Ran further diagnostics. Checked the Optic fibre cable. | Replaced Oxi PCB. Replaced Lamp unit. Lamp unit adjustment Done. Installed Option Keys and Windows 7 License label. Replaced Cal1, Cal2, Cleaning and Crea Membranes. | D1847A6598134C21 | No | Error 316 Neon Number too low and error 258 neon lamp intensity too low. The Oxi Complete test failed neon position and neon lamp intensity. The OXI Pcb did not resolved the issue. After the OXI PCB was installed the Spectrometer A D converted failed but neon number too low error disappeared. Ran further diagnostics. Checked the Optic fibre cable. | Replaced Oxi PCB. Replaced Lamp unit. Lamp unit adjustment Done. Installed Option Keys and Windows 7 License label. Replaced Cal1, Cal2, Cleaning and Crea Membranes. | Defect | HW | Oxi module | Oxi PCB
ABL815 FLEX | Membranierung mit D7707, Lot R3276, Kal u. QC i.O. | Membranierung: 22.09.2016, Status 306,7 pA mM, Kal-Status 27.09.2016: 192,5 pA mM Membranierung Lac am 27.09.2016, Kontrolle am 11.10.2016 | SWO-139921 | No | Membranierung mit D7707, Lot R3276, Kal u. QC i.O. | Membraneing: 22.09.2016, status 306.7 pA mM, Kal status 27.09.2016: 192.5 pA mM membraneing Lac on 27.09.2016, control on 11.10.2016 | Assisted Replacement | Electrode | Membrane | Lac
ABL800 BASIC | Clean all chamber perform thb calibration nopw machie is running ok. | Clean all chamber perform thb calibration nopw machie is running ok. | SWO-140994 | No | Clean all chamber perform thb calibration nopw machie is running ok. | Clean all chamber perform thb calibration nopw machie is running ok. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | 更换903-329一个正常。(by andrew) | 更换903-329一个正常。(by andrew) | SWO-139445 | No | Replace the 903-329 one normal. (by andrew) | Replace the 903-329 one normal. (by andrew) | Defect | HW | Aspiration Module | Needle Wash Unit | 813
ABL1 | Tried to Reinstalled Sp But Problem Remains, Replaced SP with New one. | Tried to Reinstalled Sp But Problem Remains, Replaced SP with New one. | SWO-140558 | No | Tried to Reinstalled Sp But Problem Remains, Replaced SP with New one. | Tried to Reinstalled Sp But Problem Remains, Replaced SP with New one.
ABL80 FLEX BASIC | センサーボード腐食のため交換 | B | SWO-142978 | No | Replacement for sensor board corrosion
ABL90 FLEX PLUS | Tubings en connector vervangen. | Tubings en connector vervangen. | SWO-140065 | No | Replace the tubings and connector. | Replace the tubings and connector. | Defect | HW | PM Kit | PM Kit | 163
ABL90 FLEX PLUS | Ersatzgerät aufgestellt092R0080N0012 um festzustellen ob die Abweichungen identisch sind, Konfiguration und Lis Kis anbindung Synlab Mode Besuch im Labor um die Übertragung und Werte zu überprüfen | Ersatzgerät aufgestellt092R0080N0012 um festzustellen ob die Abweichungen identisch sind, Konfiguration und Lis Kis anbindung Synlab Mode Besuch im Labor um die Übertragung und Werte zu überprüfen | SWO-140151 | No | Replacement device set up092R0080N0012 to determine if the deviations are identical, configuration and Li's connection Synlab Mode visit to the laboratory to check the transmission and values | Replacement device set up092R0080N0012 to determine if the deviations are identical, configuration and Li's connection Synlab Mode visit to the laboratory to check the transmission and values | Unknown | Analyzer replaced | Analyzer replaced | Analyzer replaced | Bili
ABL90 FLEX PLUS | Analysator mit Zeitserver synchronisiert, Analysator zur Überprüfung neugestartet, Uhrzeit bleibt korrekt. Kein Fehler vorhanden. | Analysator mit Zeitserver synchronisiert, Analysator zur Überprüfung neugestartet, Uhrzeit bleibt korrekt. Kein Fehler vorhanden. | SWO-140165 | No | Analyzer synchronized with time server, analyzer restarted for verification, time remains correct. No error exists. | Analyzer synchronized with time server, analyzer restarted for verification, time remains correct. No error exists. | Setup | SW | Time setting | Time setting
ABL90 FLEX PLUS | Hacken bei automatischer Sommer Winterzeit Umstellung gesetzt, Analysator mit Zeitserver synchronisiert, Einstellungen mit Commit Befehl gesichert. Analysator zur Überprüfung neugestartet, Uhrzeit bleibt korrekt. | Hacken bei automatischer Sommer Winterzeit Umstellung gesetzt, Analysator mit Zeitserver synchronisiert, Einstellungen mit Commit Befehl gesichert. Analysator zur Überprüfung neugestartet, Uhrzeit bleibt korrekt. | SWO-140144 | No | Hacking at automatic summer winter time changeover set, analyzer synchronized with time server, settings backed up with commit command. Analyzer restarted for review, time remains correct. | Hacking at automatic summer winter time changeover set, analyzer synchronized with time server, settings backed up with commit command. Analyzer restarted for review, time remains correct. | Setup | SW | Time setting | Time setting
ABL80 FLEX BASIC | Check for zeroing and found burn on sensor board. Sensor board replaced and zeroing calibration done. Sample run for same patient by three times and found all vales OK and acceptable. Analyser serviced and calibrated properly. PM service done. Working fine. Handed over in a fine working condition. | Check for zeroing and found burn on sensor board. Sensor board replaced and zeroing calibration done. Sample run for same patient by three times and found all vales OK and acceptable. Analyser serviced and calibrated properly. PM service done. Working fine. Handed over in a fine working condition. | SWO-140186 | No | Check for zeroing and found burn on sensor board. Sensor board replaced and zeroing calibration done. Sample run for same patient by three times and found all vales OK and acceptable. Analyser serviced and calibrated properly. PM service done. Working fine. Handed over in a fine working condition. | Check for zeroing and found burn on sensor board. Sensor board replaced and zeroing calibration done. Sample run for same patient by three times and found all vales OK and acceptable. Analyser serviced and calibrated properly. PM service done. Working fine. Handed over in a fine working condition.
AQT90 FLEX |  | B | SWO-139968 | No |  |  | Defect | HW | Aspiration Module | Barcode reader, Inlet
AQT90 FLEX |  | B | SWO-139965 | No |  |  | Defect | HW | Cover and Chassis | Storage Door | 4726
ABL820 FLEX | Investigated with customer who states the pump tubing was changed and the pump error intermittently returns during the PH BG calibration | Performed a pump calibration in service (calibration 71%). Replaced the pump tube and ran pump calibration (calibration 85%) . Cleaned and inspected the pump rotors. | D185A2A60E23B927 | No | Investigated with customer who states the pump tubing was changed and the pump error intermittently returns during the PH BG calibration | Performed a pump calibration in service (calibration 71%). Replaced the pump tube and ran pump calibration (calibration 85%) . Cleaned and inspected the pump rotors. | Defect | HW | pH/BG Module | Pump tubes
ABL800 BASIC | Change new printer head now machine is ok. | Change new printer head now machine is ok. | SWO-140982 | No | Change new printer head now machine is ok. | Change new printer head now machine is ok. | Defect | HW | Printer module | Printer Head
ABL835 FLEX |  | B | D18C5D3E5C170AD6 | No |  |  | Defect | HW | Fluidic manifold | Valve | 963
ABL800 BASIC | Cleaned the inlet gasket & checked equipment is funciotnal | Cleaned the inlet gasket & checked equipment is funciotnal | SWO-142786 | No | Cleaned the inlet gasket & checked equipment is funciotnal | Cleaned the inlet gasket & checked equipment is funciotnal | Cleaned | Fluidic pathway | Inlet module | Inlet Gasket
ABL825 FLEX | Found that the keyboard onboard and external will not work. | Reinstalled the windows OS with a clean 6.16 software. Copied setup from NICU analyzer and loaded it onto ICU analyzer. Configured the network settings. Went to replace analyzer CPU and received a DOA part. FSE to return with good CPU and replace. | D185BAA40E0CF8F2 | No | Found that the keyboard onboard and external will not work. | Reinstalled the windows OS with a clean 6.16 software. Copied setup from NICU analyzer and loaded it onto ICU analyzer. Configured the network settings. Went to replace analyzer CPU and received a DOA part. FSE to return with good CPU and replace. | Revisit required | HW | DMS_PC | CPU
ABL80 CO-OX | Parece problema de válvula. Se pide material. Llegada del mismo. Sustittución de la válvula del hemolizador (el gancho está bastante duro) y realización kit de tubos como medida preventiva. Puesta en marcha, tests ST, ciclo de calidad. realización primera muestra OK, se dejan pasar 10 minutos y la siguiente también OK. | cambio válvula oxi y kit de tubos | D1870E9C8434B7BD | No | Looks like a valve problem. Material is requested. Arrival of the same. Replacement of the hemolizer valve (the hook is quite hard) and kit of tubes as a preventive measure. Commissioning, ST tests, quality cycle. realization first shows OK, 10 minutes are allowed and the next also OK. | change oxy valve and tube kit
ABL80 FLEX BASIC | mit einer Spritze die Lösungen manuell gezogen, zunächst Unterdruck, hat dann funktioniert, mehrere Kal und QC | mit einer Spritze die Lösungen manuell gezogen, zunächst Unterdruck, hat dann funktioniert, mehrere Kal und QC | SWO-140483 | No | with a syringe the solutions manually pulled, first underpressure, then has worked, several Kal and QC | with a syringe the solutions manually pulled, first underpressure, then has worked, several Kal and QC
ABL80 FLEX BASIC | reinstalled sensor cassette. Then given manual 2 point calibration. Unit is working fine. | reinstalled sensor cassette. Then given manual 2 point calibration. Unit is working fine. | SWO-140541 | No | reinstalled sensor cassette. Then given manual 2 point calibration. Unit is working fine. | reinstalled sensor cassette. Then given manual 2 point calibration. Unit is working fine.
ABL80 FLEX CO-OX | ＳＣ交換 | B | SWO-143041 | No | SC Switch
ABL80 FLEX BASIC |  | B | SWO-140289 | No
ABL80 OSM | checked errors on patient samples ran coox service test, hemolyzer hardware failure cavitation absorb failing -.35 range is 5-30 checked tubing for leaks replaced internal pump head still has low cavitation values replaced hemolyzer and cavitation absorb 16.5 | replaced the internal pump and hemolyzer | D1872FF48454ABFD | No | checked errors on patient samples ran coox service test, hemolyzer hardware failure cavitation absorb failing -.35 range is 5-30 checked tubing for leaks replaced internal pump head still has low cavitation values replaced hemolyzer and cavitation absorb 16.5 | replaced the internal pump and hemolyzer
ABL80 FLEX Basic version |  | B | D1873C30840CD9BA | Yes |  |  | Re classified RMED | Installation
ABL90 FLEX | Felszívótű tömítés cseréje,Tűtartó kimosása, Átmosás, Szervíztesztek, Inicializálás, Kalibráció, QC mérés | B | SWO-140341 | No | Replacing the suction needle, cleaning the needle holder, relaying, service tests, initialization, calibration, QC measurement |  | Defect | HW | Inlet module | Inlet Gasket | 163
ABL90 FLEX | Uhrzeit nach R-Vorgabe eingestellt. | Uhrzeit nach R-Vorgabe eingestellt. | SWO-140491 | No | Time set according to R specification. | Time set according to R specification. | Setup | SW | Time setting | Time setting
ABL90 FLEX | Uhrzeit nach R-Vorgabe eingestellt. | Uhrzeit nach R-Vorgabe eingestellt. | SWO-140504 | No | Time set according to R specification. | Time set according to R specification. | Setup | SW | Time setting | Time setting
ABL90 FLEX | Uhrzeit nach R-Vorgabe eingestellt. | Uhrzeit nach R-Vorgabe eingestellt. | SWO-140508 | No | Time set according to R specification. | Time set according to R specification. | Setup | SW | Time setting | Time setting
ABL90 FLEX | Uhrzeit nach R-Vorgabe eingestellt. | Uhrzeit nach R-Vorgabe eingestellt. | SWO-140505 | No | Time set according to R specification. | Time set according to R specification. | Setup | SW | Time setting | Time setting
ABL90 FLEX | Uhrzeit nach R-Vorgabe eingestellt. | Uhrzeit nach R-Vorgabe eingestellt. | SWO-140502 | No | Time set according to R specification. | Time set according to R specification. | Setup | SW | Time setting | Time setting
ABL90 FLEX | Uhrzeit nach R-Vorgabe eingestellt. | Uhrzeit nach R-Vorgabe eingestellt. | SWO-140500 | No | Time set according to R specification. | Time set according to R specification. | Setup | SW | Time setting | Time setting
ABL90 |  | FOLLOWED CURRENT PROCEDURE TO REPLACE POWER SUPPLY. DISCONNECT MAIN POWER CORD FROM SYSTEM REAR AND RECONNECT. POWER SUPPLY NOW ALTERNATES BETWEEN BACKUP BATTERY AND POWER CORD DULY. | D186EA66845BDDA8 | No |  | FOLLOWED CURRENT PROCEDURE TO REPLACE POWER SUPPLY. DISCONNECT MAIN POWER CORD FROM SYSTEM REAR AND RECONNECT. POWER SUPPLY NOW ALTERNATES BETWEEN BACKUP BATTERY AND POWER CORD DULY. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX PLUS |  | B | D1874D12842CBD48 | Yes |  |  | Re classified RMED | Installation
AQT90 FLEX |  | B | SWO-140294 | No |  |  | Defect | HW | Cover and Chassis | Inlet Door | 7548
ABL800 FLEX | Replace the consumable & Crea membrane. Calibration sensitivity low in Lactate.Replace the Lactate membrane also & calibrate system multiple time. Calibration pass. | Replace the consumable & Crea membrane. Calibration sensitivity low in Lactate.Replace the Lactate membrane also & calibrate system multiple time. Calibration pass. | SWO-140516 | No | Replace the consumable & Crea membrane. Calibration sensitivity low in Lactate.Replace the Lactate membrane also & calibrate system multiple time. Calibration pass. | Replace the consumable & Crea membrane. Calibration sensitivity low in Lactate.Replace the Lactate membrane also & calibrate system multiple time. Calibration pass. | Assisted Replacement | Electrode | Membrane | Crea
ABL815 FLEX | fecskendő bemeneti fül cseréje, QC mérés | B | SWO-140384 | No | Syringe Input tab replacement, QC measurement |  | Defect | HW | Inlet module | Probe bend
ABL825 FLEX |  | B | D18739A4840ED9B6 | Yes |  |  | Re classified RMED
ABL825 FLEX | Lactate dropped from 10.2 to 9.4 following a membrane change. | Checked measuring chamber. Replaced connectors. Checked oring on pH elecrode. | D18707F284239F81 | No | Lactate dropped from 10.2 to 9.4 following a membrane change. | Checked measuring chamber. Replaced connectors. Checked oring on pH elecrode. | Defect | HW | El/Met Module | Electrode connector | Lac
ABL827 FLEX | VISUAL INSPECTION SHOWS MEMBRANE COMPLETELY DRY WITH NO FLUID.UNABLE TO GET SENSITIVITY ON PO2 OVER 3.8 % | FIRST TRIED REPLACING THE MEMBRANE AND SENSITIVITY TOO LOW.INSTALLED A USED PO2 ELECTRODE AND SENSITIVITY NOW IN AT 9.8%.CUSTOMER NEED TO CALL CUSTOMER SERVICE AND ORDER A NEW PO2 ELECTRODE AS THEIR ELECTRODE LOT UP03 IS UNDER WARRANTY.RAN MULTIPLE CALIBRATIONS.ALL PASSING. | D186E0B6840E72D1 | No | VISUAL INSPECTION SHOWS MEMBRANE COMPLETELY DRY WITH NO FLUID. UNABLE TO GET SENSITIVITY ON PO2 OVER 3.8 % | FIRST TRIED REPLACING THE MEMBRANE AND SENSITIVITY TOO LOW. INSTALLED A USED PO2 ELECTRODE AND SENSITIVITY NOW IN AT 9.8%. CUSTOMER NEED TO CALL CUSTOMER SERVICE AND ORDER A NEW PO2 ELECTRODE AS THEIR ELECTRODE LOT UP03 IS UNDER WARRANTY. RAN MULTIPLE CALIBRATIONS. ALL PASSING. | Revisit required | Electrode | Electrode | pO2
ABL815 FLEX | Layout bei betroffenen Messungen zu Testzwecken auf Standartgestellt - pO2 unkorrigiert wird berichtet aber hier wird auch ein Fehler angezeigt, Layout angepasst zusätzlich mit Standart Blutgaswerten | Layout bei betroffenen Messungen zu Testzwecken auf Standartgestellt - pO2 unkorrigiert wird berichtet aber hier wird auch ein Fehler angezeigt, Layout angepasst zusätzlich mit Standart Blutgaswerten | SWO-140463 | No | Layout for affected measurements for test purposes set on standard - pO2 is reported uncorrected but here also an error is displayed, layout adapted additionally with standard blood gas values | Layout for affected measurements for test purposes set on standard - pO2 is reported uncorrected but here also an error is displayed, layout adapted additionally with standard blood gas values | Setup | Patient Layout | Patient Layout | Patient Layout
ABL815 FLEX | Fehlersuche: Ventile überprüft, Ventilmembrane des BG Gasventils vom Ventilstempel ausgestanzt => Ventilmembrane ersetzt. Neustart, LS-Justage, Pumpenkal, Kal und AC5+ Messung => OK | Erneute Fehlersuche: Fließwege komplett überprüft, alle ok. Dichtigkeit der Nasssektion überprüft, ok. Lufteinlass überprüft, verstopft => frei geblasen mittels Druckluft. Erhebliche Staubbelastung des AC Moduls frei geblasen BCR Spiegel gereinigt, Einlassdichtung ersetzt. Im Servicemenü Inlet-Kal, LS-Justage, Pumpenkal durchgeführt, alle ok. Neustart, Kal und Messungen => alle ok. | SWO-140466 | No | Troubleshooting: Valves checked, valve membrane of the BG gas valve punched out by the valve stamp => Valve membrane replaced. Restart, LS-Adjustment, Pump Scale, Kal and AC5+ Measurement => OK | Re-troubleshooting: Flow paths completely checked, all ok. Tightness of the wet section checked, ok. Air inlet checked, clogged => freely blown by compressed air. Substantial dust load of the AC module freely blown BCR mirror cleaned, inlet seal replaced. In the service menu Inlet-Kal, LS-Adjustment, Pumpscale performed, all ok. Restart, Cal and Measurements => all ok. | Defect | HW | Fluidic manifold | Valve | 379
ABL815 FLEX | Erneute Fehlersuche: Fließwege komplett überprüft, alle ok. Dichtigkeit der Nasssektion überprüft, ok. Lufteinlass überprüft, verstopft => frei geblasen mittels Druckluft. Erhebliche Staubbelastung des AC Moduls frei geblasen BCR Spiegel gereinigt, Einlassdichtung ersetzt. Im Servicemenü Inlet-Kal, LS-Justage, Pumpenkal durchgeführt, alle ok. Neustart, Kal und Messungen => alle ok. | Erneute Fehlersuche: Fließwege komplett überprüft, alle ok. Dichtigkeit der Nasssektion überprüft, ok. Lufteinlass überprüft, verstopft => frei geblasen mittels Druckluft. Erhebliche Staubbelastung des AC Moduls frei geblasen BCR Spiegel gereinigt, Einlassdichtung ersetzt. Im Servicemenü Inlet-Kal, LS-Justage, Pumpenkal durchgeführt, alle ok. Neustart, Kal und Messungen => alle ok. | SWO-140509 | Yes | Re-troubleshooting: Flow paths completely checked, all ok. Tightness of the wet section checked, ok. Air inlet checked, clogged => freely blown by compressed air. Substantial dust load of the AC module freely blown BCR mirror cleaned, inlet seal replaced. In the service menu Inlet-Kal, LS-Adjustment, Pumpscale performed, all ok. Restart, Cal and Measurements => all ok. | Re-troubleshooting: Flow paths completely checked, all ok. Tightness of the wet section checked, ok. Air inlet checked, clogged => freely blown by compressed air. Substantial dust load of the AC module freely blown BCR mirror cleaned, inlet seal replaced. In the service menu Inlet-Kal, LS-Adjustment, Pumpscale performed, all ok. Restart, Cal and Measurements => all ok. | Re classified RMED | Double entry
ABL825 FLEX | The analyzer has been having numerous and various issues since it was installed in August. | Replaced the computer unit and during the first calibration, the oximetry failed. Performed power down up and ran the startup calibration again. The Hb now works as it should. Reset the time and the time date format. Changed settings to allow remote support operation. | D186E5E6840A22DA | No | The analyzer has been having numerous and various issues since it was installed in August. | Replaced the computer unit and during the first calibration, the oximetry failed. Performed power down up and ran the startup calibration again. The Hb now works as it should. Reset the time and the time date format. Changed settings to allow remote support operation. | Defect | HW | DMS_PC | CPU
AQT90 FLEX | Cartridge lock PCB replaced. analyser checked, Cartidge lock issue resolved but now next problem 9098- Inlet Barcod problem 7 come. Checked for connecting cable and barcoad- found both faulty. Need to replace Inlet barcoade and cable first. | Cartridge lock PCB replaced. analyser checked, Cartidge lock issue resolved but now next problem 9098- Inlet Barcod problem 7 come. Checked for connecting cable and barcoad- found both faulty. Need to replace Inlet barcoade and cable first. | SWO-140560 | No | Cartridge lock PCB replaced. analyser checked, Cartidge lock issue resolved but now next problem 9098- Inlet Barcod problem 7 come. Checked for connecting cable and barcoad- found both faulty. Need to replace Inlet barcoade and cable first. | Cartridge lock PCB replaced. analyser checked, Cartidge lock issue resolved but now next problem 9098- Inlet Barcod problem 7 come. Checked for connecting cable and barcoad- found both faulty. Need to replace Inlet barcoade and cable first. | Defect | HW | Storage Module | PCB, Cartridge Lock | 998
AQT90 FLEX | Perform System Cleaning Procedure, Cleaned Internal Dust from Reference Cup, restart the Analyser | Perform System Cleaning Procedure, Cleaned Internal Dust from Reference Cup, restart the Analyser | SWO-140559 | No | Perform System Cleaning Procedure, Cleaned Internal Dust from Reference Cup, restart the Analyser | Perform System Cleaning Procedure, Cleaned Internal Dust from Reference Cup, restart the Analyser | Cleaned | HW | Aspiration Module | Needle Wash Unit
ABL80 FLEX BASIC | this is user mistake no problem with instrument.in standby mode it is showing expire dates of solution pack and sensor cassette. explained the expirations of solution pack and sensor cassette. | this is user mistake no problem with instrument.in standby mode it is showing expire dates of solution pack and sensor cassette. explained the expirations of solution pack and sensor cassette. | SWO-140565 | No | this is user mistake no problem with instrument.in standby mode it is showing expire dates of solution pack and sensor cassette. explained the expirations of solution pack and sensor cassette. | this is user mistake no problem with instrument.in standby mode it is showing expire dates of solution pack and sensor cassette. explained the expirations of solution pack and sensor cassette.
ABL80 FLEX BASIC | 更换了样机上的配件903-334和902-972后修复(by andrew) | 更换了样机上的配件903-334和902-972后修复(by andrew) | SWO-140582 | No | Replacement of the fittings 903-334 and 902-972 on the prototype (by-andrew) | Replacement of the fittings 903-334 and 902-972 on the prototype (by-andrew)
ABL825 FLEX | QC Ampullen eingesetzt, leere Spüllösung ersetzt, Kal und QC ok | TEL.:Stat.Schwester - QC Ampullen seien aufgefüllt - Barcode lesbar, Laserspiegel frei pusten lassen - keine Besserung | SWO-140631 | Yes | QC ampoules inserted, empty rinsing solution replaced, Kal and QC ok | TEL.:Stat.Sister - QC ampoules are filled - barcode readable, let laser mirror blow free - no improvement | Re classified RMED | Logistical upSWO
ABL80 FLEX BASIC |  | B | SWO-140891 | No
ABL80 FLEX BASIC | CPU、CFC交換実施。 | CPU、CFC交換実施。 | SWO-140970 | No | Cpu and CFC exchange s | Cpu and CFC exchange s
ABL80 FLEX BASIC | Installed the (945-777) 200 60 cassette. Demonstrate the PICO50 sampler. it is advisable to use the radiometer PICO50 sampler. | B | SWO-140965 | No | Installed the (945-777) 200 60 cassette. Demonstrate the PICO50 sampler. it is advisable to use the radiometer PICO50 sampler.
ABL80 CO-OX | Susan called in with a possible fluid blockage. She had replaced the solution pack and sensor cassette. I checked to see if pump 3 and 4 were pumping fluid. None of the pumps were pumping fluid. I checked and found clots in the fluid line so I put on a maintenance kit. All pumps are pumps are now pumping fluid. | Installed maintenance kit. Flushed each pump. | D18A8CDBE636842E | No | Susan called in with a possible fluid blockage. She had replaced the solution pack and sensor cassette. I checked to see if pump 3 and 4 were pumping fluid. None of the pumps were pumping fluid. I checked and found clots in the fluid line so I put on a maintenance kit. All pumps are pumps are now pumping fluid. | Installed maintenance kit. Flushed each pump.
ABL90 FLEX | Neues LP (xp 03) eingesetzt | Neues LP (xp 03) eingesetzt | SWO-140805 | Yes | New LP (xp 03) inserted | New LP (xp 03) inserted | Re classified RMED | Logistical upSWO
ABL90 | Looked in data logs and Glucose and Lactate shows -- for calibration and QC. Suspect bad Sensor interface PCB | Replaced all parts from annual maintenance kit and replaced sensor interface PCB and replace fuse on Wet section PCB. Also upgraded software to MR8. | D18ACB51E62D4062 | No | Looked in data logs and Glucose and Lactate shows -- for calibration and QC. Suspect bad Sensor interface PCB | Replaced all parts from annual maintenance kit and replaced sensor interface PCB and replace fuse on Wet section PCB. Also upgraded software to MR8. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX | Korrekturen eingegeben , Zeitserver überprüft, Gerät neu gestartet. Analysator-Zeit stabil. Datensicherung durchgeführt. Servicekoffern an Med. Technik ausgehändigt, siehe Lieferschein. | Korrekturen eingegeben , Zeitserver überprüft, Gerät neu gestartet. Analysator-Zeit stabil. Datensicherung durchgeführt. Servicekoffern an Med. Technik ausgehändigt, siehe Lieferschein. | SWO-141114 | No | Corrections entered, time server checked, device restarted. Analyzer time stable. data backup. Service cases handed over to Med. Technik, see delivery note. | Corrections entered, time server checked, device restarted. Analyzer time stable. data backup. Service cases handed over to Med. Technik, see delivery note. | Setup | SW | Time setting | Time setting
ABL90 FLEX PLUS | Fehlersuche Überprüfung und Kalibrierung im Service Programm Einlaßnadel gereinigt, Dichtung erneuert,Fließwege gespült Geräteneustart SC erneuert Austausch div.Schläuche Einlaufverfahren | Fehlersuche Überprüfung und Kalibrierung im Service Programm Einlaßnadel gereinigt, Dichtung erneuert,Fließwege gespült Geräteneustart SC erneuert Austausch div.Schläuche Einlaufverfahren | SWO-140887 | No | Troubleshooting inspection and calibration in the service program inlet needle cleaned, seal renewed,flow paths flushed device restart SC renewed replacement div.hoses inlet method | Troubleshooting inspection and calibration in the service program inlet needle cleaned, seal renewed,flow paths flushed device restart SC renewed replacement div.hoses inlet method | Defect | HW | PM Kit | PM Kit
ABL90 FLEX PLUS |  | B | D18A8411E61E0AC0 | No |  |  | Assisted Replacement | SP | SP | pO2
AQT90 FLEX |  | B | SWO-140747 | No |  |  | Defect | HW | Cover and Chassis | Inlet Door
AQT90 FLEX |  | B | SWO-140827 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer
ABL825 FLEX | einer der beiden Abfallschläuche hatte sich am AC Modul verklemmt bzw. war eingeklemmt Schlauch richtig verlegt und getestet | einer der beiden Abfallschläuche hatte sich am AC Modul verklemmt bzw. war eingeklemmt Schlauch richtig verlegt und getestet | SWO-140898 | No | one of the two waste hoses had become jammed on the AC module or was clamped hose correctly laid and tested | one of the two waste hoses had become jammed on the AC module or was clamped hose correctly laid and tested | Adjustment | HW | Basic Unit | Tube aligned
ABL800 BASIC | Stripped and cleaned the entire barcode module internally and externally, re aligned barcode laser. | Stripped and cleaned the entire barcode module internally and externally, re aligned barcode laser. | SWO-141471 | No | Stripped and cleaned the entire barcode module internally and externally, re aligned barcode laser. | Stripped and cleaned the entire barcode module internally and externally, re aligned barcode laser. | Cleaned | HW | Basic Unit | Barcode reader
ABL837 FLEX |  | B | D18B115BE6214BB2 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL827 FLEX | Druckkopf und Crea. Membranen erneuert. | Druckkopf und Crea. Membranen erneuert. | SWO-141403 | No | Printhead and Crea. membranes. | Printhead and Crea. membranes. | Defect | HW | Printer module | Printer Head
ABL800 BASIC | Change Po2 and Pco2 membrane and clean all chambers,protien removal now machine is running ok. | Change Po2 and Pco2 membrane and clean all chambers,protien removal now machine is running ok. | SWO-140989 | No | Change Po2 and Pco2 membrane and clean all chambers,protien removal now machine is running ok. | Change Po2 and Pco2 membrane and clean all chambers,protien removal now machine is running ok. | Cleaned | HW | pH/BG Module | Measuring chamber
ABL825 FLEX |  | B | D18AD7D1E64A853A | No |  |  | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX |  | B | D18AD98DE63F853E | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Messkammer gereinigt, Elektrode neu membraniert, Gerät mehrfach kalibriert und mit QCs überprüfte | Messkammer gereinigt, Elektrode neu membraniert, Gerät mehrfach kalibriert und mit QCs überprüfte | SWO-141166 | No | measuring chamber cleaned, electrode re-membraneed, device calibrated several times and checked with QCs | measuring chamber cleaned, electrode re-membraneed, device calibrated several times and checked with QCs | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Fehlersuche, Lüfter abgeklemmt, weil Raum ist klimatisiert, Funktionskontrolle | Fehlersuche, Lüfter abgeklemmt, weil Raum ist klimatisiert, Funktionskontrolle | SWO-140990 | No | Troubleshooting, fan disconnected because room is air-conditioned, function control | Troubleshooting, fan disconnected because room is air-conditioned, function control | Defect | HW | Auto Check Module | FAN
ABL800 BASIC | 新品交換 | B | SWO-141186 | No | New replacement |  | Defect | HW | DMS_PC | CPU
ABL815 FLEX | Oxi Modul demontiert und die Oxi Modul PCB (alt 903-571R2000N0646 neu 903-125R8156N001) ausgetauscht, SW 6.16 neu installiert, Funktionskontrolle ok | Oxi Modul demontiert und die Oxi Modul PCB (alt 903-571R2000N0646 neu 903-125R8156N001) ausgetauscht, SW 6.16 neu installiert, Funktionskontrolle ok | SWO-141034 | No | Oxi module dismantled and replaced the Oxi module PCB (old 903-571R2000N0646 new 903-125R8156N001), SW 6.16 reinstalled, function control ok | Oxi module dismantled and replaced the Oxi module PCB (old 903-571R2000N0646 new 903-125R8156N001), SW 6.16 reinstalled, function control ok | Defect | HW | Oxi module | PCB
ABL835 FLEX |  | B | D19D83C7D024EB23 | No |  |  | Defect | HW | Oxi module | Oxi module
TCM-COMBIM 84 | Disassembled and inspected boards, connectors and battery. Ordered CF card. | B | D18AD2CFE600E621 | No | Disassembled and inspected boards, connectors and battery. Ordered CF card.
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-142857 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 FLEX | Gerät überprüft, Einlaßdichtung ausgetauscht, Lösungspack ersetzt, Sensorkassette erneuert, Kal und QC o.k. | Gerät überprüft, Einlaßdichtung ausgetauscht, Lösungspack ersetzt, Sensorkassette erneuert, Kal und QC o.k. | SWO-141650 | No | Device checked, inlet seal replaced, solution pack replaced, sensor cartridge renewed, Kal and QC ok. | Device checked, inlet seal replaced, solution pack replaced, sensor cartridge renewed, Kal and QC ok. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D18C4F355C4C0AC8 | No |  |  | Defect | HW | Oxi module | Pinch valve
ABL90 FLEX | 更换测试修复 | 更换测试修复 | SWO-141368 | No | Replacement test repair | Replacement test repair
ABL800 FLEX | clean the chamber. calibrate system multiple time. remove the air in membrane & calibrate system. multiple time. | clean the chamber. calibrate system multiple time. remove the air in membrane & calibrate system. multiple time. | SWO-142884 | No | clean the chamber. calibrate system multiple time. remove the air in membrane & calibrate system. multiple time. | clean the chamber. calibrate system multiple time. remove the air in membrane & calibrate system. multiple time. | Cleaned | Fluidic pathway | Met II | Measuring chamber
ABL835 FLEX | メンブラン交換（不良品持ち帰り後日交換） | B | SWO-141593 | No | Membrane replacement (replacement after a date after taking home defective products) |  | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL810 FLEX | pCO2 Elektrode SK-01 durch UK02 ersetzt, mehrfach kalibriert, Fehler weiter vorhanden. Beide Gas Behälter erneuert, ohne Erfolg. Elektrode zurück getauscht, pCO2 Membranen von der ITS geholt ( Lot:R0755 ) und erneuert. | pCO2 Elektrode SK-01 durch UK02 ersetzt, mehrfach kalibriert, Fehler weiter vorhanden. Beide Gas Behälter erneuert, ohne Erfolg. Elektrode zurück getauscht, pCO2 Membranen von der ITS geholt ( Lot:R0755 ) und erneuert. | SWO-141583 | No | pCO2 electrode SK-01 replaced by UK02, multiple calibrated, error still present. Both gas containers renewed, to no avail. Electrode replaced, pCO2 membranes fetched from the ITS ( Lot:R0755 ) and renewed. | pCO2 electrode SK-01 replaced by UK02, multiple calibrated, error still present. Both gas containers renewed, to no avail. Electrode replaced, pCO2 membranes fetched from the ITS ( Lot:R0755 ) and renewed. | Assisted Replacement | Electrode | Membrane | pCO2 | Rilibäk
ABL827 FLEX |  | B | D18C1A0A5C63258B | Yes |  |  | Re classified RMED | Logistical upSWO
ABL815 FLEX | Tausch des Druckerkopfes, Ausdruck wieder leserlich, ABL mehrmals kalibriert, System- und Funktionskontrolle in Ordnung. | Tausch des Druckerkopfes, Ausdruck wieder leserlich, ABL mehrmals kalibriert, System- und Funktionskontrolle in Ordnung. | SWO-141689 | No | Replacement of the printer head, printable again, ABL calibrated several times, system and function control in order. | Replacement of the printer head, printable again, ABL calibrated several times, system and function control in order. | Defect | HW | Printer module | Printer Head
AQT90 FLEX |  | B | D18C2A6D5C2F3D16 | No |  |  | Revisit required |  |  | Not Cat.
ABL80 FLEX CO-OX |  | B | SWO-141736 | Yes |  |  | Re classified RMED | Installation
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-142004 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换了信号板903-334后修复(by zhenjie) | 更换了信号板903-334后修复(by zhenjie) | SWO-141224 | No | Repair after replacing signal plate 903-334 (by zhenjie) | Repair after replacing signal plate 903-334 (by zhenjie)
ABL1 | Cleaned SC with Cleaning Solution and reinstalled, Passed QC. | Cleaned SC with Cleaning Solution and reinstalled, Passed QC. | SWO-142865 | No | Cleaned SC with Cleaning Solution and reinstalled, Passed QC. | Cleaned SC with Cleaning Solution and reinstalled, Passed QC.
ABL80 FLEX CO-OX | Gerät überprüft, Lösungspack und Sensorkassette erneuert | Gerät überprüft, Lösungspack und Sensorkassette erneuert | SWO-141864 | No | Device checked, solution pack and sensor cartridge renewed | Device checked, solution pack and sensor cartridge renewed
ABL90 FLEX | A műszer elszállításra kerül felülvizsgálat miatt, helyszinen a készülékkel megegyező típus cserekészülék üzemel | B | SWO-141701 | No | The instrument is shipped for review, the same type of replacement is in place on the |  | Unknown | Unknown | Unknown | Unknown
ABL90 FLEX | Netzteil erneuert. | Netzteil erneuert. | SWO-141817 | No | power supply is renewed. | power supply is renewed. | Defect | HW | Power Supply | Power Supply
ABL90 |  | B | D18D4441D23444E6 | Yes |  |  | Re classified RMED
ABL90 FLEX | Refitted same SP and analyser went through startup ok | Refitted same SP and analyser went through startup ok | SWO-141877 | No | Refitted same SP and analyser went through startup ok | Refitted same SP and analyser went through startup ok | Adjustment | SP | SP | SP Re install
ABL90 FLEX PLUS |  | B | D18D437DD2066750 | No
AQT90 FLEX | 更换了光学单元903-364，重新定标通过(by zhenjie) | 更换了光学单元903-364，重新定标通过(by zhenjie) | SWO-140915 | No | Replaced optical unit 903-364, re-calibration pass (by zhenjie) | Replaced optical unit 903-364, re-calibration pass (by zhenjie) | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | Luftblasen aus pCO2-Elektrode entfernt, Kal abgewartet, QK ok | Luftblasen aus pCO2-Elektrode entfernt, Kal abgewartet, QK ok | SWO-142067 | No | Air bubbles removed from pCO2 electrode, Kal waited, QK ok | Air bubbles removed from pCO2 electrode, Kal waited, QK ok | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Kontrolle der Glukoseelektrode, Austausch der verschlissenen oberen Dichtung, Glu-Membran erneuert, korrekte Kal abgewartet, QK ok | Kontrolle der Glukoseelektrode, Austausch der verschlissenen oberen Dichtung, Glu-Membran erneuert, korrekte Kal abgewartet, QK ok | SWO-142068 | No | Control of the glucose electrode, replacement of worn upper seal, glu membrane renewed, correct calf waited, QK ok | Control of the glucose electrode, replacement of worn upper seal, glu membrane renewed, correct calf waited, QK ok | Defect | Electrode | Electrode | Glu
ABL800 BASIC | Replace the ref membrane.Calibrate system multiple time. Run QC & pass. | Replace the ref membrane.Calibrate system multiple time. Run QC & pass. | SWO-142872 | No | Replace the ref membrane. Calibrate system multiple time. Run QC & pass. | Replace the ref membrane. Calibrate system multiple time. Run QC & pass. | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX | 调整仪器校准，并重启后解决故障 | 调整仪器校准，并重启后解决故障 | SWO-141996 | No | Adjust the instrument calibration and restart to resolve the problem | Adjust the instrument calibration and restart to resolve the problem | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | 0534 El Met upper liquid sensor errors. Found membrane in flat membrane valve in El Met module damaged. | Replaced flat membrane in El Met module. Performed pump cal and LS adjust. Also corrected sound settings. | D18D66C7D25E64E6 | No | 0534 El Met upper liquid sensor errors. Found membrane in flat membrane valve in El Met module damaged. | Replaced flat membrane in El Met module. Performed pump cal and LS adjust. Also corrected sound settings. | Defect | HW | El/Met Module | Valve | Membrane
ABL825 FLEX | Found errors 664, 653 and asymetery error on pHBG pump.found pump tubing to be very flat. | Replaced pump tubing.Customer stock. | D193AA152032E5D2 | No | Found errors 664, 653 and asymetery error on pHBG pump.found pump tubing to be very flat. | Replaced pump tubing. Customer stock. | Defect | HW | pH/BG Module | Pump tubes
AQT90 FLEX | Beim Eintreffen eine verbrauchte TnT Testkassette im Gerät => Überprüfung Testkassette => 8-Test Testkassette Vermutlich wurde 8-Test TnT Kit falsch bestellt Klärung mit MT | Beim Eintreffen eine verbrauchte TnT Testkassette im Gerät => Überprüfung Testkassette => 8-Test Testkassette Vermutlich wurde 8-Test TnT Kit falsch bestellt Klärung mit MT | SWO-141764 | No | On arrival a used TnT test cassette in the device => Verification test cassette => 8 test cartridge Probably was ordered 8-Test TnT Kit incorrectly Clarification with MT | On arrival a used TnT test cassette in the device => Verification test cassette => 8 test cartridge Probably was ordered 8-Test TnT Kit incorrectly Clarification with MT | User error | Logistical upSWO | Logistical upSWO | Logistical upSWO
AQT90 FLEX | 更换了处理盘传感器636-533和PCB板903-031后修复(byzhenjie) | 更换了处理盘传感器636-533和PCB板903-031后修复(byzhenjie) | SWO-141232 | No | Repair after replacing the processing disk sensor 636-533 and PCB board 903-031 (byzhenjie) | Repair after replacing the processing disk sensor 636-533 and PCB board 903-031 (byzhenjie) | Defect | HW | Process Module | PCB, Cup transport - Rear
ABL80 FLEX CO-OX | Schlauchsystem gespült, Analysator aus und wieder eingeschaltet, Kal, QC | Schlauchsystem gespült, Analysator aus und wieder eingeschaltet, Kal, QC | SWO-142286 | No | Hose system flushed, analyzer off and back on, Kal, QC | Hose system flushed, analyzer off and back on, Kal, QC
ABL80 FLEX BASIC | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-142336 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL90 | CONFIRMED ISSUE OVER PHONE. | CUSTOMER HAD ALREADY REPLACED SC BUT ERROR REMAINED. GUIDED CUSTOMER TO REPLACE SC AGAIN, INSPECT (GOLD) CONTACT PINS ARE CLEAN AND STRAIGHT, REBOOT ANALYZER. | D18E8418480CFDCB | No | CONFIRMED ISSUE OVER PHONE. | CUSTOMER HAD ALREADY REPLACED SC BUT ERROR REMAINED. GUIDED CUSTOMER TO REPLACE SC AGAIN, INSPECT (GOLD) CONTACT PINS ARE CLEAN AND STRAIGHT, REBOOT ANALYZER.
ABL90 FLEX PLUS |  | B | D18ED0514843D8EB | No
AQT90 FLEX | 更换902-924一个正常。(by zhenjie) | 更换902-924一个正常。(by zhenjie) | SWO-141755 | No | Replace the 902-924 one normal. (by zhenjie) | Replace the 902-924 one normal. (by zhenjie) | Defect | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX |  | B | SWO-142101 | No |  |  | Defect | HW | Incubator module | PCB, Incubator
ABL825 FLEX | Nadel justiert und Einlaßdichtung gewechselt Initialisierung abgewartet mehrere 2Pkt. Kal. | Nadel justiert und Einlaßdichtung gewechselt Initialisierung abgewartet mehrere 2Pkt. Kal. | SWO-142261 | No | Needle adjusted and inlet seal changed initialization waited several 2pkt. Kal. | Needle adjusted and inlet seal changed initialization waited several 2pkt. Kal. | Defect | HW | Inlet module | Probe bend
ABL815 FLEX | Fehlersuche, Ref-Elektrode ersetzt (alt: 07-12, neu: 47-04), alle Elektrolyte membraniert, akz. Kal abgewartet, Kal i.O., QC i.O. | Fehlersuche, Ref-Elektrode ersetzt (alt: 07-12, neu: 47-04), alle Elektrolyte membraniert, akz. Kal abgewartet, Kal i.O., QC i.O. | SWO-142248 | No | Troubleshooting, ref electrode replaced (old: 07-12, new: 47-04), all electrolytes membraned, akz. Kal waited, Kal i.O., QC i.O. | Troubleshooting, ref electrode replaced (old: 07-12, new: 47-04), all electrolytes membraned, akz. Kal waited, Kal i.O., QC i.O. | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX |  | B | D18EA07C482B6757 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL800 BASIC | Reconnected all connection in oxy module and then restart analyzer. Problem is solved. Instrument is working fine, | Reconnected all connection in oxy module and then restart analyzer. Problem is solved. Instrument is working fine, | SWO-142850 | No | Reconnected all connection in oxy module and then restart analyzer. Problem is solved. Instrument is working fine, | Reconnected all connection in oxy module and then restart analyzer. Problem is solved. Instrument is working fine, | Unknown | HW | Oxi module | Oxi module
ABL820 FLEX | found that the lamp unit needed to be adjusted again. | adjusted the lamp unit on the analyzer corrected the issue. | D18EA04E4856B5E9 | No | found that the lamp unit needed to be adjusted again. | adjusted the lamp unit on the analyzer corrected the issue. | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Vison controll upgare kit installed to analyser and check working fine. Check for service programes. Analyser is in ready mode. Sample run but aborted after 2 min. Analyser open and check for sample run found cartredge and punch out stuking and making nolice. and in automatic recovery again come OK. Doubt on Storage wheel PCB- Need to replace first. | Vison controll upgare kit installed to analyser and check working fine. Check for service programes. Analyser is in ready mode. Sample run but aborted after 2 min. Analyser open and check for sample run found cartredge and punch out stuking and making nolice. and in automatic recovery again come OK. Doubt on Storage wheel PCB- Need to replace first. | SWO-142470 | No | Vison controll upgare kit installed to analyser and check working fine. Check for service programes. Analyser is in ready mode. Sample run but aborted after 2 min. Analyser open and check for sample run found cartredge and punch out stuking and making nolice. and in automatic recovery again come OK. Doubt on Storage wheel PCB- Need to replace first. | Vison controll upgare kit installed to analyser and check working fine. Check for service programes. Analyser is in ready mode. Sample run but aborted after 2 min. Analyser open and check for sample run found cartredge and punch out stuking and making nolice. and in automatic recovery again come OK. Doubt on Storage wheel PCB- Need to replace first. | Defect | HW | Aspiration Module | Barcode reader, Inlet
ABL80 FLEX CO-OX | Fliesswege gespült (Hemolyzer), Nullpunktkal mithilfe zersing Kassette durchgeführt. | B | SWO-142458 | No | Flow paths rinsed (hemolyzer), zero-point scale performed using zersing cassette.
ABL80 FLEX Basic version | Revisión de consumibles y tubos. El pack de soluciones está obstruido y no sale líquido por uno de los conectores. Se observa suciedad en salida de desperdicios. Se sustituye pack y piezas de kit anual y se comprueba funcionamiento con controles de calidad. | B | D19390BB203AA937 | No | Review of consumables and tubes. The solution pack is clogged and no liquid is released from one of the connectors. Dirt is observed in the exit of waste. Annual kit and kit parts are replaced and operation is checked with quality controls.
ABL80 CO-OX | Checked data logs. No errors evident. Customer stated some samples were lost due to the error popping up. | Replaced i button probe base, cable and board. | D18FD2CEBE2AE7BA | No | Checked data logs. No errors evident. Customer stated some samples were lost due to the error popping up. | Replaced i button probe base, cable and board.
ABL90 FLEX | LP gewechselt | B | SWO-142650 | No | LP changed |  | User error | SP empty | SP empty | SP empty
ABL90 | Unit failing levels 1 and 3 Internal.Level 1 wil be low, and level 3 will be high.Intermittent Calibration errors.Unable to logon to unit as administrator.Suspect Silicon contamination to SC. | Performed Complete Software Install .Reseated all connections from Wet Section Bd. Flushed system.Inspected Fluid path. | D18FB79ABE2667DF | No | Unit failing levels 1 and 3 Internal.Level 1 wil be low, and level 3 will be high. Intermittent Calibration errors. Unable to logon to unit as administrator. Suspect Silicon contamination to SC. | Performed Complete Software Install . Reseated all connections from Wet Section Bd. Flushed system. Inspected Fluid path. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Silicone
ABL90 | Client advised analyzer not switching back from batter to AC after move | OSPD found battery switch not detecting AC current when plugged back in to power source. Swapped power cables, tried different outlet, no change. fse to order new Battery Option. | D18FCCF1BE20E406 | No | Client advised analyzer not switching back from batter to AC after move | OSPD found battery switch not detecting AC current when plugged back in to power source. Swapped power cables, tried different outlet, no change. fse to order new Battery Option.
ABL90 | FOUND SYSTEM NORMAL ALL LOGS NORMAL | FLUSHED SYSTEM | D1904FE6BE0E3CAE | No | FOUND SYSTEM NORMAL ALL LOGS NORMAL | FLUSHED SYSTEM
ABL90 FLEX PLUS |  | B | D18FE27FBE322104 | Yes |  |  | Re classified RMED | Logistical upSWO | PER
ABL90 FLEX PLUS | Nettoyage du système fluidique e.o Test divers à l'aide du programme de service e.o Ajustement du système introduction e.o Calibration e.o contrôle de qualité e.o | Nettoyage du système fluidique e.o Test divers à l'aide du programme de service e.o Ajustement du système introduction e.o Calibration e.o contrôle de qualité e.o | SWO-142478 | No | Cleaning the fluid system e.o Various test using the service program e.o Adjustment of the system introduction e.o Calibration e.o quality control e.o | Cleaning the fluid system e.o Various test using the service program e.o Adjustment of the system introduction e.o Calibration e.o quality control e.o | Assisted Replacement | SP | SP | Leak current
AQT90 FLEX | 更换了903-364一个后修复(by zhenjie) | 更换了903-364一个后修复(by zhenjie) | SWO-142032 | No | Replaced 903-364 a post-fix (by zhenjie) | Replaced 903-364 a post-fix (by zhenjie) | Defect | HW | Process Module | Optical Unit | 129
ABL837 FLEX | Performed OXI test fails lamp voltage | Performed complete OXI aligment. Failed Crea on level 4 Changed the c rea membranes and took 5 2Pt Cals | D18FB6AEBE57255B | No | Performed OXI test fails lamp voltage | Performed complete OXI aligment. Failed Crea on level 4 Changed the c rea membranes and took 5 2Pt Cals | Adjustment | HW | Oxi module | Lamp unit
ABL90 FLEX PLUS | Replaced solution pack. Replaced sensor cassette. | Replaced solution pack. Replaced sensor cassette. | SWO-143058 | No | Replaced solution pack. Replaced sensor cassette. | Replaced solution pack. Replaced sensor cassette. | Adjustment | SW | Off/On | Off/On | 145
ABL800 BASIC | found repalced the Ref membrane calibrated found Na results are OK. Insturment is functional. | found repalced the Ref membrane calibrated found Na results are OK. Insturment is functional. | SWO-144538 | No | found repalced the Ref membrane calibrated found Na results are OK. Insturment is functional. | found repalced the Ref membrane calibrated found Na results are OK. Insturment is functional. | Assisted Replacement | Electrode | Electrode | Reference | LEH ex
ABL80 FLEX BASIC | 更换902-972修复(by zhenjie) | 更换902-972修复(by zhenjie) | SWO-142887 | No | Replacement 902-972 repair (by zhenjie) | Replacement 902-972 repair (by zhenjie)
AQT90 FLEX | CİHAZIN STROGE WHEEL HOME SENSÖRLERİ TEMİZLENİP,CUP DRYER FİLİTRESİ DEĞİŞTİRİLEREK SORUN GİDERİLDİ. | CİHAZIN STROGE WHEEL HOME SENSÖRLERİ TEMİZLENİP,CUP DRYER FİLİTRESİ DEĞİŞTİRİLEREK SORUN GİDERİLDİ. | SWO-142949 | Yes | CİHAZIN STROGE WHEEL HOME SENSORS HAVE BEEN SECURED AND THE PROBLEM HAS BEEN ADDRESSED BY CHANGING THE CUP DRYER FILM. | CİHAZIN STROGE WHEEL HOME SENSORS HAVE BEEN SECURED AND THE PROBLEM HAS BEEN ADDRESSED BY CHANGING THE CUP DRYER FILM.
ABL90 | Upon arrival it was discovered that the time was correct and they were not using a time stamp and the send tab was available to send results. | Waiting for further instructions. | D193A86B200A211B | No | Upon arrival it was discovered that the time was correct and they were not using a time stamp and the send tab was available to send results. | Waiting for further instructions. | No fault found | SW | Time setting | Time setting
ABL90 | Client advised Sensor Bay door is not opening closing properly and they are unable to swap sensor cassette. | OSPD determined potential loose cable as door began to work after analyzer was moved on the counter. Conditioned calibrated QCd new Sensor Cassette for client. FSE ordered and replaced Stepper motor for sensor bay door. | D19386A02041E40D | No | Client advised Sensor Bay door is not opening closing properly and they are unable to swap sensor cassette. | OSPD determined potential loose cable as door began to work after analyzer was moved on the counter. Conditioned calibrated QCd new Sensor Cassette for client. FSE ordered and replaced Stepper motor for sensor bay door. | Defect | HW | Cover and Chassis | Stepper Motor, SC door
ABL90 FLEX | Straightened bent inlet probe. Replaced Sensor Cassette. | Straightened bent inlet probe. Replaced Sensor Cassette. | SWO-143275 | No | Straightened bent inlet probe. Replaced Sensor Cassette. | Straightened bent inlet probe. Replaced Sensor Cassette. | Adjustment | HW | Inlet module | Probe bend
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-145592 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | SP | 189
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-145590 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | SP | 181
AQT90 FLEX | 更换903-364一个正常。by jesson | 更换903-364一个正常。by jesson | SWO-142035 | No | Replace 903-364 one normal. by jesson | Replace 903-364 one normal. by jesson | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | Replaced lactate membrane. Cleaned measuring chamber. | Replaced lactate membrane. Cleaned measuring chamber. | SWO-143371 | No | Replaced lactate membrane. Cleaned measuring chamber. | Replaced lactate membrane. Cleaned measuring chamber. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL825 FLEX | Remplacement de l'imprimante par un second main. Mis en route de l'ABL et essai : OK Cette intervention est sous garantie. | Remplacement de l'imprimante par un second main. Mis en route de l'ABL et essai : OK Cette intervention est sous garantie. | SWO-143180 | No | Replacing the printer with a second hand. Started the ABL and test: OK This intervention is under warranty. | Replacing the printer with a second hand. Started the ABL and test: OK This intervention is under warranty. | Defect | HW | Printer module | Printer Head
ABL825 FLEX | Verified insert ranges that was being barcoded in was matching instrument and they were not. Upon further review I found the altitude was set to meters. | Set altitude to 5280ft. | D1937333201B4F52 | No | Verified insert ranges that was being barcoded in was matching instrument and they were not. Upon further review I found the altitude was set to meters. | Set altitude to 5280ft. | User error | SW | Setup | altitude
ABL815 FLEX | Membran Air Ventil PH BG Modul sichtbar gerissen - Ventil ausgebaut Membran ersetzt, Membran Air Ventil E-Lyte Modul überprüft - Membran ebenfalls def - Membran ersetzt, Kal und QC durchgeführt | Membran Air Ventil PH BG Modul sichtbar gerissen - Ventil ausgebaut Membran ersetzt, Membran Air Ventil E-Lyte Modul überprüft - Membran ebenfalls def - Membran ersetzt, Kal und QC durchgeführt | SWO-143161 | No | Membrane Air Valve PH BG Module visibly ripped - Valve removed membrane replaced, Membrane Air Valve E-Lyte module checked - membrane also def - membrane replaced, Kal and QC performed | Membrane Air Valve PH BG Module visibly ripped - Valve removed membrane replaced, Membrane Air Valve E-Lyte module checked - membrane also def - membrane replaced, Kal and QC performed | Defect | HW | pH/BG Module | Valve, air | Membrane
ABL80 FLEX BASIC | 检查机器发现由于测试卡漏夜，造成902-272短路，引起自动关机，更换新的902-272后，机器不能正常工作。更换903-334后修复 by jesson | 检查机器发现由于测试卡漏夜，造成902-272短路，引起自动关机，更换新的902-272后，机器不能正常工作。更换903-334后修复 by jesson | SWO-142935 | No | Check the machine found that due to the test card leakage night, resulting in 902-272 short circuit, causing automatic shutdown, replacement of the new 902-272, the machine does not work properly. Repair by jesson after replacing 903-334 | Check the machine found that due to the test card leakage night, resulting in 902-272 short circuit, causing automatic shutdown, replacement of the new 902-272, the machine does not work properly. Repair by jesson after replacing 903-334
ABL80 FLEX CO-OX OSM | Solution Pack memory chip clean & position adjust. Pre-Maintenance 진행 Liquid flow line clean. Sensor cassette clean. System cycle check. Data delete work. 담당자 변경으로 인해 User Training 실시함. | B | SWO-143391 | No | Solution Pack memory chip clean & position adjust. Pre-Maintenance 진행 Liquid flow line clean. Sensor cassette clean. System cycle check. Data delete work. 담당자 변경으로 인해 User Training 실시함.
ABL80 CO-OX | Inspected SC found pin to be bent over. | Replaced SC. | D1A19095325ABEF9 | No | Inspected SC found pin to be bent over. | Replaced SC.
ABL80 OSM | Verify issue. Found blockage in tubing waste line. For the loose monitor - Found screws that holds the monitor are loose. | For the fluid blockage replaced tubing waste line, pump head and waste manifold. For the loose monitor tighten screws holding the monitor. | D195132A962C73DA | No | Verify issue. Found blockage in tubing waste line. For the loose monitor - Found screws that holds the monitor are loose. | For the fluid blockage replaced tubing waste line, pump head and waste manifold. For the loose monitor tighten screws holding the monitor.
ABL90 FLEX | Kal aktiviert danach SC (44-11) aus Kondigerät eingesetzt | Kal aktiviert danach SC (44-11) aus Kondigerät eingesetzt | SWO-143413 | No | Kal activates then SC (44-11) used from confectionery | Kal activates then SC (44-11) used from confectionery | Assisted Replacement | SC | SC | Glu/Lac
ABL90 | Opened up printer lid and took out the paper and inspected the printer for paper jam or something broken and did not find anything. | Closed printer door and it works. | D194ED239652B8DF | No | Opened up printer lid and took out the paper and inspected the printer for paper jam or something broken and did not find anything. | Closed printer door and it works. | No fault found | HW | Printer module | Printer module
ABL90 FLEX | Spild adaptor - slange - ventilslange - forbindelsespakning skiftet. Kal. og QC ok | Spild adaptor - slange - ventilslange - forbindelsespakning skiftet. Kal. og QC ok | SWO-143441 | No | Spillage adaptor-Hose-valve Hose-connection packing switched. Kal. and QC OK | Spillage adaptor-Hose-valve Hose-connection packing switched. Kal. and QC OK | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX | Fehlersuche, Reinigung Oximetrieventil, Neustart, Funktionskontrolle | Fehlersuche, Reinigung Oximetrieventil, Neustart, Funktionskontrolle | SWO-143508 | No | Troubleshooting, cleaning oximetry valve, restart, function control | Troubleshooting, cleaning oximetry valve, restart, function control | Defect | HW | Oxi module | Pinch Valve
ABL90 | Troubleshot and found that the connection settings had reverted back to DCHP. | Inserted analyzer static IP settings | D194B0E2962D2A16 | No | Troubleshot and found that the connection settings had reverted back to DCHP. | Inserted analyzer static IP settings | Setup | SW | His/Lis | Radiance
ABL90 FLEX PLUS | Analysator geht nach Neustart am 31.10 23:06 wieder um 1 Std. nach. Letzte Zeitserver Synchronisation 31.10 23:26. Einstellungen alle korrekt, nach Neustart bleibt die Uhrzeit korrekt, Zeitserver Synchronisation klappt auch. Datalogs aufgenommen. | Analysator geht nach Neustart am 31.10 23:06 wieder um 1 Std. nach. Letzte Zeitserver Synchronisation 31.10 23:26. Einstellungen alle korrekt, nach Neustart bleibt die Uhrzeit korrekt, Zeitserver Synchronisation klappt auch. Datalogs aufgenommen. | SWO-143889 | No | Analyzer goes back to 1 hr after restart on 31.10 23:06. Last time server synchronization 31.10 23:26. Settings all correct, after restart the time remains correct, time server synchronization also works. datalogs. | Analyzer goes back to 1 hr after restart on 31.10 23:06. Last time server synchronization 31.10 23:26. Settings all correct, after restart the time remains correct, time server synchronization also works. datalogs. | Setup | SW | Time setting | Time setting
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-145591 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | SP | 1244
AQT90 FLEX | 更换光学单元 | 更换光学单元 | SWO-143357 | No | Replace the optical unit | Replace the optical unit | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 室内温度过低，升高室内温度后，正常。 | 室内温度过低，升高室内温度后，正常。 | SWO-143064 | No | The indoor temperature is too low, after raising the room temperature, normal. | The indoor temperature is too low, after raising the room temperature, normal. | Defect | HW | Process Module | PCB, Cup dryer | 168
AQT90 FLEX | 更换配件903-210后，重新校准后修复 by jesson | 更换配件903-210后，重新校准后修复 by jesson | SWO-142936 | No | Repair by jesson after recalibration after replacing accessory 903-210 | Repair by jesson after recalibration after replacing accessory 903-210 | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL825 FLEX | Lac membraniert, Funktionskontrolle | Lac membraniert, Funktionskontrolle | SWO-143435 | No | Lake membraniert, Funktionskontrolle | Lake membraniert, Funktionskontrolle | Assisted Replacement | Electrode | Membrane | Lac | Unstable
ABL837 FLEX | Performed OXI test in srvice, fails lamp voltage | Adjusted lamp volrage | D194C1519628920B | No | Performed OXI test in srvice, fails lamp voltage | Adjusted lamp volrage | Adjustment | HW | Oxi module | Lamp unit | Halogen
ABL825 FLEX | Konfiguration geprüft und geändert, Funktionskontrolle ok | Konfiguration geprüft und geändert, Funktionskontrolle ok | SWO-143681 | No | Configuration checked and changed, function control ok | Configuration checked and changed, function control ok | Setup | SW | Roma.ini | Print layout
ABL80 FLEX BASIC | 更换测试卡和试剂包修复 | 更换测试卡和试剂包修复 | SWO-143644 | No | Replacement test card and reagent kit repair | Replacement test card and reagent kit repair
ABL80 FLEX | tried to reinstalled Sc problem remains, need to replace SC with new one. | tried to reinstalled Sc problem remains, need to replace SC with new one. | SWO-144479 | No | tried to reinstalled Sc problem remains, need to replace SC with new one. | tried to reinstalled Sc problem remains, need to replace SC with new one.
ABL80 FLEX CO-OX | Habe den Einlass gereinigt. Mehrfach den Analysator. Sich spülen lassen. Danach Kalibration und QC durchgeführt. Im Anschluss noch THB Kalibration durchgeführt. Auch hier keine Fehlermeldung. Es lief keine Flüssigkeit aus. Es gab auch keine andere Fehlermeldung. Nach durchgeführten Maßnahmen befand sich der Analysator im Messbereiten Zustand. | Habe den Einlass gereinigt. Mehrfach den Analysator. Sich spülen lassen. Danach Kalibration und QC durchgeführt. Im Anschluss noch THB Kalibration durchgeführt. Auch hier keine Fehlermeldung. Es lief keine Flüssigkeit aus. Es gab auch keine andere Fehlermeldung. Nach durchgeführten Maßnahmen befand sich der Analysator im Messbereiten Zustand. | SWO-143820 | No | Have cleaned the inlet. Multiple times the analyzer. Let yourself be rinsed. Then calibrated and QC performed. Subsequently, THB calibration was carried out. Again, no error message. There was no liquid leaking. There was no other error message. After the measures were carried out, the analyzer was in a ready-to-measure state. | Have cleaned the inlet. Multiple times the analyzer. Let yourself be rinsed. Then calibrated and QC performed. Subsequently, THB calibration was carried out. Again, no error message. There was no liquid leaking. There was no other error message. After the measures were carried out, the analyzer was in a ready-to-measure state.
ABL80 FLEX BASIC |  | B | SWO-143547 | No
ABL80 FLEX BASIC | Sensor board foung burn and SC too. try to clean and install but same problem reported. Need to replace Sensor Board- Zeroing calibration than SC. | Sensor Board Request and zeroing cassette request sent. Standby analyser arranged from Indore will be receiving tomorrow. Service report handed over. | SWO-143827 | No | Sensor board foung burn and SC too. try to clean and install but same problem reported. Need to replace Sensor Board- Zeroing calibration than SC. | Sensor Board Request and zeroing cassette request sent. Standby analyser arranged from Indore will be receiving tomorrow. Service report handed over.
ABL80 CO-OX | Paula went to change the solution pack because it was time and found fluid in the SP cavity. I checked for leaks and did not see any.I checked the waste manifold and it was clear. I believe the leak was coming from the other SP. The solution was not going in to the waste port. | Paula put on another SP. | D19625270C26C3AA | No | Paula went to change the solution pack because it was time and found fluid in the SP cavity. I checked for leaks and did not see any. I checked the waste manifold and it was clear. I believe the leak was coming from the other SP. The solution was not going in to the waste port. | Paula put on another SP.
ABL80 FLEX BASIC | done zeroing of ABL80. and also provide a training about the preanalytical. performed the samples now the analyser is working satisfactorily. | done zeroing of ABL80. and also provide a training about the preanalytical. performed the samples now the analyser is working satisfactorily. | SWO-144372 | No | done zeroing of ABL80. and also provide a training about the preanalytical. performed the samples now the analyser is working satisfactorily. | done zeroing of ABL80. and also provide a training about the preanalytical. performed the samples now the analyser is working satisfactorily.
ABL90 |  | B | D196559A0C423B88 | No |  |  | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Überprüfung des Analysators im Service. Werte innerhalb Grenzwerte in Ordnung. Keine Beanstandung an Wartungsteilen. Anwender sehr unzufrieden mit Situation der BGA Messung. Falsche Tastatur hinterlegt, korriegiert auf Windowsebene. Konditionierung der Sensorkassette braucht zu lange. Metabolite nach 1,5 Stunden nicht stabil. Kalibrationen in Ordnung Qualitätskontrolle bei Metabolite nicht in Ordnung. | Überprüfung des Analysators im Service. Werte innerhalb Grenzwerte in Ordnung. Keine Beanstandung an Wartungsteilen. Anwender sehr unzufrieden mit Situation der BGA Messung. Falsche Tastatur hinterlegt, korriegiert auf Windowsebene. Konditionierung der Sensorkassette braucht zu lange. Metabolite nach 1,5 Stunden nicht stabil. Kalibrationen in Ordnung Qualitätskontrolle bei Metabolite nicht in Ordnung. | SWO-143806 | No | Checking the analyzer in the service. Values within limits in order. No complaint about maintenance parts. Users very dissatisfied with the situation of BGA measurement. Incorrect keyboard stored, corrieled at the Windows level. Conditioning of the sensor cartridge takes too long. metabolites are not stable after 1.5 hours. Calibrations in order quality control with metabolites not in order. | Checking the analyzer in the service. Values within limits in order. No complaint about maintenance parts. Users very dissatisfied with the situation of BGA measurement. Incorrect keyboard stored, corrieled at the Windows level. Conditioning of the sensor cartridge takes too long. metabolites are not stable after 1.5 hours. Calibrations in order quality control with metabolites not in order. | Assisted Replacement | SC | SC | Start up
ABL90 FLEX | manuell gespült LP getauscht ohne Erfolg SC getauscht manuell kalibriert System IO | manuell gespült LP getauscht ohne Erfolg SC getauscht manuell kalibriert System IO | SWO-144101 | No | manually rinsed LP exchanged without success SC exchanged manually calibrated system IO | manually rinsed LP exchanged without success SC exchanged manually calibrated system IO | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Power supply is not switching back to main power when plugged back in to AC. | Replaced power supply. | D1961AA10C52DFB6 | No | Power supply is not switching back to main power when plugged back in to AC. | Replaced power supply. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX PLUS | Auf der Windows Oberfläche die Aktuelle Uhrzeit eingestellt und mittels USB Stick mit dem Befehl Commit2CDrive gespeichert. Den ABL mehrmals neu gestartet die Aktuelle Uhrzeit bleibt bestehen. | Auf der Windows Oberfläche die Aktuelle Uhrzeit eingestellt und mittels USB Stick mit dem Befehl Commit2CDrive gespeichert. Den ABL mehrmals neu gestartet die Aktuelle Uhrzeit bleibt bestehen. | SWO-144041 | No | Set the current time on the Windows interface and save it using the COMMIT2CDrive command using a USB stick. The ABL restarts the current time several times. | Set the current time on the Windows interface and save it using the COMMIT2CDrive command using a USB stick. The ABL restarts the current time several times. | Setup | SW | Time setting | Time setting
ABL90 FLEX PLUS | Auf der Windows Oberfläche die Aktuelle Uhrzeit eingestellt und mittels USB Stick mit dem Befehl Commit2CDrive gespeichert. Den ABL neu gestartet die Aktuelle Uhrzeit bleibt bestehen. | Auf der Windows Oberfläche die Aktuelle Uhrzeit eingestellt und mittels USB Stick mit dem Befehl Commit2CDrive gespeichert. Den ABL neu gestartet die Aktuelle Uhrzeit bleibt bestehen. | SWO-144006 | No | Set the current time on the Windows interface and save it using the COMMIT2CDrive command using a USB stick. The ABL restarts the current time remains. | Set the current time on the Windows interface and save it using the COMMIT2CDrive command using a USB stick. The ABL restarts the current time remains. | Setup | SW | Time setting | Time setting
ABL800 BASIC | Found that problem with Ref electrode replaced the standby electrode and membrane calibrated performed the QC found Na results are OK, instrument is functional, 954-603- Lot: 3714 | Found that problem with Ref electrode replaced the standby electrode and membrane calibrated performed the QC found Na results are OK, instrument is functional, 954-603- Lot: 3714 | SWO-144537 | No | Found that problem with Ref electrode replaced the standby electrode and membrane calibrated performed the QC found Na results are OK, instrument is functional, 954-603- Lot: 3714 | Found that problem with Ref electrode replaced the standby electrode and membrane calibrated performed the QC found Na results are OK, instrument is functional, 954-603- Lot: 3714 | Assisted Replacement | Electrode | Electrode | Reference | LEH ex
ABL825 FLEX | checked data and cal logs for errors, checked valves and sensors, checked electrode and membrane, checked Zero current | removed modules and cleaned valves, performed zero current cal, removed and replaced lactate electrode and membrane, | D1968D980C4152F9 | No | checked data and cal logs for errors, checked valves and sensors, checked electrode and membrane, checked Zero current | removed modules and cleaned valves, performed zero current cal, removed and replaced lactate electrode and membrane, | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL825 FLEX | Monitor ausgetauscht zusammen mit Jonne Marcher von RMED. Test des neuen Monitors, OK | Zu letztem SWO Arrival Departure Time nachgetragen - keine weiteren Aktionen. | SWO-143743 | No | Monitor exchanged together with Jonne Marcher from RMED. Test of the new monitor, OK | Added to the last SWO Arrival Departure Time - no further actions. | Defect | HW | DMS_PC | Touch screen | Frozen
ABL825 FLEX | Zu letztem SWO Arrival Departure Time nachgetragen - keine weiteren Aktionen. | Zu letztem SWO Arrival Departure Time nachgetragen - keine weiteren Aktionen. | SWO-143746 | Yes | Added to the last SWO Arrival Departure Time - no further actions. | Added to the last SWO Arrival Departure Time - no further actions. | Re classified RMED | Gemba visit
ABL825 FLEX | Besuch dieses und weiterer Geräte bezüglich der frozen screen Problematik zusammen mit Jonne Marcher von RMED. | Besuch dieses und weiterer Geräte bezüglich der frozen screen Problematik zusammen mit Jonne Marcher von RMED. | SWO-143886 | Yes | Visit this and other devices regarding the frozen screen problem together with Jonne Marcher from RMED. | Visit this and other devices regarding the frozen screen problem together with Jonne Marcher from RMED. | Re classified RMED | Gemba visit
ABL800 BASIC | Replaced Neon Lamp, performed complete oximetry test successfully. 2 point calibration performed successfully | Replaced Neon Lamp, performed complete oximetry test successfully. 2 point calibration performed successfully | SWO-143595 | No | Replaced Neon Lamp, performed complete oximetry test successfully. 2 point calibration performed successfully | Replaced Neon Lamp, performed complete oximetry test successfully. 2 point calibration performed successfully | Defect | HW | Oxi module | Lamp unit | Neon - Low
AQT90 FLEX | Netzwerkdaten und Aktivitätenlog überprüft- seit 01.01.2016 13.12Uhr fehlt die Verbindung zu Aqure - Netzwerkdaten fehlten komplett, Daten wurden am 31.10. korrekt eingeben und abgespeichert, erneut eingegeben und auf Datenverlustnach Neustart überprüft, i.O. | Netzwerkdaten und Aktivitätenlog überprüft- seit 01.01.2016 13.12Uhr fehlt die Verbindung zu Aqure - Netzwerkdaten fehlten komplett, Daten wurden am 31.10. korrekt eingeben und abgespeichert, erneut eingegeben und auf Datenverlustnach Neustart überprüft, i.O. | SWO-143876 | No | Network data and activity log checked - since 01.01.2016 13.12o the connection to Aqure is missing completely - network data was completely missing, data was entered and saved correctly on 31.10. , re-entered and checked for data loss after restart, i.O. | Network data and activity log checked - since 01.01.2016 13.12o the connection to Aqure is missing completely - network data was completely missing, data was entered and saved correctly on 31.10. , re-entered and checked for data loss after restart, i.O. | Setup | SW | His/Lis | His/Lis
TCM TOSCA | STK durchgeführt. GW nicht überschritten. Funktionskontrolle und Sichtkontrolle ok. | STK durchgeführt. GW nicht überschritten. Funktionskontrolle und Sichtkontrolle ok. | SWO-143786 | No | STK. GW not exceeded. Functional control and visual inspection ok. | STK. GW not exceeded. Functional control and visual inspection ok.
ABL80 FLEX | After Checking found that leakage from Sensor Cassette,Tried to installed New SC but Problem Remains, Need to replace Sensor Board Connector with new one. | B | SWO-144227 | No | After Checking found that leakage from Sensor Cassette,Tried to installed New SC but Problem Remains, Need to replace Sensor Board Connector with new one.
ABL80 FLEX BASIC | Sensor Board Request and zeroing cassette request sent. Standby analyser arranged from Indore will be receiving tomorrow. Service report handed over. | Sensor Board Request and zeroing cassette request sent. Standby analyser arranged from Indore will be receiving tomorrow. Service report handed over. | SWO-143920 | No | Sensor Board Request and zeroing cassette request sent. Standby analyser arranged from Indore will be receiving tomorrow. Service report handed over. | Sensor Board Request and zeroing cassette request sent. Standby analyser arranged from Indore will be receiving tomorrow. Service report handed over.
ABL90 | Client advised of error 1002: pO2 Dark Data is out of range | ospd found pO2 channel failing sensor simulator test, verified alignment. Replaced sensor assembly PCB. Replaced inlet gasket to fix failed vaccum test. | D1978AA2825AE419 | No | Client advised of error 1002: pO2 Dark Data is out of range | ospd found pO2 channel failing sensor simulator test, verified alignment. Replaced sensor assembly PCB. Replaced inlet gasket to fix failed vaccum test. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 | DATA LOGS SHOW INHOMOGENOUS CAL 3 SOLUTION | FLUSHED CASETTE AND ANALYZER,PERFORMED FLUID TRANSPORT ADJUSTMENT.PRIMED WITH BLOOD.IF PROBLEM RETURNS REPLACE SOLUTION PACK | D1977E33820A5FA3 | No | DATA LOGS SHOW INHOMOGENOUS CAL 3 SOLUTION | FLUSHED CASETTE AND ANALYZER,PERFORMED FLUID TRANSPORT ADJUSTMENT. PRIMED WITH BLOOD. IF PROBLEM RETURNS REPLACE SOLUTION PACK | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | tHb results on ABL90 are 0.6 to 0.9 g dL higher than on ABL80 and Beckman Coulter hematology analyzers. | Aligned ABL80 and 90 and ran comparisons. All results were 0,6 g dL higher. Replaced hemolyzer unit and results improved to 0.3 g dL higher. Installed Win 7 software and restored data and setup. Calibrated touch screen. Replaced solution pack to correct pO2 QC errors. Will fax form in for replacement. Performed fluid transport programs. | D1977B508203F830 | No | tHb results on ABL90 are 0.6 to 0.9 g dL higher than on ABL80 and Beckman Coulter hematology analyzers. | Aligned ABL80 and 90 and ran comparisons. All results were 0,6 g dL higher. Replaced hemolyzer unit and results improved to 0.3 g dL higher. Installed Win 7 software and restored data and setup. Calibrated touch screen. Replaced solution pack to correct pO2 QC errors. Will fax form in for replacement. Performed fluid transport programs. | Defect | HW | Oxi module | Hemolyzer
ABL90 | CONFIRMED ISSUE. | REPLACED FLASH CARD AND RELOADED BOTH WIINDOWS XPe AND ANALYZER SW. | D1979562823ACD55 | No | CONFIRMED ISSUE. | REPLACED FLASH CARD AND RELOADED BOTH WIINDOWS XPe AND ANALYZER SW. | Defect | HW | DMS_PC | CF Card
ABL90 FLEX PLUS |  | B | D19F61CE462B4AA9 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Mail
ABL90 FLEX PLUS | Auf der Windows Oberfläche die Aktuelle Uhrzeit eingestellt und mittels USB Stick mit dem Befehl Commit2CDrive gespeichert. Den ABL mehrmals neu gestartet die Aktuelle Uhrzeit bleibt bestehen. | Auf der Windows Oberfläche die Aktuelle Uhrzeit eingestellt und mittels USB Stick mit dem Befehl Commit2CDrive gespeichert. Den ABL mehrmals neu gestartet die Aktuelle Uhrzeit bleibt bestehen. | SWO-143984 | No | Set the current time on the Windows interface and save it using the COMMIT2CDrive command using a USB stick. The ABL restarts the current time several times. | Set the current time on the Windows interface and save it using the COMMIT2CDrive command using a USB stick. The ABL restarts the current time several times.
AQT90 FLEX |  | 更换光学单元903-364一个后修复 | SWO-143944 | No |  | Replacement of optical unit 903-364 one post-repair | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 安装测试卡时发现测试卡不能正常推入到机器，检查发现886-625损坏，不能正常固定在存储盘上。更换886-625后修复 by jesson | 安装测试卡时发现测试卡不能正常推入到机器，检查发现886-625损坏，不能正常固定在存储盘上。更换886-625后修复 by jesson | SWO-143663 | No | When the test card is installed, it is found that the test card can not be pushed into the machine properly, the inspection found that 886-625 is damaged, can not be properly fixed on the storage disk. Repair by jesson after replacing 886-625 | When the test card is installed, it is found that the test card can not be pushed into the machine properly, the inspection found that 886-625 is damaged, can not be properly fixed on the storage disk. Repair by jesson after replacing 886-625 | Defect | HW | Storage Module | Cartridge guide
AQT90 FLEX | 清洁机器内部灰尘，更换定标卡944-216后修复 by jesson | 清洁机器内部灰尘，更换定标卡944-216后修复 by jesson | SWO-143546 | No | Clean the dust inside the machine and repair by jesson after replacing the calibration card 944-216 | Clean the dust inside the machine and repair by jesson after replacing the calibration card 944-216 | Cleaned | HW | Storage Module | Process wheel | 161
ABL825 FLEX |  | B | D1975EAD8253FC2B | Yes |  |  | Re classified RMED | Refurbish Used Analyzer
ABL825 FLEX |  | B | D1975D5B8223ABA5 | Yes |  |  | Re classified RMED | Refurbish Used Analyzer
ABL827 FLEX |  | B | D19C2E2D5A3DABAC | Yes |  |  | Re classified RMED | Used analyzer Refurbished
ABL827 FLEX |  | B | D1975CE5822DFC28 | Yes |  |  | Re classified RMED | Used analyzer Refurbished
ABL835 FLEX | Dekonterminering udført Begge målekamre renset Kal. & QC | Dekonterminering udført Begge målekamre renset Kal. & QC | SWO-143924 | No | Decontermination performed both measuring chambers cleaned Kal. & QC | Decontermination performed both measuring chambers cleaned Kal. & QC | Cleaned | HW | pH/BG Module | Measuring chamber
ABL835 FLEX | pO2 and pCO2 has excessive drift every 5 to 7 days pCO2 has a drift of 420 to 38 intermittently, pO2 has a negative drift -6.3 and below intermittently--- Travel time from Vacaville and then back to Vallejo | replaced bsoth gas 1 and gas 2 regulators replaced the gas chassi | D1974A9D822633F7 | No | pO2 and pCO2 has excessive drift every 5 to 7 days pCO2 has a drift of 420 to 38 intermittently; pO2 has a negative drift -6.3 and below intermittently--- Travel time from Vacaville and then back to Vallejo | replaced bsoth gas 1 and gas 2 regulators replaced the gas chassi | Defect | HW | Gas module | Gas valve 1 and 2
ABL815 FLEX | Abfallschlauch gewechselt, mehrmals gespült, Oxi Modul durchgefallen Gerät neugestartet neon Lamp Unit Intensität zu gering. Flüssigkeit wurde frühzeitig in Oxi Modul erkannt nach 2. kal. Behoben. Oxi Test erneut durchgefallen. Oxi Modul wird bestellt und ausgetauscht. | Abfallschlauch gewechselt, mehrmals gespült, Oxi Modul durchgefallen Gerät neugestartet neon Lamp Unit Intensität zu gering. Flüssigkeit wurde frühzeitig in Oxi Modul erkannt nach 2. kal. Behoben. Oxi Test erneut durchgefallen. Oxi Modul wird bestellt und ausgetauscht. | SWO-144231 | No | Waste hose changed, rinsed several times, Oxi module restarted device neon lamp unit intensity too low. Liquid was detected early in Oxi module after 2nd cal. Fixed. Oxi test failed again. Oxi module is ordered and replaced. | Waste hose changed, rinsed several times, Oxi module restarted device neon lamp unit intensity too low. Liquid was detected early in Oxi module after 2nd cal. Fixed. Oxi test failed again. Oxi module is ordered and replaced. | Defect | HW | Oxi module | Tubing
ABL805 FLEX | Lampenintensität liegt bei 3800 => LampUnit im Servicemenü auf den Sollwert justiert und Kal 1 zur Verifikation durchgeführt. | Lampenintensität liegt bei 3800 => LampUnit im Servicemenü auf den Sollwert justiert und Kal 1 zur Verifikation durchgeführt. | SWO-144007 | No | Lamp intensity is 3800 => LampUnit in the service menu adjusted to the setpoint and Kal 1 is performed for verification. | Lamp intensity is 3800 => LampUnit in the service menu adjusted to the setpoint and Kal 1 is performed for verification. | Adjustment | HW | Oxi module | Lamp unit | Halogen
ABL820 FLEX | found that the user keeps failing qc low outside of packet insert for level 2 s7345 | Replaced CO2 electrode checked all the chambers and orings on electrodes. I also remembraned ref to correct the issue. Also came back next day to verify it worked after sitting a while. | D197509382297657 | No | found that the user keeps failing qc low outside of packet insert for level 2 s7345 | Replaced CO2 electrode checked all the chambers and orings on electrodes. I also remembraned ref to correct the issue. Also came back next day to verify it worked after sitting a while. | Assisted Replacement | Electrode | Electrode | pCO2 | Level 2
ABL825 FLEX | Konnektor Gluc Kontakt def nach durch Hotline angeleitete Maßnahmen durch den Anwender - Konnektor ersetzt, Gluc neu membraniert Empf weiter nur..., Gluc Elektrode ersetzt Empf. bei 800 , Ref Elektrode wegen QC Fehler ELyte neu membraniert, Kontakt des Konnektors gereinigt, Kal und QC durchgeführt | Konnektor Gluc Kontakt def nach durch Hotline angeleitete Maßnahmen durch den Anwender - Konnektor ersetzt, Gluc neu membraniert Empf weiter nur..., Gluc Elektrode ersetzt Empf. bei 800 , Ref Elektrode wegen QC Fehler ELyte neu membraniert, Kontakt des Konnektors gereinigt, Kal und QC durchgeführt | SWO-144180 | No | Connector Gluc contact def after hotline-guided measures replaced by the user - connector, Gluc re-membranes Empf continues only..., Gluc electrode replaces Empf. at 800 , ref electrode re-logged due to QC error ELyte, contact of the connector cleaned, Kal and QC performed | Connector Gluc contact def after hotline-guided measures replaced by the user - connector, Gluc re-membranes Empf continues only..., Gluc electrode replaces Empf. at 800 , ref electrode re-logged due to QC error ELyte, contact of the connector cleaned, Kal and QC performed | Assisted Replacement | Electrode | Electrode | Glu
ABL825 FLEX | Konfiguration im Bereich des Bildschirms angepasst, Funktionskontrolle ok | mit Herr Döhler gesprochen, Töne müssen über Windows angepasst werden | SWO-144038 | No | Configuration adjusted in the area of the screen, function control ok | spoken to Mr. Döhler, sounds need to be adjusted via Windows | Setup | SW | User interface | Audio
AQT90 FLEX | Cleaned the Filter. Has to change the Needle wash unit. 903-329 : Needle wash unit, complete (Hct). | B | SWO-144151 | No | Cleaned the Filter. Has to change the Needle wash unit. 903-329 : Needle wash unit, complete (Hct). |  | Defect | HW | Aspiration Module | Needle Wash Unit
AQT90 FLEX | Fehlersuche, Bildschirmkabel aus der Halterung gerutscht, Gerät geöffnet, Kabelführung optimiert, im Serviceprogr. div. Einstellungen bzgl. d. Trocknung vorgenommen, QC i.O. | Fehlersuche; Bildschirmkabel aus der Halterung gerutscht; Gerät geöffnet; Kabelführung optimiert; im Serviceprogr. div. Einstellungen bzgl. d. Trocknung vorgenommen; QC i.O. | SWO-144030 | No | Troubleshooting; screen cable slipped out of the bracket; Device open; Cable routing optimized; in the service progr. div. Settings made with regard to drying; QC in the first | Troubleshooting; screen cable slipped out of the bracket; Device open; Cable routing optimized; in the service progr. div. Settings made with regard to drying; QC in the first | Defect | HW | Analyzer Control Module | Cable
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-144571 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-145619 | No | Replacement test card repair | Replacement test card repair
ABL90 FLEX | LP gewechselt | LP gewechselt | SWO-144314 | No | LP changed | LP changed | User error | SP | SP Expired | SP Expired
ABL90 |  | B | D1986131F80A75D9 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Removed blockage from hemolyser. | Removed blockage from hemolyser. | SWO-144460 | No | Removed blockage from hemolyser. | Removed blockage from hemolyser. | Cleaned | Fluidic pathway | Oxi module | Hemolyzer
ABL90 FLEX PLUS |  | B | D1986F00F843E4CD | Yes |  |  | Re classified RMED | PER Like upSWO
ABL90 FLEX PLUS | Gerät überprüft, Einlaßnadel gesäubert, Dichtung ausgetauscht | B | SWO-144411 | No | Device checked, inlet needle cleaned, seal replaced |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D1984AB5F846E826 | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
TCM4 | Electrodo trocado em garantia | B | D1987DBFF84F64C8 | No | Trocado electrode in guarantee
ABL90 FLEX | 更换配件903-183后修复 by jesson | 更换配件903-183后修复 by jesson | SWO-143552 | No | Repair by jesson after replacing accessory 903-183 | Repair by jesson after replacing accessory 903-183 | Defect | HW | Basic Unit | Sensor Interface PCB
AQT90 FLEX |  | B | SWO-144213 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
ABL815 FLEX | tHb Kalibration erneut durchgeführt, Messwert in Ordnung, ohne Beanstandung. Einsatz einer neuen pCO2 Elektrode. Empfindlichkeit stark außer Grenzen. Erneuter Einsatz der alten Elektrode. Messkammern der Glukose und pH Elektrode gereinigt und Elektrodenkontakte gebürstet. Unterer Elektrodendichring bei der pH Elektrode erneuert. Kalibrationswerte der pCO2 Elektrode in Ordnung. Neue Barcodes für die neue Software Version eingespielt. Kalibrationen und Qualitätskontrolle in Ordnung. | tHb Kalibration erneut durchgeführt, Messwert in Ordnung, ohne Beanstandung. Einsatz einer neuen pCO2 Elektrode. Empfindlichkeit stark außer Grenzen. Erneuter Einsatz der alten Elektrode. Messkammern der Glukose und pH Elektrode gereinigt und Elektrodenkontakte gebürstet. Unterer Elektrodendichring bei der pH Elektrode erneuert. Kalibrationswerte der pCO2 Elektrode in Ordnung. Neue Barcodes für die neue Software Version eingespielt. Kalibrationen und Qualitätskontrolle in Ordnung. | SWO-144428 | No | tHb calibration performed again, measured value in order, without complaint. Use of a new pCO2 electrode. Sensitivity strongly out of bounds. Renewed use of the old electrode. Measuring chambers of glucose and pH electrode cleaned and electrode contacts brushed. Lower electrode tight ring renewed at the pH electrode. Calibration values of the pCO2 electrode in order. New barcodes for the new software version. Calibrations and quality control in order. | tHb calibration performed again, measured value in order, without complaint. Use of a new pCO2 electrode. Sensitivity strongly out of bounds. Renewed use of the old electrode. Measuring chambers of glucose and pH electrode cleaned and electrode contacts brushed. Lower electrode tight ring renewed at the pH electrode. Calibration values of the pCO2 electrode in order. New barcodes for the new software version. Calibrations and quality control in order. | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL820 FLEX | found that tube had come disconnected. | reconnected tube and everything is fine | D198A263F810B16F | No | found that tube had come disconnected. | reconnected tube and everything is fine | Adjustment | HW | Basic Unit | Tubing | Unkonwn tube, Disconnect
ABL90 FLEX | 910-324缺货取消，改借910-298一个。 | 更换910-298一个正常。by jesson | SWO-144489 | No | 910-324 out of stock cancelled, instead of 910-298. | Replace 910-298 one normal. by jesson | Revisit required | HW | DMS_PC | Touch screen
ABL825 FLEX | checked power supply, checked boot tool USB, checked electrodes and membranes | removed DMS and replaced Hard drive, replaced PCO2 and Glucose elctrodes, replaced electrode membranes, removed Oxi module and replacd Oxi control board, installed boot bios, win7, registry file and 6.16 software | D19A5D9D6E61DE73 | Yes | checked power supply, checked boot tool USB, checked electrodes and membranes | removed DMS and replaced Hard drive, replaced PCO2 and Glucose elctrodes, replaced electrode membranes, removed Oxi module and replacd Oxi control board, installed boot bios, win7, registry file and 6.16 software | Re classified RMED | Installation
ABL90 FLEX | 更换910-298一个正常。by jesson | 更换910-298一个正常。by jesson | SWO-144492 | No | Replace 910-298 one normal. by jesson | Replace 910-298 one normal. by jesson | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX | 重新定标正常。by jesson | 重新定标正常。by jesson | SWO-144477 | No | Recalibration is normal. by jesson | Recalibration is normal. by jesson | Defect | Consumables | Test Kit | Cartridge, Cal
ABL80 FLEX BASIC |  | B | SWO-144595 | No
ABL80 OSM | Spectrum errors 581 on blood samples | Replaced Spectrometer Hemolyzer Ran THB | D19C90785A169FBD | No | Spectrum errors 581 on blood samples | Replaced Spectrometer Hemolyzer Ran THB
ABL90 FLEX | Débouchage système complète et le cassette. Nettoyage du joint d'intro. Maintenant tout les test sont OK Calibrations et CQ OK | Débouchage système complète et le cassette. Nettoyage du joint d'intro. Maintenant tout les test sont OK Calibrations et CQ OK | SWO-144921 | No | Complete system unblocking and cassette. Cleaning the intro joint. Now all the tests are OK Calibrations and CQ OK | Complete system unblocking and cassette. Cleaning the intro joint. Now all the tests are OK Calibrations and CQ OK | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | GTに従い流路洗浄で血液排出有。以降改善。 | GTに従い流路洗浄で血液排出有。以降改善。 | SWO-144956 | No | Blood discharge is available by flow path cleaning according to GT. Improvements since. | Blood discharge is available by flow path cleaning according to GT. Improvements since. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Client advised major gap between glucose readings between 2 analyzers. | OSPD found analyzer failing Calibration on Glucose parameter. Status above 3,000. Replaced sensor cassette, status normalized, client to run correlation on new samples. | D19C5FB05A3CA974 | No | Client advised major gap between glucose readings between 2 analyzers. | OSPD found analyzer failing Calibration on Glucose parameter. Status above 3,000. Replaced sensor cassette, status normalized, client to run correlation on new samples. | Assisted Replacement | SC | SC | Glu | Gap
ABL90 FLEX | Vérification dans le menu service + joints et connections Tout les test sont OK : appareille fonctionne J'ai constaté que un pack de solution était pas bien stocké (vertical). Ça peux cause les erreurs de homogénéité . C'est très importante pour gardé les SP dans le bon position! | Vérification dans le menu service + joints et connections Tout les test sont OK : appareille fonctionne J'ai constaté que un pack de solution était pas bien stocké (vertical). Ça peux cause les erreurs de homogénéité . C'est très importante pour gardé les SP dans le bon position! | SWO-144922 | No | Checking in the service menu - seals and connections All the tests are OK: device works I found that a solution pack was not well stored (vertical). It can cause the errors of homogeneity. This is very important to keep The SPs in the right position! | Checking in the service menu - seals and connections All the tests are OK: device works I found that a solution pack was not well stored (vertical). It can cause the errors of homogeneity. This is very important to keep The SPs in the right position! | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | Einlassdichtungen und Einlasssone erneuert. | Einlassdichtungen und Einlasssone erneuert. | SWO-145228 | No | Inlet seals and inlet tones are renewed. | Inlet seals and inlet tones are renewed. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Funktion OK. Durch Fehlzuordnung Hotline wurde der falsche Kunde besucht. | Funktion OK. Durch Fehlzuordnung Hotline wurde der falsche Kunde besucht. | SWO-144853 | Yes | Function OK. Due to misallocation hotline, the wrong customer was visited. | Function OK. Due to misallocation hotline, the wrong customer was visited. | Re classified RMED | error in Doc from Hotline
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-145627 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | SP | 184
AQT90 FLEX |  | B | SWO-144573 | No |  |  | Defect | HW | Cover and Chassis | Inlet Door | 4587
AQT90 FLEX | 更换903--138，636-536各一个正常。---jack.chu | 更换903--138，636-536各一个正常。---jack.chu | SWO-144483 | No | Replace 903-138,636-536 each normal. ---jack.chu | Replace 903-138,636-536 each normal. ---jack.chu | Defect | HW | Incubator module | PCB, Incubator | 2168
AQT90 FLEX | 更换910-339修复---jack.chu | 更换910-339修复---jack.chu | SWO-144470 | No | Replace 910-339 Repair ---jack.chu | Replace 910-339 Repair ---jack.chu | Defect | HW | Aspiration Module | Barcode reader, Inlet | 998
AQT90 FLEX |  | B | SWO-144564 | No |  |  | Defect | HW | Aspiration Module | PCB, Vacu vision
ABL800 FLEX | Performed protein removal & complete oximetry test then 2 point calibration & QC done within the range | Performed complete oximetry test & 2 point calibration & QC done values within the range | SWO-147666 | No | Performed protein removal & complete oximetry test then 2 point calibration & QC done within the range | Performed complete oximetry test & 2 point calibration & QC done values within the range | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL827 FLEX |  | B | D19C83735A61FC3F | Yes |  |  | Re classified RMED | Used analyzer Refurbished
ABL800 BASIC | Check and clean auto check scanner and barcode scanner,clean and check all tubings for any blockage,check all pump tubings,clean inlet Gasket,waste line,all chambers and un protein removal and done 2-point calibration which is OK,Check auto check scanner in which some places is working.Run auto check QC which is also OK and sample esults are satisfactory. | Check and clean auto check scanner and barcode scanner,clean and check all tubings for any blockage,check all pump tubings,clean inlet Gasket,waste line,all chambers and un protein removal and done 2-point calibration which is OK,Check auto check scanner in which some places is working.Run auto check QC which is also OK and sample esults are satisfactory. | SWO-145844 | No | Check and clean auto check scanner and barcode scanner,clean and check all tubings for any blockage,check all pump tubings,clean inlet Gasket,waste line,all chambers and un protein removal and done 2-point calibration which is OK,Check auto check scanner in which some places is working. Run auto check QC which is also OK and sample esults are satisfactory. | Check and clean auto check scanner and barcode scanner,clean and check all tubings for any blockage,check all pump tubings,clean inlet Gasket,waste line,all chambers and un protein removal and done 2-point calibration which is OK,Check auto check scanner in which some places is working. Run auto check QC which is also OK and sample esults are satisfactory. | Cleaned | HW | Auto Check Module | Barcode reader, AC
ABL825 FLEX | 更换电极修复 | 更换电极修复 | SWO-145615 | No | Replacement electrode repair | Replacement electrode repair | Assisted Replacement | Electrode | Electrode | pO2
ABL825 FLEX | 更换945-613一个正常。---jack.chu | 更换945-613一个正常。---jack.chu | SWO-145039 | No | Replace 945-613 one normal. ---jack.chu | Replace 945-613 one normal. ---jack.chu | Assisted Replacement | Electrode | Electrode | pO2
ABL835 FLEX | Intarmittant thb errors. | Repalced OXI module module PCB. Completed all service tests. | D19C6F605A3F1E48 | No | Intarmittant thb errors. | Repalced OXYGEN module Module PCB. Completed all service tests. | Defect | HW | Oxi module | Oxi PCB
ABL820 FLEX | I checked the fuses on the power supply and 1 of the 4A primary fuses was blown. I took one from the power supply that I had and it blew again. I put the good fuse in the power supply and put in the new power supply. I also moved the analyzer power cord to a different outlet on their UPS. it powered on. | Replaced power supply. | D19C2D905A2BC3B1 | No | I checked the fuses on the power supply and 1 of the 4A primary fuses was blown. I took one from the power supply that I had and it blew again. I put the good fuse in the power supply and put in the new power supply. I also moved the analyzer power cord to a different outlet on their UPS. it powered on. | Replaced power supply. | Defect | HW | Power supply | Power supply
ABL825 FLEX | COULD NOT IDENTIFY AIR LEAK. | REPLACED MEASURING CHAMBER, TUBINGS AND SWAPPED LIQ SENSORS | D19C61FE5A3F2A39 | No | COULD NOT IDENTIFY AIR LEAK. | REPLACED MEASURING CHAMBER, TUBINGS AND SWAPPED LIQ SENSORS | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | Revisión de fallo del analizador (presenta problemas en el lector de código de barras interno), limpieza de los contactos del fusible de la PCB Scan y se comprueba el correcto funcionamiento del mismo mediante QC ´s y muestras. | B | D19C76CE5A3644CC | No | Analyzer failure review (presents problems in the internal barcode reader); CLEANING the PCB Scan fuse contacts and checking the correct operation of the PCB scan by MEANS of QC's and samples. |  | Cleaned | HW | Storage Module | PCB, Cup transport - Front
AQT90 FLEX | Optische Einheit ausgebaut, alle cups entfernt, Abfallkassette XP-68 | Datenübertragung einrichten. | SWO-144896 | No | Optical unit removed, all cups removed, waste cartridge XP-68 | Set up data transfer. | Defect | Consumables | SP | Cup Stack
ABL80 FLEX Basic version |  | B | D19D8F0ED0076BE8 | No
ABL80 FLEX | Skiftet Kassette *4 og væske pakke *2 og skiftet slanger 905-896 | Skiftet Kassette *4 og væske pakke *2 og skiftet slanger 905-896 | SWO-145189 | No | Changed cassette * 4 and Liquid Package * 2 and changed tubing 905-896 | Changed cassette * 4 and Liquid Package * 2 and changed tubing 905-896
ABL90 FLEX | A készülék visszaszállítása és beüzemelése. Kalibrálás, QC mérés | A műszer elszállításra kerül felülvizsgálat miatt, helyszinen a készülékkel megegyező típus cserekészülék üzemel | SWO-145043 | No | Return and commissioning of the device. Calibration, QC measurement | The instrument is shipped for review, the same type of replacement is in place on the | Unknown | Analyzer replaced | Inlet module | Inlet Gasket
ABL90 FLEX | Performed manual fluid transport calibrations in service program, all passed ok | Performed manual fluid transport calibrations in service program, all passed ok | SWO-145093 | No | Performed manual fluid transport calibrations in service program, all passed ok | Performed manual fluid transport calibrations in service program, all passed ok | Defect | HW | Inlet module | Inlet Gasket | 1184
ABL90 FLEX | Replaced inlet gasket with new, left spare with customer | Replaced inlet gasket with new, left spare with customer | SWO-145092 | No | Replaced inlet gasket with new, left spare with customer | Replaced inlet gasket with new, left spare with customer | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Run through fluid transport manual calibration programs and all passed OK, flushed analyser and refitted SP | Run through fluid transport manual calibration programs and all passed OK, flushed analyser and refitted SP | SWO-145123 | No | Run through fluid transport manual calibration programs and all passed OK, flushed analyser and refitted SP | Run through fluid transport manual calibration programs and all passed OK, flushed analyser and refitted SP | Defect | HW | Inlet module | Inlet Gasket
ABL825 FLEX | 04 07 2016: REVISIÓN DE HISTÓRICOS, CAMBIO DE MEMBRANAS, TUBOS PERISTALTICAS, AGUJA DE ENTRADA. CALIBRACIONES Y CONTROLES. NO PASA CONTROL NIVEL 3. 07 07 2016: SUSTITUIR TUBOS EN MODULO OXIMETRÍA, AJUSTE DE LAMPARA, CALIBRACIÓN DE THB. CALIBRACIONES Y CONTROLES. NO PASA CONTROL NIVEL 3. 08 07 2016: REEMPLAZAR HEMOLIZADOR EN MODULO OXIMETRÍA. AJUSTE LAMPARA, CALIBRACIÓN THB. CALIBRACIONES Y CONTROLES. | 04 07 2016: REVISIÓN DE HISTÓRICOS, CAMBIO DE MEMBRANAS, TUBOS PERISTALTICAS, AGUJA DE ENTRADA. CALIBRACIONES Y CONTROLES. NO PASA CONTROL NIVEL 3. 07 07 2016: SUSTITUIR TUBOS EN MODULO OXIMETRÍA, AJUSTE DE LAMPARA, CALIBRACIÓN DE THB. CALIBRACIONES Y CONTROLES. NO PASA CONTROL NIVEL 3. 08 07 2016: REEMPLAZAR HEMOLIZADOR EN MODULO OXIMETRÍA. AJUSTE LAMPARA, CALIBRACIÓN THB. CALIBRACIONES Y CONTROLES. | D19DB8B3D022D214 | No | 04 07 2016: HISTORICAL REVIEW, MEMBRANE CHANGE, PERISTALTIC TUBOS, INPUT NEEDLE. CALIBRATIONS AND CONTROLS. NO LEVEL 3 CONTROL PASS. 07 07 2016: SUBSTITUTE TUBES IN OXIMETRY MODULE, LAMPARATION ADJUSTMENT, THB CALIBRATION. CALIBRATIONS AND CONTROLS. NO LEVEL 3 CONTROL PASS. 08 07 2016: REPLACE HEMOLIZER IN OXIMETRY MODULE. LAMPARA SETTING, THB CALIBRATION. CALIBRATIONS AND CONTROLS. | 04 07 2016: HISTORICAL REVIEW, MEMBRANE CHANGE, PERISTALTIC TUBOS, INPUT NEEDLE. CALIBRATIONS AND CONTROLS. NO LEVEL 3 CONTROL PASS. 07 07 2016: SUBSTITUTE TUBES IN OXIMETRY MODULE, LAMPARATION ADJUSTMENT, THB CALIBRATION. CALIBRATIONS AND CONTROLS. NO LEVEL 3 CONTROL PASS. 08 07 2016: REPLACE HEMOLIZER IN OXIMETRY MODULE. LAMPARA SETTING, THB CALIBRATION. CALIBRATIONS AND CONTROLS. | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Elektroden entnommen und gereinigt, Proteinentferner-Programm, Elyt.Elektr. membraniert, Modul-Ventile ausgebaut und gereinigt, Elyt.Modul Pumpenschlauch ersetzt, Einlaßdichtung ersetzt Pumpenkal., Kal. ok | Elektroden entnommen und gereinigt, Proteinentferner-Programm, Elyt.Elektr. membraniert, Modul-Ventile ausgebaut und gereinigt, Elyt.Modul Pumpenschlauch ersetzt, Einlaßdichtung ersetzt Pumpenkal., Kal. ok | SWO-145256 | No | Electrodes removed and cleaned, protein remover program, Elyt.Elektr. membraneed, module valves removed and cleaned, Elyt.Module pump hose replaced, inlet seal replaces pump scale, Kal. Ok | Electrodes removed and cleaned, protein remover program, Elyt.Elektr. membraneed, module valves removed and cleaned, Elyt.Module pump hose replaced, inlet seal replaces pump scale, Kal. Ok | User error | Electrode | Membrane | Reference
ABL835 FLEX |  | B | D19EB617461C4788 | No |  |  | Defect | HW | Inlet module | Inlet Gasket
ABL815 FLEX | Display erneuert, Display kalibriert, Kal u. QC i.O. HDD bereinigt | Display erneuert, Display kalibriert, Kal u. QC i.O. HDD bereinigt | SWO-145246 | No | Display renewed, display calibrated, Kal & QC i.O. HDD cleaned | Display renewed, display calibrated, Kal & QC i.O. HDD cleaned | Defect | HW | DMS_PC | Touch screen
ABL800 BASIC | Performed decontimation now its working fine. | Performed decontimation now its working fine. | SWO-145944 | No | Performed decontimation now its working fine. | Performed decontimation now its working fine. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL835 FLEX | Service-Pumpenkal., Proteinentferner, LS-Justage, Probemessungen | Service-Pumpenkal., Proteinentferner, LS-Justage, Probemessungen | SWO-145168 | No | Service pump scale, protein remover, LS-adjustment, sample measurements | Service pump scale, protein remover, LS-adjustment, sample measurements | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | Gerät ausgetauscht, AQT 90 838R0059N009 kalibriert und aufgestellt, QC's durchgeführt, | Datenübertragung einrichten. | SWO-145202 | No | Device replaced, AQT 90 838R0059N009 calibrated and installed, QC's performed, | Set up data transfer. | Revisit required | Analyzer replaced | Analyzer replaced | Analyzer replaced
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-145581 | No | Repair after replacing test card | Repair after replacing test card
ABL80 FLEX BASIC |  | B | SWO-145321 | No
ABL80 FLEX BASIC | cross checked with another analyser. CPU unit was not working. CPU-945 is required. | cross checked with another analyser. CPU unit was not working. CPU-945 is required. | SWO-146069 | No | cross checked with another analyser. CPU unit was not working. CPU-945 is required. | cross checked with another analyser. CPU unit was not working. CPU-945 is required.
ABL80 FLEX BASIC | 更换902-972修复---jack.chu | 更换902-972修复---jack.chu | SWO-144712 | No | Replace 902-972 Repair ---jack.chu | Replace 902-972 Repair ---jack.chu
ABL90 FLEX | LS justage durchgeführt, Pumpenkalibrierung nicht bestanden, Drucktest nicht bestanden, Lösungspacktest nicht bestanden. Es wird kein Druck im Gerät aufgebaut. Ventil ausgebaut und gereinigt. Tests erneuert durchgeführt - weiterhin kein Druckaufbau im Gerät. Fehlersuche durchgeführt - Einlassdichtung und Abfallschlauch undicht - ausgetauscht. Kalibriert und QC's durchgeführt. | LS justage durchgeführt, Pumpenkalibrierung nicht bestanden, Drucktest nicht bestanden, Lösungspacktest nicht bestanden. Es wird kein Druck im Gerät aufgebaut. Ventil ausgebaut und gereinigt. Tests erneuert durchgeführt - weiterhin kein Druckaufbau im Gerät. Fehlersuche durchgeführt - Einlassdichtung und Abfallschlauch undicht - ausgetauscht. Kalibriert und QC's durchgeführt. | SWO-145621 | No | LS adjustment performed, pump calibration failed, pressure test failed, solution pack test failed. No pressure is set up in the device. Valve removed and cleaned. Tests performed renewed - still no pressure build-up in the device. Troubleshooting performed - inlet seal and waste hose leaking - replaced. Calibrated and QC's performed. | LS adjustment performed, pump calibration failed, pressure test failed, solution pack test failed. No pressure is set up in the device. Valve removed and cleaned. Tests performed renewed - still no pressure build-up in the device. Troubleshooting performed - inlet seal and waste hose leaking - replaced. Calibrated and QC's performed. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Hemolyzer, Ventil und Ventilschlauch erneuert. | Hemolyzer, Ventil und Ventilschlauch erneuert. | SWO-145450 | No | Hemolyzer, valve and valve hose. | Hemolyzer, valve and valve hose. | Defect | HW | Oxi module | Hemolyzer
ABL90 | Unit disabled upon arrival.Unit wanting replacement of SP. Flushed system. | Replaced SC and SP. | D1A7CFAE80146A6D | No | Unit disabled upon arrival. Unit wanting replacement of SP. Flushed system. | Replaced SC and SP. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D19F6955462A4AC2 | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Analysator gespült - allgemeine Überprüfung - tHb KAL - Erläuterungen Gerätebeauftragter - Funktionsüberprüfung etc. ohne Beanstandung ! | Analysator gespült - allgemeine Überprüfung - tHb KAL - Erläuterungen Gerätebeauftragter - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-145408 | No | Analyzer rinsed - general check - tHb KAL - Explanations device officer - function check etc. without complaint ! | Analyzer rinsed - general check - tHb KAL - Explanations device officer - function check etc. without complaint ! | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS |  | B | D19F62A9463A4AAE | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Lösungspack aus Kundenbestand ausgetauscht, Anwender eingewiesen, Funktionskontrolle ok | Lösungspack aus Kundenbestand ausgetauscht, Anwender eingewiesen, Funktionskontrolle ok | SWO-145446 | No | Solution pack replaced from customer stock, user instructed, function control ok | Solution pack replaced from customer stock, user instructed, function control ok | Assisted Replacement | SP | SP | SP | Leak current
ABL90 FLEX PLUS | Fitted new power supply and fitted battery. | B | SWO-145400 | No | Fitted new power supply and fitted battery. |  | Defect | HW | Power Supply | Power Supply
TCM4 |  | B | SWO-144722 | No
ABL90 FLEX | 更换测试卡修复 | 更换测试卡修复 | SWO-145634 | No | Replacement test card repair | Replacement test card repair
ABL90 | CHEQUEO DE NEUMATICA, COMPROBACIÓN DE SONDA, GASKET, INLET GASKET, DESOBSTRUCCIÓN DE SISTEMA NEUMATICO, SUSTITUCIÓN DE SC TOMADO EN PRESTAMO DEL LABORATORIO, PASADOS TEST DE SERVICIO, CALIBRACIÓN Y ESTABILIZACIÓN. | DESOBSTRUCCIÓN DE SISTEMA NEUMÁTICO, SUSTITUCIÓN DE SC. REVISIÓN PREVENTIVA SEGÚN PROCEDIMIENTO | D19EF85446150541 | No | NEUMATIC CHECK, SONDA CHECK, GASKET, INLET GASKET, NEUMATIC SYSTEM DEOBSTRUCTION, SC SUBSTITUTE TAKEN IN LABORATORY PROVISION, SERVICE TEST PASS, CALIBRATION AND STABILIZATION. | DEOBSTRUCTION OF TIRE SYSTEM, SC SUBSTITUTE. PREEMPTIVE REVIEW AS PROCEDURE | Defect | HW | PM Kit | PM Kit | FT
AQT90 FLEX | 更换903-364后修复---jack.chu | 更换903-364后修复---jack.chu | SWO-144964 | No | Repair ---jack.chu after replacing 903-364 | Repair ---jack.chu after replacing 903-364 | Defect | HW | Process Module | Optical Unit | 124
ABL725 | Udskiftet hemolyzeren, opjusteret og kørt komplet check - fejler på neon lamp, udskiftet PCB'et men vil ikke tage softwaren. Sagen overgivet til ALL RASPE. | Udskiftet hemolyzeren, opjusteret og kørt komplet check - fejler på neon lamp, udskiftet PCB'et men vil ikke tage softwaren. Sagen overgivet til ALL RASPE. | SWO-145658 | No | Replaced the Hemolyzeren, upgraded and run complete check-fails on the neon lamp, replaced the PCB but will not take the software. The case surrendered to ALL RASPE. | Replaced the Hemolyzeren, upgraded and run complete check-fails on the neon lamp, replaced the PCB but will not take the software. The case surrendered to ALL RASPE.
ABL835 FLEX |  | B | D19F669846064ABC | No |  |  | Assisted Replacement | Electrode | Electrode | Na+
ABL825 FLEX | Jeg har udskiftet AutoCheck modulet. Det udskiftet modul er testet og fundet ok. Der er kørt 1 og 2 punkts kal. Der er Kørt QCer. Alt er OK. | Jeg har udskiftet AutoCheck modulet. Det udskiftet modul er testet og fundet ok. Der er kørt 1 og 2 punkts kal. Der er Kørt QCer. Alt er OK. | SWO-145437 | No | I replaced the Auto Check module. The replaced module is tested and found OK. 1 and 2 point Kal have been run. QCer has been run. Everything is OK. | I replaced the Auto Check module. The replaced module is tested and found OK. 1 and 2 point Kal have been run. QCer has been run. Everything is OK. | Defect | HW | Auto Check Module | FAN | 735
ABL800 BASIC | Replaced lactate membrane | Replaced lactate membrane | SWO-147684 | No | Replaced lactate membrane | Replaced lactate membrane | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Gerät geprüft, Inlet Leihmodul installiert, Referenzelektrode neu membraniert, Funktionskontrolle ok | Gerät geprüft, Inlet Leihmodul installiert, Referenzelektrode neu membraniert, Funktionskontrolle ok | SWO-145533 | No | Device tested, Inlet rental module installed, reference electrode re-membraneed, function control ok | Device tested, Inlet rental module installed, reference electrode re-membraneed, function control ok | Assisted Replacement | Electrode | Membrane | Reference
ABL800 BASIC | Found inlet gasket leak, replaced the inlet gasket restarted instrument found instrument is functional. | Found inlet gasket leak, replaced the inlet gasket restarted instrument found instrument is functional. | SWO-145898 | No | Found inlet gasket leak, replaced the inlet gasket restarted instrument found instrument is functional. | Found inlet gasket leak, replaced the inlet gasket restarted instrument found instrument is functional. | Defect | HW | Inlet module | Inlet Gasket
ABL825 FLEX | Startphase QC und Kalibration. | Startphase QC und Kalibration. | SWO-145681 | No | Start phase QC and calibration. | Start phase QC and calibration. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Proteinentfernung durchgeführt - Messkammer gereinigt Lac Elektrode gemäß Reinigungsverfahren gesäubert Glu-, Lac- und Ref.-Membran erneuert, tHb kalibriert Gerät kalibriert und überprüft | Proteinentfernung durchgeführt - Messkammer gereinigt Lac Elektrode gemäß Reinigungsverfahren gesäubert Glu-, Lac- und Ref.-Membran erneuert, tHb kalibriert Gerät kalibriert und überprüft | SWO-145371 | No | Protein removal performed - measuring chamber cleaned lac electrode according to cleaning method cleaned glu, lac and ref. membrane renewed, tHb calibrated device calibrated and checked | Protein removal performed - measuring chamber cleaned lac electrode according to cleaning method cleaned glu, lac and ref. membrane renewed, tHb calibrated device calibrated and checked | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL835 FLEX | Proteinentfernung durchgeführt - Messkammer gereinigt Lac Elektrode gemäß Reinigungsverfahren gesäubert Glu-, Lac- und Ref.-Membran erneuert, tHb kalibriert Gerät kalibriert und überprüft | Proteinentfernung durchgeführt - Messkammer gereinigt Lac Elektrode gemäß Reinigungsverfahren gesäubert Glu-, Lac- und Ref.-Membran erneuert, tHb kalibriert Gerät kalibriert und überprüft | SWO-145370 | No | Protein removal performed - measuring chamber cleaned lac electrode according to cleaning method cleaned glu, lac and ref. membrane renewed, tHb calibrated device calibrated and checked | Protein removal performed - measuring chamber cleaned lac electrode according to cleaning method cleaned glu, lac and ref. membrane renewed, tHb calibrated device calibrated and checked | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL800 BASIC | ｵｰﾄﾁｪｯｸﾓｼﾞｭｰﾙ（代替品）へ入替え作業。 | B | SWO-149956 | No | The replacement work to the auto check module (replacement).
AQT90 FLEX |  | B | D19EC348462B7260 | No |  |  | Defect | HW | Process Module | Cable | Reconnect
AQT90 FLEX | Storage wheel PCB replaced, analyser swithed on but error comes. Check in service program found Error 13 for punch out. Incubator wheel open and clean the sensor. Incubator wheel cleaned and Process wheel cleaned. Reference cup checked and cleaned. Analyzer On and Install new Cartridge - System check done. Test perform for PCT, found well. Analyser Serviced properly and calibrated. Handed over in a fine working condition. | Storage wheel PCB replaced, analyser swithed on but error comes. Check in service program found Error 13 for punch out. Incubator wheel open and clean the sensor. Incubator wheel cleaned and Process wheel cleaned. Reference cup checked and cleaned. Analyzer On and Install new Cartridge - System check done. Test perform for PCT, found well. Analyser Serviced properly and calibrated. Handed over in a fine working condition. | SWO-145441 | No | Storage wheel PCB replaced, analyser swithed on but error comes. Check in service program found Error 13 for punch out. Incubator wheel open and clean the sensor. Incubator wheel cleaned and Process wheel cleaned. Reference cup checked and cleaned. Analyzer On and Install new Cartridge - System check done. Test perform for PCT, found well. Analyser Serviced properly and calibrated. Handed over in a fine working condition. | Storage wheel PCB replaced, analyser swithed on but error comes. Check in service program found Error 13 for punch out. Incubator wheel open and clean the sensor. Incubator wheel cleaned and Process wheel cleaned. Reference cup checked and cleaned. Analyzer On and Install new Cartridge - System check done. Test perform for PCT, found well. Analyser Serviced properly and calibrated. Handed over in a fine working condition. | Defect | HW | Process Module | PCB, Cup transport - Front
AQT90 FLEX | Testmessungen durchgeführt- kein Erfolg! Abdeckung des Storage Wheel entfernt, lockeren Abstandsbolzen unterlegt und befestigt, da zuvor schleifende Geräusche an der Abdeckung zu bemerken waren. Needle Assembly ausgebaut, ließ sich sehr leicht aufschrauben, Dichtung und Sitz gereinigt und wieder eingebaut. Einlaufverfahren und Systemtest + 5QCs fehlerfrei durchgeführt . | Testmessungen durchgeführt- kein Erfolg! Abdeckung des Storage Wheel entfernt, lockeren Abstandsbolzen unterlegt und befestigt, da zuvor schleifende Geräusche an der Abdeckung zu bemerken waren. Needle Assembly ausgebaut, ließ sich sehr leicht aufschrauben, Dichtung und Sitz gereinigt und wieder eingebaut. Einlaufverfahren und Systemtest + 5QCs fehlerfrei durchgeführt . | SWO-145853 | No | Test measurements carried out - no success! Removed the Storage Wheel cover, loose spacer bolts underlaid and fastened, as previously grinding noises were noticeable on the cover. Needle Assembly removed, could be screwed very easily, seal and seat cleaned and reinstalled. Inlet procedure and system test + 5QCs performed error-free. | Test measurements carried out - no success! Removed the Storage Wheel cover, loose spacer bolts underlaid and fastened, as previously grinding noises were noticeable on the cover. Needle Assembly removed, could be screwed very easily, seal and seat cleaned and reinstalled. Inlet procedure and system test + 5QCs performed error-free. | Adjustment | HW | Storage Module | Storage wheel
ABL80 FLEX CO-OX | Förderung von Lösungen nicht möglich - Fließweg verstopft, Fließweg überprüft - ein Schlauchstück ist geknickt, sodass keine Lösung mehr durch gefördert werden kann - Schlauch ausgetauscht. Sensorkassette neu initialisiert. Kalibriert und QC's durchgeführt. | Frau Schnick möchte einen Vorortservice. | SWO-145915 | No | Promotion of solutions not possible - flow path clogged, flow path checked - a hose piece is bent, so that no solution can be conveyed by - hose replaced. sensor cartridge is reinitialized. Calibrated and QC's performed. | Ms. Schnick wants a suburban service.
ABL80 FLEX BASIC | temperature was out of range. Kept the cassettes and analyser in cool and reinitiated. calibration passed with lactate. | temperature was out of range. Kept the cassettes and analyser in cool and reinitiated. calibration passed with lactate. | SWO-145979 | No | temperature was out of range. Kept the cassettes and analyser in cool and reinitiated. calibration passed with lactate. | temperature was out of range. Kept the cassettes and analyser in cool and reinitiated. calibration passed with lactate.
ABL90 FLEX | Inlet gasket vervangen. | Inlet gasket vervangen. | SWO-145990 | No | Inlet gasket vervangen. | Inlet gasket vervangen. | Defect | HW | Inlet module | Inlet Gasket
ABL90 | The analyzer is failing the pressure test. Specifically, the vacuum value, which normally reads between 0.25-0.75. The vacuum value is reading 0.91 | Replaced the inlet gasket and performed successful pressure tests. Values are observed near 0.47. Analyzer re-initialized without error. A credit claim will be submitted on the customer's behalf. The analyzer prompted replacement of a recently installed solution pack. Analyzer is within acceptable operating specifications. | D1A01B1FBC43C22E | No | The analyzer is failing the pressure test. Specifically, the vacuum value, which normally reads between 0.25-0.75. The vacuum value is reading 0.91 | Replaced the inlet gasket and performed successful pressure tests. Values are observed near 0.47. Analyzer re-initialized without error. A credit claim will be submitted on the customer's behalf. The analyzer prompted replacement of a recently installed solution pack. Analyzer is within acceptable operating specifications. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Replaced inlet gasket, inlet probe and inlet connector gasket. | Replaced inlet gasket, inlet probe and inlet connector gasket. | SWO-145846 | No | Replaced inlet gasket, inlet probe and inlet connector gasket. | Replaced inlet gasket, inlet probe and inlet connector gasket. | Defect | HW | Inlet module | Inlet Gasket | & probe + connector gasket
ABL90 |  | B | D19FEEB1BC44B0AB | No |  |  | Setup | sw | His/Lis | His/Lis
AQT90 FLEX |  | 收到备件外包装是886-639，内部货物与上次借用886-638一致，无法使用。需要重新借件。---jack.chu | SWO-145626 | No |  | The package received for spare parts is 886-639, and the internal goods are consistent with the previous loan of 886-638 and cannot be used. A re-borrowing is required. ---jack.chu | Defect | HW | Cover and Chassis | Inlet Door
ABL827 FLEX | Tried cleaning electrode and different lot of membranes-No change | Replaced lactate electrode | D1A028D3BC391A8E | No | Tried cleaning electrode and different lot of membranes-No change | Replaced lactate electrode | Assisted Replacement | Electrode | Electrode | Lac
ABL827 FLEX | Tried cleaning the electrode and different lot of membranes-No change | Replaced lactate electrode | D1A025A1BC491A85 | No | Tried cleaning the electrode and different lot of membranes-No change | Replaced lactate electrode | Assisted Replacement | Electrode | Electrode | Lac
ABL800 BASIC | Clean pco2 chamber chane membrane prefrom QC all results are now ok. | Clean pco2 chamber chane membrane prefrom QC all results are now ok. | SWO-145832 | No | Clean pco2 chamber chane membrane prefrom QC all results are now ok. | Clean pco2 chamber chane membrane prefrom QC all results are now ok. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL820 FLEX | 比較メンブラン交換。CO2メンブランにピンホールあり交換。 | 比較メンブラン交換。CO2メンブランにピンホールあり交換。 | SWO-146204 | No | Comparison membrane replacement. There is a pinhole in THE CO2 membrane and it is exchanged. | Comparison membrane replacement. There is a pinhole in THE CO2 membrane and it is exchanged. | Assisted Replacement | Electrode | Membrane | pCO2 | Pinholes
ABL805 FLEX | AutoCheck Deckel testweise mit Desinfektionstuch, alkoholischer Desinfektionslösung, =R= Reinigungslösung sowie mit Hypochloridlösung versucht zu entfärben, ohne Erfolg, daher Deckel ersetzt. | AutoCheck Deckel testweise mit Desinfektionstuch, alkoholischer Desinfektionslösung, =R= Reinigungslösung sowie mit Hypochloridlösung versucht zu entfärben, ohne Erfolg, daher Deckel ersetzt. | SWO-145814 | Yes | AutoCheck lid tested with disinfectant cloth, alcoholic disinfectant solution, =R= cleaning solution as well as with hypochloride solution tries to discolor, without success, therefore lid replaced. | AutoCheck lid tested with disinfectant cloth, alcoholic disinfectant solution, =R= cleaning solution as well as with hypochloride solution tries to discolor, without success, therefore lid replaced. | Re classified RMED | Cosmetic issue
ABL800 BASIC | Check inlet Gasket and found clot in Gasket,clean and flush the Gasket and check for QC which is OK then run some sample which is also Ok without query. | Check inlet Gasket and found clot in Gasket,clean and flush the Gasket and check for QC which is OK then run some sample which is also Ok without query. | SWO-145847 | No | Check inlet Gasket and found clot in Gasket,clean and flush the Gasket and check for QC which is OK then run some sample which is also Ok without query. | Check inlet Gasket and found clot in Gasket,clean and flush the Gasket and check for QC which is OK then run some sample which is also Ok without query. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL800 BASIC | perform complete oximetry calibration. tHb calibration now machine is ok. | perform complete oximetry calibration. tHb calibration now machine is ok. | SWO-145830 | No | perform complete oximetry calibration. tHb calibration now machine is ok. | perform complete oximetry calibration. tHb calibration now machine is ok. | Adjustment | HW | Oxi module | Lamp unit
ABL800 BASIC | We need to perform tHb calibration. | We need to perform tHb calibration. | SWO-145833 | Yes | We need to perform tHb calibration. | We need to perform tHb calibration. | Re classified RMED | Double entry
ABL825 FLEX | Error 331. Found very minor leak around pH electrode. | Replace seal ring for pH electrode. Replace solution pump tubing. Check for leaks. | D1A01EB6BC5F6EA3 | No | Error 331. Found very minor leak around pH electrode. | Replace seal ring for pH electrode. Replace solution pump tubing. Check for leaks. | Assisted Replacement | Electrode | Electrode | Sealing Ring
ABL835 FLEX |  | B | D1A033E5BC515EC4 | No |  |  | Adjustment | HW | Inlet module | Probe calibration
AQT90 FLEX | Replaced the Needle wash unit. 903-329 : Needle wash unit, complete (Hct). | Replaced the Needle wash unit. 903-329 : Needle wash unit, complete (Hct). | SWO-145796 | No | Replaced the Needle wash unit. 903-329 : Needle wash unit, complete (Hct). | Replaced the Needle wash unit. 903-329 : Needle wash unit, complete (Hct). | Defect | HW | Aspiration Module | Needle Wash Unit
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-146543 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC | Sensor Board Replaced. Zeroing Calibration Done. PM Service done. Analyser Serviced and calibrated Properly. QC performed and found all well with in range. Stand by analyser couriered to Indore Back. Analyser handed over in a fine working condition. | Sensor Board Replaced. Zeroing Calibration Done. PM Service done. Analyser Serviced and calibrated Properly. QC performed and found all well with in range. Stand by analyser couriered to Indore Back. Analyser handed over in a fine working condition. | SWO-145910 | No | Sensor Board Replaced. Zeroing Calibration Done. PM Service done. Analyser Serviced and calibrated Properly. QC performed and found all well with in range. Stand by analyser couriered to Indore Back. Analyser handed over in a fine working condition. | Sensor Board Replaced. Zeroing Calibration Done. PM Service done. Analyser Serviced and calibrated Properly. QC performed and found all well with in range. Stand by analyser couriered to Indore Back. Analyser handed over in a fine working condition.
ABL80 FLEX BASIC | Tausch des Netzteils. Analysator boothet wieder normal. Funktion ohne Beanstandung. Alter Lösungspack verfallen, erneuert. Kalibration und Qualitätskontrolle in Ordnung. | Tausch des Netzteils. Analysator boothet wieder normal. Funktion ohne Beanstandung. Alter Lösungspack verfallen, erneuert. Kalibration und Qualitätskontrolle in Ordnung. | SWO-146136 | No | Replacing the power supply. Analyzer boothet again normal. Function without complaint. Old solution pack forfeited, renewed. Calibration and quality control in order. | Replacing the power supply. Analyzer boothet again normal. Function without complaint. Old solution pack forfeited, renewed. Calibration and quality control in order.
ABL90 | Checked for paper jams and none was found. | Replaced printer. | D1A15D22322ED846 | No | Checked for paper jams and none was found. | Replaced printer. | Defect | HW | Printer module | Printer module
ABL90 FLEX | Lösungspack erneut ersetzt, Alle Schläuche erstzt, Anschlüsse gespült, Test mit neuem SP erfolgreich, Neustart, Kalibration, QC Kontrolle, QC Fehler vorhanden daher in der Ambulanz bescheid gesagt das evtl. die Sensorkassette ersetzt werden muss. | Lösungspack erneut ersetzt, Alle Schläuche erstzt, Anschlüsse gespült, Test mit neuem SP erfolgreich, Neustart, Kalibration, QC Kontrolle, QC Fehler vorhanden daher in der Ambulanz bescheid gesagt das evtl. die Sensorkassette ersetzt werden muss. | SWO-146164 | No | Solution pack replaced again, All hoses rupts, connections flushed, test with new SP successful, restart, calibration, QC control, QC error present therefore in the ambulance decision said that possibly the sensor cartridge needs to be replaced. | Solution pack replaced again, All hoses rupts, connections flushed, test with new SP successful, restart, calibration, QC control, QC error present therefore in the ambulance decision said that possibly the sensor cartridge needs to be replaced. | Assisted Replacement | SP | SP | Flow selector | Manifold
ABL90 FLEX PLUS | Collected and delivered Sensor cassette to customer. | Collected and delivered Sensor cassette to customer. | SWO-146367 | Yes | Collected and delivered Sensor cassette to customer. | Collected and delivered Sensor cassette to customer. | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | LP durch Anwender erneuert. Fehler überprüft, Datenlogs erstellt, defekten LP abgezogen. | LP durch Anwender erneuert. Fehler überprüft, Datenlogs erstellt, defekten LP abgezogen. | SWO-146178 | No | LP renewed by users. Error checked, data logs created, defective LP subtracted. | LP renewed by users. Error checked, data logs created, defective LP subtracted. | Assisted Replacement | SP | SP | Inconsistante database
TCM4 |  | 更换945-377后报气瓶消耗快。需要再借902-778一个。---jack.chu | SWO-145874 | No |  | After replacing the 945-377, the cylinder is consumed quickly. Another 902-778 is required. ---jack.chu
ABL825 FLEX | PC Modul ausgebaut Drucker erneuert, Modul eingebaut Gerät überprüft | PC Modul ausgebaut Drucker erneuert, Modul eingebaut Gerät überprüft | SWO-145998 | No | PC module removed printer renewed, module built-in device checked | PC module removed printer renewed, module built-in device checked | Defect | HW | Printer module | Printer head
ABL815 FLEX |  | B | D1A4F2EE940D01F2 | No |  |  | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX |  | B | SWO-145994 | No |  |  | Assisted Replacement | Electrode | Electrode | pH
ABL800 BASIC | Changed membranes for Ca and Lac. Then given 1 point and 2 point calibration. Now instrument is working ok. | Changed membranes for Ca and Lac. Then given 1 point and 2 point calibration. Now instrument is working ok. | SWO-145986 | No | Changed membranes for Ca and Lac. Then given 1 point and 2 point calibration. Now instrument is working ok. | Changed membranes for Ca and Lac. Then given 1 point and 2 point calibration. Now instrument is working ok. | Assisted Replacement | Electrode | Membrane | Lac | To replace menbrane
AQT90 FLEX | Geräteabholung für Reparatur in GS17 | Kundengerät nach GS17 Reparatur zurück zum Kunden. Systemüberprüfung und QC o.k. | SWO-148367 | No | Device pickup for repair in GS17 | Customer device after GS17 repair back to the customer. System check and QC ok | LEH Mail |  |  | Not Cat. | LEH Mail
ABL90 FLEX | LP gewechselt | LP gewechselt | SWO-146341 | No | LP changed | LP changed | User error | SP empty | SP empty | SP empty
ABL800 BASIC | replaced the flat membrane for flat membrane valve. did calibration, QC and samples all results found with in the range. | replaced the flat membrane for flat membrane valve. did calibration, QC and samples all results found with in the range. | SWO-146910 | No | replaced the flat membrane for flat membrane valve. did calibration, QC and samples all results found with in the range. | replaced the flat membrane for flat membrane valve. did calibration, QC and samples all results found with in the range. | Defect | HW | pH/BG Module | Valve, air | Membrane
ABL827 FLEX | Amanda said the pCO2 will not pass the calibrations. She has changed the membrane 3 times. It is failing the sensitivity and drift. I asked her to make sure it had no bubbles in the membrane. I suggested replacing the gases to see if it would help. I called back and it had no change. I told her to just leave it and let it cal and I checked back later with no change. I suggested a new electrode. They will need to order one Monday. | No correction. | D1A431DA1E53C3D0 | Yes | Amanda said the pCO2 will not pass the calibrations. She has changed the membrane 3 times. It is failing the sensitivity and drift. I asked her to make sure it had no bubbles in the membrane. I suggested replacing the gases to see if it would help. I called back and it had no change. I told her to just leave it and let it cal and I checked back later with no change. I suggested a new electrode. They will need to order one Monday. | No correction. | Re classified RMED | Phone support
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-147225 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | pCO2
ABL80 FLEX BASIC | Replace the sample probe from old sc and reinstall and calibrate and give again proper training for Installation of SC & SP and sample processing to the staff. | Replace the sample probe from old sc and reinstall and calibrate and give again proper training for Installation of SC & SP and sample processing to the staff. | SWO-147881 | No | Replace the sample probe from old sc and reinstall and calibrate and give again proper training for Installation of SC & SP and sample processing to the staff. | Replace the sample probe from old sc and reinstall and calibrate and give again proper training for Installation of SC & SP and sample processing to the staff.
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-150662 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX | Gerät überprüft, Kal Daten und QC Ergebnisse kontrolliert, Vergleichsmessungen mit dem ABL 90 Station 14 ohne Beanstandung | Gerät überprüft, Kal Daten und QC Ergebnisse kontrolliert, Vergleichsmessungen mit dem ABL 90 Station 14 ohne Beanstandung | SWO-146559 | No | Device checked, Kal data and QC results checked, comparison measurements with the ABL 90 Station 14 without complaint | Device checked, Kal data and QC results checked, comparison measurements with the ABL 90 Station 14 without complaint | No fault found | Measurement | Measurement | Measurement
ABL90 FLEX |  | B | SWO-146284 | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-146315 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Lac | 131
AQT90 FLEX |  | B | SWO-146282 | No |  |  | Defect | HW | Process Module | Sensor, safe | 5784
ABL800 BASIC | LS test & pump calibration pass. check for leakage. Replace the inlet gasket & pump tubing . replace the needle tubing also. still fail to aspirate.check for check the fluidic valve. replace the rinse & air valve still fail to aspirate. | LS test & pump calibration pass. check for leakage. Replace the inlet gasket & pump tubing . replace the needle tubing also. still fail to aspirate.check for check the fluidic valve. replace the rinse & air valve still fail to aspirate. | SWO-146521 | No | LS test & pump calibration pass. check for leakage. Replace the inlet gasket & pump tubing . replace the needle tubing also. still fail to aspirate.check for check the fluidic valve. replace the rinse & air valve still fail to aspirate. | LS test & pump calibration pass. check for leakage. Replace the inlet gasket & pump tubing . replace the needle tubing also. still fail to aspirate.check for check the fluidic valve. replace the rinse & air valve still fail to aspirate. | Revisit required | HW | Inlet module | Unknown
ABL835 FLEX | 更换902-844一个 by jesson | 更换902-844一个 by jesson | SWO-146224 | Yes | Replace 902-844 a by jesson | Replace 902-844 a by jesson | Re classified RMED | OOB
ABL827 FLEX | Skiftet referenceelektrodeconnectoren og bestilt en referenceelektrode, slutafprøvet. | Skiftet referenceelektrodeconnectoren og bestilt en referenceelektrode, slutafprøvet. | SWO-146481 | No | Switched Referenceelektrodeconnectoren and ordered a reference electrode, final tested. | Switched Referenceelektrodeconnectoren and ordered a reference electrode, final tested. | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX |  | B | D1A51BCE9461FC61 | No |  |  | Defect | HW | Inlet module | Friction detector unit
ABL825 FLEX | Screen Froze while customer was trying to run a sample. It has happened twice in the past couple weeks. | Reinstalled WIN 7 and Roma 6.16. Customer to observe over the next couple weeks if the screen freeze issue comes again. Verified SEND key. | D1A6606D0A5A8016 | No | Screen Froze while customer was trying to run a sample. It has happened twice in the past couple weeks. | Reinstalled WIN 7 and Roma 6.16. Customer to observe over the next couple weeks if the screen freeze issue comes again. Verified SEND key. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Computer board unresponsive. Could not get anything on startup. | Replaced CPU unitl complete. Successfully loaded WIN 7 and 6.16 software. | D1A548AB94354738 | No | Computer board unresponsive. Could not get anything on startup. | Replaced CPU unitl complete. Successfully loaded WIN 7 and 6.16 software. | Defect | HW | DMS_PC | CPU
ABL80 CO-OX |  | B | D1A6CABD0A25EF94 | No
ABL80 CO-OX | Customer had this problem. They got a replacement sensor and it has the same error message. Analyzer was installed 2 months ago | Replaced analog PCB and sensor connector PCB. Performed zero calibration. Problem remains. | D1A65AE10A4BB1D5 | No | Customer had this problem. They got a replacement sensor and it has the same error message. Analyzer was installed 2 months ago | Replaced analog PCB and sensor connector PCB. Performed zero calibration. Problem remains.
AQURE | お客様の要望により、赤点滅ではなく赤点灯に変更。 | お客様の要望により、赤点滅ではなく赤点灯に変更。 | SWO-146887 | Yes | At the request of the customer, it was changed to red light instead of flashing red. | At the request of the customer, it was changed to red light instead of flashing red.
ABL90 FLEX | Műszer hypós tisztítása, Folyadékáramlási tesztek, Felszívótű csere, Páciens minták sorozatos tesztje (7 egymás követő mérés rendben volt), Kalibrálás, QC mérés | B | SWO-146857 | No | Hypós cleaning of the instrument, liquid flow tests, absorption needle replacement, sequential test of patterns (7 consecutive measurements were okay), calibration, QC measurement |  | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Solutionpack erneut aktiviert, Neustart nun erfolgreich. | Solutionpack erneut aktiviert, Neustart nun erfolgreich. | SWO-146854 | No | Solutionpack reactivated, restart now successful. | Solutionpack reactivated, restart now successful. | Assisted Replacement | SP | SP | CAL3 empty
ABL90 FLEX | Da mehrmals Lösungspack und Sensor getauscht wurden,Test mit Simulator tester durchgeführt, alle Tests im Service ok, complete pO2 test erfolgreich, nochmals Solutionpack und Sensorkassette ersetzt. Neustart, Kalibration, QC Kontrolle | Da mehrmals Lösungspack und Sensor getauscht wurden,Test mit Simulator tester durchgeführt, alle Tests im Service ok, complete pO2 test erfolgreich, nochmals Solutionpack und Sensorkassette ersetzt. Neustart, Kalibration, QC Kontrolle | SWO-146853 | No | Since the solution pack and sensor were exchanged several times, test with simulator tester was carried out, all tests in the service ok, complete pO2 test successfully, again solutionpack and sensor cassette replaced. Restart, Calibration, QC Control | Since the solution pack and sensor were exchanged several times, test with simulator tester was carried out, all tests in the service ok, complete pO2 test successfully, again solutionpack and sensor cassette replaced. Restart, Calibration, QC Control | Assisted Replacement | SC | SC | pO2
ABL90 FLEX | 技術部にて、調整。 | 技術部にて、調整。 | SWO-149073 | No | In the technical department, adjustment. | In the technical department, adjustment. | Assisted Replacement | SP | SP | FT
ABL90 FLEX | Nach SC Wechsel pCO2 Empf. < 10%. SC nochmals ersetzt Reklamationsausdrucke an RDEU gemailt. | Nach SC Wechsel pCO2 Empf. < 10%. SC nochmals ersetzt Reklamationsausdrucke an RDEU gemailt. | SWO-146949 | Yes | After SC change pCO2 empf. • 10%. SC again replaces complaint expressions to RDEU emailed. | After SC change pCO2 empf. • 10%. SC again replaces complaint expressions to RDEU emailed. | Re classified RMED | Phone support
ABL90 FLEX | SC Wechsel < 4h, Sytem kalibriert und QC wiederholt. Systemzeit eingestellt, Zeitzone konfiguriert. Anwender hatte SC bereits gewechselt, danach pCO2 Probleme. Reklamationsausdruck an RDEU gemailt. | SC Wechsel < 4h, Sytem kalibriert und QC wiederholt. Systemzeit eingestellt, Zeitzone konfiguriert. Anwender hatte SC bereits gewechselt, danach pCO2 Probleme. Reklamationsausdruck an RDEU gemailt. | SWO-146948 | No | SC Change of 4h, system calibrated and QC repeated. System time set, time zone configured. User had already changed SC, then pCO2 problems. Complaint statement sent to RDEU. | SC Change of 4h, system calibrated and QC repeated. System time set, time zone configured. User had already changed SC, then pCO2 problems. Complaint statement sent to RDEU. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS |  | B | D1A669D50A535FD3 | Yes |  |  | Re classified RMED | SWU
ABL825 FLEX | Checked halogen lamp voltage and found it was reading near the upper limit. | Adjusted halogen lamp voltage. | D1A640E40A1B2724 | No | Checked halogen lamp voltage and found it was reading near the upper limit. | Adjusted halogen lamp voltage. | Adjustment | HW | Oxi module | Lamp unit
ABL835 FLEX | Lüfter getauscht, Kal | Lüfter getauscht, Kal | SWO-146849 | No | Fan replaced, Kal | Fan replaced, Kal | Defect | HW | Auto Check Module | FAN | Noise
AQT90 FLEX | Storage Wheel demontiert, Kugellager gereinigt. Nach Neustart Fehler mit der HCT Einheiten aufgetreten, HCT Einheiten demontiert und gereinigt, Flachbandkabel korrekt eingesteckt. Systemreinigung durchgeführt. Zeitzone korrigiert. LQC Barcodes wurden nicht erkannt, internen BCR demontiert und gereinigt, LQCs i.O. Wartung notwendig. | Storage Wheel demontiert, Kugellager gereinigt. Nach Neustart Fehler mit der HCT Einheiten aufgetreten, HCT Einheiten demontiert und gereinigt, Flachbandkabel korrekt eingesteckt. Systemreinigung durchgeführt. Zeitzone korrigiert. LQC Barcodes wurden nicht erkannt, internen BCR demontiert und gereinigt, LQCs i.O. Wartung notwendig. | SWO-146819 | No | Storage Wheel dismantled, ball bearing cleaned. After restart error occurred with the HCT units, HCT units disassembled and cleaned, ribbon cable plugged in correctly. system cleaning. Time zone corrected. LQC barcodes were not detected, internal BCR dismantled and cleaned, LQCs i.O. maintenance necessary. | Storage Wheel dismantled, ball bearing cleaned. After restart error occurred with the HCT units, HCT units disassembled and cleaned, ribbon cable plugged in correctly. system cleaning. Time zone corrected. LQC barcodes were not detected, internal BCR dismantled and cleaned, LQCs i.O. maintenance necessary. | Cleaned | HW | Storage Module | Storage wheel | 151
AQT90 FLEX |  | B | D1A6C9BC0A18EF91 | No |  |  | Cleaned | HW | Storage Module | Sensor, storage wheel
AQT90 FLEX | Reparatur Ersatzgerät AQT90R0052N010 installiert. QC und Testmessungen durchgeführt. Alles o.k. | Reparatur Ersatzgerät AQT90R0052N010 installiert. QC und Testmessungen durchgeführt. Alles o.k. | SWO-147284 | No | Repair replacement unit AQT90R0052N010 installed. QC and test measurements. Everything OK. | Repair replacement unit AQT90R0052N010 installed. QC and test measurements. Everything OK. | Revisit required | Analyzer replaced | Analyzer replaced | Analyzer replaced
AQT90 FLEX | Main Board ern., Cups entfernt, Software aufgespielt, srec Datei eingespielt, Gerät neu gestartet | div. Sensoren gereinigt, Lichtschranke der Optischen Einheit ern., Gerät neu gestartet, QC´s durchgef. | SWO-146755 | No | Main board ern., cups removed, software played up, srec file recorded, device restarted | div. Sensors cleaned, light barrier of the optical unit ern., device restarted, QC's through. | Defect | HW | Analyzer Control Module | PCB, Main board
AQT90 FLEX | Service Dump gezogen und Flows überprüft. Die Flows sind erhöht und zwei Werte über dem Grenzwert von 1600. Wartungsmassnahme fällig, durchgeführt, siehe extra Service Report. Tausch der Pumpenköpfe P2 und P4. Pumpenachsen stark verschmutzt mit hellen Kristallen, gereinigt. Tausch von mehreren Schläuchen, Nummer 4 und 6 und Schlauch zur Cup Wash Unit. Einlassnadel ausgebaut und mit Lupe überprüft. War verschmutzt, gründlich gereinigt. Einlassnadel neu auf den Washblock kalibriert. Incubator wheel unten verchmutzt durch getrocknete Spüllösung, gereinigt. Blankmessungen durchgeführt, Werte überprüft. Die meisten liegen nun unter 1400 Werten und damit o. K. Qualitätskontrolle in Ordnung. | Service Dump gezogen und Flows überprüft. Die Flows sind erhöht und zwei Werte über dem Grenzwert von 1600. Wartungsmassnahme fällig, durchgeführt, siehe extra Service Report. Tausch der Pumpenköpfe P2 und P4. Pumpenachsen stark verschmutzt mit hellen Kristallen, gereinigt. Tausch von mehreren Schläuchen, Nummer 4 und 6 und Schlauch zur Cup Wash Unit. Einlassnadel ausgebaut und mit Lupe überprüft. War verschmutzt, gründlich gereinigt. Einlassnadel neu auf den Washblock kalibriert. Incubator wheel unten verchmutzt durch getrocknete Spüllösung, gereinigt. Blankmessungen durchgeführt, Werte überprüft. Die meisten liegen nun unter 1400 Werten und damit o. K. Qualitätskontrolle in Ordnung. | SWO-146851 | No | Service dump dragged and flows checked. The flows are increased and two values above the limit of 1600. Maintenance measure due, see extra Service Report. Replacement of pump heads P2 and P4. Pump axes heavily polluted with bright crystals, cleaned. Exchange of several hoses, number 4 and 6 and hose to the Cup Wash Unit. Inlet needle removed and checked with magnifying glass. Was dirty, thoroughly cleaned. recalibrated to the washblock. Incubator wheel below suspected by dried rinsing solution, cleaned. Blank measurements performed, values checked. Most of them are now below 1400 values and thus o.K. quality control is fine. | Service dump dragged and flows checked. The flows are increased and two values above the limit of 1600. Maintenance measure due, see extra Service Report. Replacement of pump heads P2 and P4. Pump axes heavily polluted with bright crystals, cleaned. Exchange of several hoses, number 4 and 6 and hose to the Cup Wash Unit. Inlet needle removed and checked with magnifying glass. Was dirty, thoroughly cleaned. recalibrated to the washblock. Incubator wheel below suspected by dried rinsing solution, cleaned. Blank measurements performed, values checked. Most of them are now below 1400 values and thus o.K. quality control is fine. | Cleaned | HW | Aspiration Module | Pump - P4
AQT90 FLEX | AQT geöffnet und Staub entfernt, insbesondere am Scanner des Storagewheels, Probelauf | AQT geöffnet und Staub entfernt, insbesondere am Scanner des Storagewheels, Probelauf | SWO-146801 | No | AQT opened and dust removed, especially on the scanner of the storage wheel, test run | AQT opened and dust removed, especially on the scanner of the storage wheel, test run | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX BASIC |  | B | SWO-146953 | No
ABL90 FLEX | Készülék az előző napi szerzív után ugyanzat a hibajelenséget produkálta, ezért a helyszinen nem javítható. A helyszinen cserekészülék üzemel a javítás ideéig. Fogyóanyagok cseréje, Kalibrálás, QC mérés | B | SWO-147071 | No | After the previous day's contract, the device produced the same problem, so it cannot be repaired on the site. There is a replacement for the site. Replacement of consumables, calibration, QC measurement |  | Unknown | Analyzer replaced | Analyzer replaced | Analyzer replaced
ABL90 FLEX PLUS | TouchScreen erneuert Kalibration und QC durchgeführt | TouchScreen erneuert Kalibration und QC durchgeführt | SWO-146979 | No | TouchScreen renews calibration and QC performed | TouchScreen renews calibration and QC performed | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX | 更换试剂包。 | 更换试剂包。 | SWO-147226 | No | Replace the reagent kit. | Replace the reagent kit. | Assisted Replacement | SP | SP | FT | 1245
AQT90 FLEX | 检查进样针的位置，发现传感器的状态值不稳定。更换636-549 修复 by jesson | 检查进样针的位置，发现传感器的状态值不稳定。更换636-549 修复 by jesson | SWO-146923 | No | Check the position of the sample pin and find that the sensor's status value is not stable. Replace 636-549 Fix by jesson | Check the position of the sample pin and find that the sensor's status value is not stable. Replace 636-549 Fix by jesson | Defect | HW | Aspiration Module | Needle Sensor | 4794
ABL820 FLEX | checked analyzer screen for function tried to replicate the problem | ran a screen calibration | D1A78C31805A55ED | No | checked analyzer screen for function tried to replicate the problem | ran a screen calibration | Adjustment | HW | DMS_PC | Touch screen
ABL827 FLEX | Found that the patient profile was not setup correctly. | Removed the ranges and linked the type to calibration verification under patient profile and setup. | D1A7E2558024B7A6 | No | Found that the patient profile was not setup correctly. | Removed the ranges and linked the type to calibration verification under patient profile and setup. | Setup | SW | Patient Layout | Patient Layout
ABL825 FLEX | THB Kal durchgeführt, Messkammer auf Undichtigkeiten überprüft und OK. OXI Ventil gereinigt, OXI Modul erneuert. Gerät kalibriert. | THB Kal durchgeführt, Messkammer auf Undichtigkeiten überprüft und OK. OXI Ventil gereinigt, OXI Modul erneuert. Gerät kalibriert. | SWO-147213 | No | THB Kal performed, measuring chamber checked for leaks and OK. OXI valve cleaned, OXI module renewed. device calibrated. | THB Kal performed, measuring chamber checked for leaks and OK. OXI valve cleaned, OXI module renewed. device calibrated.
ABL80 FLEX BASIC |  | B | SWO-147315 | No
ABL80 FLEX CO-OX | Sensor connector board stark verschmutzt mit Flüssigkeit ( rot ) ABL zur Rep mitgenommen | ABL nach Rep. Wieder aufgestellt --> Kal. Und QC i.O | SWO-147440 | No | Sensor connector board heavily polluted with liquid ( red ) ABL taken to Rep | ABL to Rep. Repositioned --> Kal. And QC i.O
ABL80 FLEX CO-OX | ABL nach Rep. Wieder aufgestellt --> Kal. Und QC i.O | ABL nach Rep. Wieder aufgestellt --> Kal. Und QC i.O | SWO-147439 | No | ABL to Rep. Repositioned --> Kal. And QC i.O | ABL to Rep. Repositioned --> Kal. And QC i.O
ABL90 FLEX | Allgemeine Überprüfung - Analysator gespült - Undichtigkeit an Schlauch für Quetschventil - Erneuerung Schlauch - Funktionsüberprüfung etc. | Allgemeine Überprüfung - Analysator gespült - Undichtigkeit an Schlauch für Quetschventil - Erneuerung Schlauch - Funktionsüberprüfung etc. | SWO-147470 | No | General inspection - Analyzer rinsed - Leakage to hose for crushing valve - Renewal hose - Function check etc. | General inspection - Analyzer rinsed - Leakage to hose for crushing valve - Renewal hose - Function check etc. | Defect | HW | Oxi module | Pinch valve tube
ABL90 | Confirmed problem. | Replaced the solution pack and the inlet gasket. | D1A97F41F63E2B7D | No | Confirmed problem. | Replaced the solution pack and the inlet gasket. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Showed the Sr. how to activate the Solution pack correctly. | Showed the Sr. how to activate the Solution pack correctly. | SWO-150077 | No | Showed the Sr. how to activate the Solution pack correctly. | Showed the Sr. how to activate the Solution pack correctly.
AQT90 FLEX |  | B | SWO-147233 | No |  |  | Defect | HW | Aspiration Module | Cup wash unit
ABL800 BASIC | Restarted the instrument and calibrated the touch screen found working properly instrument under observation. | Restarted the instrument and calibrated the touch screen found working properly instrument under observation. | SWO-148009 | No | Restarted the instrument and calibrated the touch screen found working properly instrument under observation. | Restarted the instrument and calibrated the touch screen found working properly instrument under observation. | Defect | HW | DMS_PC | Touch screen | Revisit required??
ABL825 FLEX | Examined data logs and Oxi module. | Performed complete oxi test. All oxi components failed. Replaced oxi PCB and lamp assembly. Performed service tHB calibration. Verified operation of ABL825. | D1A90CADF658EF47 | No | Examined data logs and Oxi module. | Performed complete oxi test. All oxi components failed. Replaced oxi PCB and lamp assembly. Performed service tHB calibration. Verified operation of ABL825. | Defect | HW | Oxi module | Oxi module
AQT90 FLEX | Realigned the bottom cartridge barcode scanner. Scanned in 4 D-dimmer test cartridges. New TnT LOT cartridge ordered for site. | Realigned the bottom cartridge barcode scanner. Scanned in 4 D-dimmer test cartridges. New TnT LOT cartridge ordered for site. | SWO-147505 | No | Realigned the bottom cartridge barcode scanner. Scanned in 4 D-dimmer test cartridges. New TnT LOT cartridge ordered for site. | Realigned the bottom cartridge barcode scanner. Scanned in 4 D-dimmer test cartridges. New TnT LOT cartridge ordered for site. | Adjustment | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX Basic version | Substituição da placa de sensor. O problema mantem-se. Substituição da cassete de electrodos. Inicialização. Verificação do correcto funcionamento. | B | D1AA19F96C23BDC4 | No | Replacing the sensor plate. The problem remains. Replacing the electro-cassette. Startup. Verification of proper operation.
ABL80 CO-OX | Internal and external QC problems with QC falling out of insert ranges resulting in frequent supply changes. | Checked zero settings on both analyzers. OK. | D1AA12B56C124D5E | No | Internal and external QC problems with QC falling out of insert ranges resulting in frequent supply changes. | Checked zero settings on both analyzers. OK.
ABL80 FLEX CO-OX | Das beschriebene Problem trat am 7. 8.11.2016 gehäuft auf (von 21 Messungen wurden 9 angezweifelt), am 17.11.2016 von 10 Messungen eine, PO2 Kanal überprüft, Nullpunktjustierung, Funktionskontrolle | B | SWO-147610 | No | The described problem occurred on 7.11.2016 (out of 21 measurements 9 were questioned), on 17.11.2016 one of 10 measurements, PO2 Channel checked, zero adjustment, function control
ABL90 | Analyzer is locked. Instrument calibration and internal QC's are questioned. | Ran an instrument calibration, followed by the necessary internal QC's. This unlocked the analyzer and turned all parameters green. Upon running external QC, inhomogenous sample errors and calibration errors were present. Replaced the SC and re-ran all necessary QC's and calibrations successfully. | D1AA0F0E6C19270A | No | Analyzer is locked. Instrument calibration and internal QC's are questioned. | Ran an instrument calibration, followed by the necessary internal QC's. This unlocked the analyzer and turned all parameters green. Upon running external QC, inhomogenous sample errors and calibration errors were present. Replaced the SC and re-ran all necessary QC's and calibrations successfully. | Assisted Replacement | SC | SC | Lac
ABL90 | Analyzer is locked. Instrument calibration and internal QC's are questioned. Multiple calibration failures are present | Replaced the SC and re-ran all necessary QC's and calibrations successfully. Also ran whole blood through the analyzer to assist with the conditioning. | D1AA2DBF6C03270D | No | Analyzer is locked. Instrument calibration and internal QC's are questioned. Multiple calibration failures are present | Replaced the SC and re-ran all necessary QC's and calibrations successfully. Also ran whole blood through the analyzer to assist with the conditioning. | Assisted Replacement | SC | SC | Ca++
ABL90 FLEX | Fehlersuche, das ursprüngliche Problem konnte durch Erneuerung des LP nicht behoben werden, da dieser defekt ist (wird vom Kunden reklamiert), LP erneuert, auf Serviceprogrammebene wird nun der Pressure Test, Solution Pack Check etc. bestanden, Startphase ok System-und Funktionskontrolle i.O. | Fehlersuche, das ursprüngliche Problem konnte durch Erneuerung des LP nicht behoben werden, da dieser defekt ist (wird vom Kunden reklamiert), LP erneuert, auf Serviceprogrammebene wird nun der Pressure Test, Solution Pack Check etc. bestanden, Startphase ok System-und Funktionskontrolle i.O. | SWO-147783 | No | Troubleshooting, the original problem could not be solved by renewing the LP, because it is defective (complained by the customer), LP is renewed, at the service program level the Pressure Test, Solution Pack Check etc. is passed, start phase ok System and function control in the | Troubleshooting, the original problem could not be solved by renewing the LP, because it is defective (complained by the customer), LP is renewed, at the service program level the Pressure Test, Solution Pack Check etc. is passed, start phase ok System and function control in the | Assisted Replacement | SP | SP | FT
ABL90 | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again | replaced power supply | D1AA56FB6C01905A | No | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again | replaced power supply | Defect | HW | Power Supply | Power Supply
ABL90 | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | D1AA58626C1A9060 | No | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | Defect | HW | Power Supply | Power Supply
ABL90 | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | D1AA590C6C4A9063 | No | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | Defect | HW | Power Supply | Power Supply
ABL90 | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | D1AA59E96C4C9068 | No | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | Defect | HW | Power Supply | Power Supply
ABL90 | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | D1AA5A9E6C3D906C | No | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | Defect | HW | Power Supply | Power Supply
ABL90 | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | D1AA5CA06C4E9073 | No | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | Defect | HW | Power Supply | Power Supply
ABL90 | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | D1AA5D656C009077 | No | disconnected the mains power from an analyzer with a battery installed, the analyzer would still run on battery when mains power is connected again. | replaced power supply | Defect | HW | Power Supply | Power Supply
AQT90 FLEX |  | B | SWO-147562 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer
ABL825 FLEX |  | B | D1AA3A996C2121F8 | No |  |  | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL825 FLEX |  | B | D1AE01B6CE17EA42 | No |  |  | Assisted Replacement | Electrode | Electrode | Lac | Sensitivity
ABL800 BASIC | replaced analyzer control PCB but its not working its showing DMS not connected to analyzer we need to replace again new analyzer control PCB for this case. | B | SWO-149398 | No | replaced analyzer control PCB but its not working its showing DMS not connected to analyzer we need to replace again new analyzer control PCB for this case. |  | Revisit required | HW | Basic Unit | Analyzer control PCB
ABL825 FLEX | Schlauch BG Modul ist am Dual Valve Housing mehrfach abgesprungen ohne erkennbaren Grund - Dual Valve Housing ersetzt, Ventil Fehlschaltungen Air Ventil BG Modul im Speicher - Ventil ersetzt, Kal und QC durchgeführt | Schlauch BG Modul ist am Dual Valve Housing mehrfach abgesprungen ohne erkennbaren Grund - Dual Valve Housing ersetzt, Ventil Fehlschaltungen Air Ventil BG Modul im Speicher - Ventil ersetzt, Kal und QC durchgeführt | SWO-147837 | No | Tube BG module has jumped off the Dual Valve Housing several times for no apparent reason - Dual Valve Housing replaced, valve faulty circuits Air valve BG module replaced in memory valve, Kal and QC performed | Tube BG module has jumped off the Dual Valve Housing several times for no apparent reason - Dual Valve Housing replaced, valve faulty circuits Air valve BG module replaced in memory valve, Kal and QC performed | Defect | HW | pH/BG Module | Dual valve housing
AQT90 FLEX |  | B | D1AA3F7E6C26EFA0 | No |  |  | Defect | Consumables | SP | SP
AQT90 FLEX |  | B | D1AA3BF06C0BB9B0 | Yes |  |  | Re classified RMED | PM
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-149069 | No | Replacement test card repair | Replacement test card repair
ABL815 FLEX | Bemeneti kapilláris csö visszahelyezése, Elszívó csőszakaszok tisztítása, Bemeneti egység átmosása, Tű egyesítése és hypós tisztítása, Levegőző szelep tisztítás, Levegőszűrő tiszt. Minta tesztelés (8x), Kalibrációk (Inlet, LS, Pumpák, Elektródák) QC Mérés (Level1 & Level2) | B | SWO-147904 | No | Replacing the inlet capillary tube, cleaning suction pipe sections, washing input unit, uniting the needle and cleaning the hypto, ventilation valve cleaning, air filter officer. Sample testing (8x), calibrates (Inlet, LS, pumps, electrodes) QC measurement (Level1 & Level2) |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL80 FLEX BASIC |  | B | SWO-147980 | No
ABL80 FLEX BASIC |  | B | SWO-147981 | No
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-148127 | No | Replace the test card | Replace the test card
ABL90 FLEX | Einlassdichtung erneuert. Gerät im Servicemenu überprüft. Startphase danach 2Std. | Einlassdichtung erneuert. Gerät im Servicemenu überprüft. Startphase danach 2Std. | SWO-148368 | No | Inlet seal renewed. device in the service menu. Start phase after 2hrs | Inlet seal renewed. device in the service menu. Start phase after 2hrs | Defect | HW | Inlet module | Inlet Gasket
AQT90 FLEX | 更换910-343一个正常。 | B | SWO-147974 | No | Replace 910-343 one normal. |  | Defect | HW | Analyzer Control Module | CPU
ABL800 FLEX | Replaced membrane and gases still showing the same issue. Checked the pressure transducer and solenoid valve according to the service program. Its found ok. Has to change the PCO2 electrode. Part Number: 945-612-E788 pCO2 electrode | B | SWO-149547 | No | Replaced membrane and gases still showing the same issue. Checked the pressure transducer and solenoid valve according to the service program. Its found ok. Has to change the PCO2 electrode. Part Number: 945-612-E788 pCO2 electrode |  | Revisit required | Electrode | Electrode | pCO2
ABL837 FLEX | No send button is present to allow for patient data transfer. The activity log also indicates that the LIS HIS communication key is present. | Reinstalled software keys and verified the presence of the HIS LIS communication key, activated it and the send button was once again present. Patient samples were run and data successfully transferred. Analyzer is within acceptable operating specifications. | D1AE23B6CE11ED3F | No | No send button is present to allow for patient data transfer. The activity log also indicates that the LIS HIS communication key is present. | Reinstalled software keys and verified the presence of the HIS LIS communication key, activated it and the send button was once again present. Patient samples were run and data successfully transferred. Analyzer is within acceptable operating specifications. | Setup | SW | His/Lis | His/Lis
ABL827 FLEX |  | B | D1AF542A4447B980 | No |  |  | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | REVEIWED DATA LOG AND NO OUTSTANDING ISSUES FOUND. LED IN BOTH MEASUREMENT CHAMBERS ARE CLEARLY VISIBLE SO NO NEED TO BE CHANGED. | REPLACED ALL PARTS IN PM KIT PER CURRENT PROCEDURE. CLEANED AND REMOVED DUST ACCUMULATED INSIDE SYSTEM. | D1ADF24BCE6057A6 | Yes | REVEIWED DATA LOG AND NO OUTSTANDING ISSUES FOUND. LED IN BOTH MEASUREMENT CHAMBERS ARE CLEARLY VISIBLE SO NO NEED TO BE CHANGED. | REPLACED ALL PARTS IN PM KIT PER CURRENT PROCEDURE. CLEANED AND REMOVED DUST ACCUMULATED INSIDE SYSTEM. | Re classified RMED | PM
AQT90 FLEX | Kundengerät nach GS17 Reparatur zurück zum Kunden. Systemüberprüfung und QC o.k. | Kundengerät nach GS17 Reparatur zurück zum Kunden. Systemüberprüfung und QC o.k. | SWO-148366 | No | Customer device after GS17 repair back to the customer. System check and QC ok | Customer device after GS17 repair back to the customer. System check and QC ok | LEH Mail |  |  | Not Cat. | LEH Mail
ABL80 FLEX BASIC | 更换试剂包后修复 | 更换试剂包后修复 | SWO-149031 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit
ABL80 FLEX BASIC |  | B | SWO-148384 | No
ABL80 FLEX BASIC |  | B | SWO-148372 | No
ABL90 FLEX | LP erneuert | LP erneuert | SWO-148477 | No | LP renewed | LP renewed | User error | SP | SP | SP empty
ABL90 | Confirmed problem. | Replaced all of the tubing, inlet connector gasket and the inlet gasket with holder using a maint kit. Flushed the ensor cassette with warm water three times. | D1AFAC024433670E | No | Confirmed problem. | Replaced all of the tubing, inlet connector gasket and the inlet gasket with holder using a maint kit. Flushed the ensor cassette with warm water three times. | Defect | HW | PM Kit | PM Kit
ABL90 FLEX PLUS | Found inlet slipping on coupling, retightened, now ok. | Found inlet slipping on coupling, retightened, now ok. | SWO-148620 | No | Found inlet slipping on coupling, retightened, now ok. | Found inlet slipping on coupling, retightened, now ok. | Adjustment | HW | Auto Inlet module | Motor assembly | TU issued
ABL90 FLEX PLUS | Replaced inlet gasket, now ok. | Replaced inlet gasket, now ok. | SWO-148621 | No | Replaced inlet gasket, now ok. | Replaced inlet gasket, now ok. | Defect | HW | Auto Inlet module | Inlet Gasket
ABL825 FLEX | Replaced all original parts. New spectrometer unit needed | Fitted new spectrometer unit. ran complete oxi test. All now passsing Tested and checked. All working ok | SWO-149084 | No | Replaced all original parts. New spectrometer unit needed | Fitted new spectrometer unit. ran complete oxi test. All now passsing Tested and checked. All working ok | Revisit required | HW | Oxi module | Spectrometer | Good ex.
ABL800 BASIC | check the fluidic module. check the valve. replace the valve. Need top check the fluidic module PCB. | Replace the inlet liquid sensor.Replace the PH BG module air & rinse valve. check the fluidic module. valve. rinse valve replace. the valve movement error replace the fluidic PCB. the Fluidic PCB faulty causing issue in sample aspiration. Need to replace the fludic PCB. | SWO-149527 | No | check the fluidic module. check the valve. replace the valve. Need top check the fluidic module PCB. | Replace the inlet liquid sensor. Replace the PH BG module air & rinse valve. check the fluidic module. valve. rinse valve replace. the valve movement error replace the fluidic PCB. the Fluidic PCB faulty causing issue in sample aspiration. Need to replace the fludic PCB. | Revisit required | HW | Fluidic manifold | PCB, Fludig PCB
ABL800 BASIC | Replaced reference electrode. | Replaced reference electrode. | SWO-148687 | No | Replaced reference electrode. | Replaced reference electrode. | Assisted Replacement | Electrode | Electrode | Reference
ABL800 BASIC | replaced new oxi module now its its working fine | B | SWO-149402 | No | Replaced new OXI module now its its working fine |  | Defect | HW | Oxi module | Oxi module
ABL835 FLEX |  | B | D1AF50624456B0E9 | No |  |  | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | STK durchgeführt. Zählerstand: 15777 | STK durchgeführt. Zählerstand: 15777 | SWO-148586 | Yes | STK. Counter reading: 15777 | STK. Counter reading: 15777 | Re classified RMED | Not upSWO
ABL825 FLEX | Ersatzteil bestellt. Reparatur voraussichtlich in KW 48 nach Lieferung des Ersatzteiles. | Ersatzteil bestellt. Reparatur voraussichtlich in KW 48 nach Lieferung des Ersatzteiles. | SWO-148453 | No | spare part ordered. Repair expected in KW 48 after delivery of the spare part. | spare part ordered. Repair expected in KW 48 after delivery of the spare part. | Defect | HW | FlexQ Module | LED
ABL825 FLEX | Nadel ist ganz gleichmäßig nach oben gebogen, Anwenderfehler sind nach diversen Schulungen und aufgrund der gleichmäßigen Biegung der Nadel nahezu auszuschließen! Nadel ersetzt, weiteres Verfahren im Rahmen der Halbjahreswartung. | Nadel ist ganz gleichmäßig nach oben gebogen; Anwenderfehler sind nach diversen Schulungen und aufgrund der gleichmäßigen Biegung der Nadel nahezu auszuschließen! Nadel ersetzt, weiteres Verfahren im Rahmen der Halbjahreswartung. | SWO-148454 | No | Needle is evenly bent upwards; User errors can be almost ruled out after various training sessions and due to the uniform bending of the needle! needle replaced, further procedures in the context of half-year maintenance. | Needle is evenly bent upwards; User errors can be almost ruled out after various training sessions and due to the uniform bending of the needle! needle replaced, further procedures in the context of half-year maintenance. | Defect | HW | Inlet module | Probe bend
AQT90 FLEX |  | B | D1AF121F442AE965 | No |  |  | Defect | HW | Process Module | Optical Unit
AQT90 FLEX | Nadel ausgebaut und gereinigt, Fließwege ausgewertet Flow < 1300. Funktion zur Zeit OK. | Nadel ausgebaut und gereinigt, Fließwege ausgewertet Flow < 1300. Funktion zur Zeit OK. | SWO-148589 | No | Needle removed and cleaned, flow paths evaluated flow 1300. Function currently OK. | Needle removed and cleaned, flow paths evaluated flow 1300. Function currently OK. | Defect | SW | Analyzer Control Module | XML format | 35
AQT90 FLEX |  | B | D1AF0D9B442628F9 | Yes |  |  | Re classified RMED | Techline doc
AQT90 FLEX |  | B | D1AF10A6444628FB | Yes |  |  | Re classified RMED | Techline doc
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-148639 | No | Replacement test card repair | Replacement test card repair
ABL90 FLEX | Replace bent probe and damaged gasket with new Provided Jennifer Brown with spare gaskets and probe | Replace bent probe and damaged gasket with new Provided Jennifer Brown with spare gaskets and probe | SWO-148873 | No | Replace bent probe and damaged gasket with new Provided Jennifer Brown with spare gaskets and probe | Replace bent probe and damaged gasket with new Provided Jennifer Brown with spare gaskets and probe | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Fitted new sensor cassette, still failing Na and Ca. Checked pins, found a number of them bent. On closer inspection found that the sensor cassette is misaligned with the pins. Went into service mode and found the Ca measured impedance at -0.86. Tried to reposition bay and straighten pins. Will need setting correctly with the correct tool and will also need bent pins replacing. | Fitted new sensor cassette, still failing Na and Ca. Checked pins, found a number of them bent. On closer inspection found that the sensor cassette is misaligned with the pins. Went into service mode and found the Ca measured impedance at -0.86. Tried to reposition bay and straighten pins. Will need setting correctly with the correct tool and will also need bent pins replacing. | SWO-148923 | No | Fitted new sensor cassette, still failing Na and Ca. Checked pins, found a number of them bent. On closer inspection found that the sensor cassette is misaligned with the pins. Went into service mode and found the Ca measured impedance at -0.86. Tried to reposition bay and straighten pins. Will need setting correctly with the correct tool and will also need bent pins replacing. | Fitted new sensor cassette, still failing Na and Ca. Checked pins, found a number of them bent. On closer inspection found that the sensor cassette is misaligned with the pins. Went into service mode and found the Ca measured impedance at -0.86. Tried to reposition bay and straighten pins. Will need setting correctly with the correct tool and will also need bent pins replacing. | Revisit required | HW | Basic Unit | Sensor bay
ABL90 FLEX | Gerät gereigt, verstopfung lokalisiert und ansaugenadel ersetzt. Gerät neu kalibriert I.O. | Gerät gereigt, verstopfung lokalisiert und ansaugenadel ersetzt. Gerät neu kalibriert I.O. | SWO-148794 | No | Device relocated, clogging localized and suction superseded. Device recalibrated I.O. | Device relocated, clogging localized and suction superseded. Device recalibrated I.O. | Cleaned | Fluidic pathway | Fluidic pathway | Probe
ABL90 FLEX | Zeiteinstellungen überprüft auf der Windows und auf der Analysatorebene, Commit-2-C-Drive Befehl ausgeführt mit Neustart, Zeiteinstellung nun ok. | Zeiteinstellungen überprüft auf der Windows und auf der Analysatorebene, Commit-2-C-Drive Befehl ausgeführt mit Neustart, Zeiteinstellung nun ok. | SWO-148775 | No | Time settings checked on Windows and at the analyzer level, commit-2-C-Drive command executed with restart, time setting now ok. | Time settings checked on Windows and at the analyzer level, commit-2-C-Drive command executed with restart, time setting now ok. | Setup | SW | Time setting | Time setting
AQT90 FLEX |  | B | SWO-148712 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL800 BASIC | Replaced inlet gasket. Cleaned printer head. | Replaced inlet gasket. Cleaned printer head. | SWO-149007 | No | Replaced inlet gasket. Cleaned printer head. | Replaced inlet gasket. Cleaned printer head. | Defect | HW | Inlet module | Inlet Gasket
ABL835 FLEX | 18-11-2016: CPU hjemtaget til rep. 21-11-2016: CPU skiftet pga. 3x gentart i weekenden 22-11-2016: CPU board skiftet 23-11-2016: CPU monteret i kundens apparat | 18-11-2016: CPU hjemtaget til rep. 21-11-2016: CPU skiftet pga. 3x gentart i weekenden 22-11-2016: CPU board skiftet 23-11-2016: CPU monteret i kundens apparat | SWO-148556 | No | 18-11-2016: CPU Home taken to Rep. 21-11-2016: CPU Change due to 3x Gentart on weekends 22-11-2016: CPU Board switched 23-11-2016: CPU mounted in the customer's device | 18-11-2016: CPU Home taken to Rep. 21-11-2016: CPU Change due to 3x Gentart on weekends 22-11-2016: CPU Board switched 23-11-2016: CPU mounted in the customer's device | Defect | HW | DMS_PC | CPU
ABL805 FLEX | Datei Export zur Weitersendung an RDEU Service zur Untersuchung des Problems. | Datei Export zur Weitersendung an RDEU Service zur Untersuchung des Problems. | SWO-148846 | Yes | File export to RDEU Service to investigate the problem. | File export to RDEU Service to investigate the problem. | Re classified RMED | Collecting data logs
ABL815 FLEX | Gerät überprüft, zurzeit grün OK. Datenexport zur Weiterleitung an RDEU Service gezogen und aufbereitet. | Gerät überprüft, zurzeit grün OK. Datenexport zur Weiterleitung an RDEU Service gezogen und aufbereitet. | SWO-148830 | No | Device checked, currently green OK. Data export for forwarding to RDEU Service. | Device checked, currently green OK. Data export for forwarding to RDEU Service. | Unknown | Electrode | Electrode | pCO2
ABL827 FLEX |  | B | D1B0BDEDBA206715 | Yes |  |  | Re classified RMED | Not upSWO
ABL827 FLEX | Proteinentfernung durchgeführt, Service Dump und DataLogs gezogen, AU Probleme ausgelesen und gespeichert, mehrere Parameter membraniert - Fälligkeit war erreicht, Messkammern gereinigt, Lufteinlass gespült, System- und Funktionskontrolle in Ordnung. | Proteinentfernung durchgeführt, Service Dump und DataLogs gezogen, AU Probleme ausgelesen und gespeichert, mehrere Parameter membraniert - Fälligkeit war erreicht, Messkammern gereinigt, Lufteinlass gespült, System- und Funktionskontrolle in Ordnung. | SWO-148816 | No | Protein removal performed, service dump and DataLogs pulled, AU problems read out and stored, several parameters membraned - maturity was reached, measuring chambers cleaned, air inlet flushed, system and function control in order. | Protein removal performed, service dump and DataLogs pulled, AU problems read out and stored, several parameters membraned - maturity was reached, measuring chambers cleaned, air inlet flushed, system and function control in order. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX |  | B | D1B0A52FBA619D8F | No |  |  | Cleaned | HW | Incubator module | PCB, Incubator
ABL80 FLEX BASIC | 更换件933-306，905-914和903-381后修复 | 更换件933-306，905-914和903-381后修复 | SWO-148682 | No | Replacement parts 933-306, 905-914 and 903-381 after repair | Replacement parts 933-306, 905-914 and 903-381 after repair
ABL90 FLEX | Fehlersuche Geräteneustart LP, SC und Einlaßdichtung wurden durch den Kunden bereits erneuert , Blutgerinnsel aus Einlaßnadel entfernt Überprüfung im Service Programm SC erneuert, da Lac und Glu Fehler bei Kal. Einlaufkalibrierung | Fehlersuche Geräteneustart LP, SC und Einlaßdichtung wurden durch den Kunden bereits erneuert , Blutgerinnsel aus Einlaßnadel entfernt Überprüfung im Service Programm SC erneuert, da Lac und Glu Fehler bei Kal. Einlaufkalibrierung | SWO-149062 | No | Troubleshooting Device restart LP, SC and inlet seal have already been renewed by the customer, Blood clot removed from inlet needle Review in the service program SC renewed as Lac and Glu errors at Kal. Inlet calibration | Troubleshooting Device restart LP, SC and inlet seal have already been renewed by the customer, Blood clot removed from inlet needle Review in the service program SC renewed as Lac and Glu errors at Kal. Inlet calibration | Cleaned | Fluidic pathway | Fluidic pathway | Probe | Clot
ABL90 FLEX PLUS | Nieuwe verwarmingsweerstanden en regelweerstand gemonteerd. Nu weer een correcte regeling. Instrument opgestart. | Opgestart, radiance IP-adres ingegeven. THb calibratie uitgevoerd en hierna een bloedmonster gemeten. | SWO-149093 | No | New heating resistors and control resistor mounted. Now again a correct arrangement. Instrument is started. | Started, Radiance IP address entered. THb calibration performed and then a blood sample is measured. | Defect | HW | Basic Unit | Thermistor w cable
ABL90 FLEX PLUS | Opgestart, radiance IP-adres ingegeven. THb calibratie uitgevoerd en hierna een bloedmonster gemeten. | Opgestart, radiance IP-adres ingegeven. THb calibratie uitgevoerd en hierna een bloedmonster gemeten. | SWO-149127 | Yes | Started, Radiance IP address entered. THb calibration performed and then a blood sample is measured. | Started, Radiance IP address entered. THb calibration performed and then a blood sample is measured. | Re classified RMED
AQT90 FLEX |  | B | SWO-149000 | No |  |  | Defect | HW | Storage Module | Stepper Motor, punch out
AQT90 FLEX |  | B | SWO-148944 | No |  |  | Defect | HW | Aspiration Module | Barcode reader, Inlet
AQT90 FLEX |  | B | SWO-148992 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit
AQT90 FLEX | 更换903-186后修复(by andrew) | 更换903-186后修复(by andrew) | SWO-148630 | No | Repair after replacement 903-186 (by andrew) | Repair after replacement 903-186 (by andrew) | Defect | HW | Aspiration Module | Barcode reader, Inlet
AQT90 FLEX | 更换886-625一个正常。（by andrew） | 更换886-625一个正常。（by andrew） | SWO-148707 | No | Replace 886-625 one normal. (by andrew) | Replace 886-625 one normal. (by andrew) | Defect | HW | Storage Module | Cartridge guide
ABL800 FLEX | Replace the Crea consumable. Calibrate system & calibration pass. | Replace the Crea consumable. Calibrate system & calibration pass. | SWO-149520 | No | Replace the Crea consumable. Calibrate system & calibration pass. | Replace the Crea consumable. Calibrate system & calibration pass. | Assisted Replacement | Electrode | Solution bottle | Crea
ABL825 FLEX | Fitted new spectrometer unit. ran complete oxi test. All now passsing Tested and checked. All working ok | Fitted new spectrometer unit. ran complete oxi test. All now passsing Tested and checked. All working ok | SWO-149085 | No | Fitted new spectrometer unit. ran complete oxi test. All now passsing Tested and checked. All working ok | Fitted new spectrometer unit. ran complete oxi test. All now passsing Tested and checked. All working ok | Defect | HW | Oxi module | Spectrometer | Good ex.
ABL800 BASIC | Printer head required to rectify the error | Printer head required to rectify the error | SWO-149054 | No | Printer head required to rectify the error | Printer head required to rectify the error | Defect | HW | Printer module | Printer head
ABL825 FLEX | Remplacé la pompe péristaltique e.o Test étanchéité de la section fluidique e.o Calibration e.o contrôle de qualité e.o | Remplacé la pompe péristaltique e.o Test étanchéité de la section fluidique e.o Calibration e.o contrôle de qualité e.o | SWO-149038 | No | Replaced the peristaltic pump e.o Waterproofing test of the fluidic section e.o Calibration e.o quality control e.o | Replaced the peristaltic pump e.o Waterproofing test of the fluidic section e.o Calibration e.o quality control e.o | Defect | HW | pH/BG Module | Pump rotor
ABL815 FLEX |  | B | D1B1DE2F303AC7CA | No |  |  | Defect | SW | Win XP | Win XP
ABL80 FLEX CO-OX | Gerät zerlegt, Flüssigkeitswege durchgespült, keine Stenose feststellbar. Nach Neustart Kal und QC in Ordnung. Gerät heruntergefahren. LP und SC auf Undichtigkeiten überprüft, kein Fehler feststellbar. Neustart in Ordnung. | Gerät zerlegt, Flüssigkeitswege durchgespült, keine Stenose feststellbar. Nach Neustart Kal und QC in Ordnung. Gerät heruntergefahren. LP und SC auf Undichtigkeiten überprüft, kein Fehler feststellbar. Neustart in Ordnung. | SWO-149542 | No | device disassembled, fluid pathways rinsed through, no stenosis detectable. After rebooting Kal and QC fine. Device shut down. LP and SC checked for leaks, no error detectable. Restart okay. | device disassembled, fluid pathways rinsed through, no stenosis detectable. After rebooting Kal and QC fine. Device shut down. LP and SC checked for leaks, no error detectable. Restart okay.
ABL80 FLEX CO-OX | Fehlersuche, Reinigung Flüssigkeitssensor, Funktionskontrolle | Fehlersuche, Reinigung Flüssigkeitssensor, Funktionskontrolle | SWO-149431 | No | Troubleshooting, cleaning liquid sensor, function control | Troubleshooting, cleaning liquid sensor, function control
ABL90 FLEX | Gerät überprüft, Lösungspack erneuert, tHb Kal durchgeführt, QC o.k. | Gerät überprüft, Lösungspack erneuert, tHb Kal durchgeführt, QC o.k. | SWO-149578 | No | Device checked, solution pack renewed, tHb Kal performed, QC ok | Device checked, solution pack renewed, tHb Kal performed, QC ok | Assisted Replacement | SP | SP | QC
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-149977 | No | Repair after replacing test card | Repair after replacing test card
ABL825 FLEX | Inletmodul demontiert und die Friction Detector Unit ausgetauscht, Funktionskontrolle ok | Inletmodul demontiert und die Friction Detector Unit ausgetauscht, Funktionskontrolle ok | SWO-149423 | No | Inlet module disassembled and the Friction Detector Unit replaced, function control ok | Inlet module disassembled and the Friction Detector Unit replaced, function control ok | Defect | HW | Inlet module | Friction detector unit
ABL825 FLEX | Einlassnadel offenbar ohne äußeren Einfluss bananenkrumm und gleichmäßig nach oben gebogen => Nadel ersetzt und zur Untersuchung an RDEU gesendet. | Einlassnadel offenbar ohne äußeren Einfluss bananenkrumm und gleichmäßig nach oben gebogen => Nadel ersetzt und zur Untersuchung an RDEU gesendet. | SWO-149397 | No | Inlet needle apparently without external influence banana crumb and evenly bent upwards => needle replaced and sent to RDEU for examination. | Inlet needle apparently without external influence banana crumb and evenly bent upwards => needle replaced and sent to RDEU for examination. | Defect | HW | Inlet module | Probe bend | Foto?
ABL825 FLEX |  | B | D1B2F33CA61BAE8D | No |  |  | Unknown | HW | Unknown | Unknown
ABL825 FLEX | Monitor (910-320r1628n1148) geprüft und ausgetauscht, Funktionskontrolle ok | Monitor (910-320r1628n1148) geprüft und ausgetauscht, Funktionskontrolle ok | SWO-149422 | No | Monitor (910-320r1628n1148) checked and replaced, function control ok | Monitor (910-320r1628n1148) checked and replaced, function control ok | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX | 更换903-364一个正常。(by andrew) | 更换903-364一个正常。(by andrew) | SWO-149057 | Yes | Replace 903-364 one normal. (by andrew) | Replace 903-364 one normal. (by andrew) | Re classified RMED | HW | Process Module | Optical Unit | 131
ABL800 BASIC | Replace the inlet liquid sensor.Replace the PH BG module air & rinse valve. check the fluidic module. valve. rinse valve replace. the valve movement error replace the fluidic PCB. the Fluidic PCB faulty causing issue in sample aspiration. Need to replace the fludic PCB. | Replace the inlet liquid sensor.Replace the PH BG module air & rinse valve. check the fluidic module. valve. rinse valve replace. the valve movement error replace the fluidic PCB. the Fluidic PCB faulty causing issue in sample aspiration. Need to replace the fludic PCB. | SWO-149572 | No | Replace the inlet liquid sensor. Replace the PH BG module air & rinse valve. check the fluidic module. valve. rinse valve replace. the valve movement error replace the fluidic PCB. the Fluidic PCB faulty causing issue in sample aspiration. Need to replace the fludic PCB. | Replace the inlet liquid sensor. Replace the PH BG module air & rinse valve. check the fluidic module. valve. rinse valve replace. the valve movement error replace the fluidic PCB. the Fluidic PCB faulty causing issue in sample aspiration. Need to replace the fludic PCB. | Revisit required | HW | Fluidic manifold | PCB, Fludic PCB
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-149583 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | pO2
ABL80 FLEX BASIC |  | B | SWO-149669 | No
ABL80 FLEX BASIC |  | B | SWO-149671 | No
ABL80 FLEX BASIC |  | B | SWO-149673 | No
ABL80 FLEX BASIC |  | B | SWO-149685 | No
ABL90 FLEX | Gerät überprüft, Sensorkassette erneuert, Lösungspack ausgetauscht | Gerät überprüft, Sensorkassette erneuert, Lösungspack ausgetauscht | SWO-149927 | No | Device checked, sensor cartridge renewed, solution pack replaced | Device checked, sensor cartridge renewed, solution pack replaced | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | 2mal SC getauscht | 2mal SC getauscht | SWO-150005 | No | 2 times SC exchanged | 2 times SC exchanged | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS |  | B | D1B6DAC108510757 | No |  |  | Assisted Replacement | SC | SC | Reference | 1361
ABL90 FLEX PLUS | System überprüft und durchgespült. Im Servicemenu sämtliche Tests mehrmals durchgeführt. SC neu installiert. Kal., QC und System i.o. Funktionskontrolle i.o. | System überprüft und durchgespült. Im Servicemenu sämtliche Tests mehrmals durchgeführt. SC neu installiert. Kal., QC und System i.o. Funktionskontrolle i.o. | SWO-149853 | No | system checked and flushed through. Performed all tests several times in the service menu. SC reinstalled. Kal., QC and System i.o. Functional control i.o. | system checked and flushed through. Performed all tests several times in the service menu. SC reinstalled. Kal., QC and System i.o. Functional control i.o. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | Sensor Simulator, Sensorkassette und Lösungspack aus der GS geholt. Mit dem Sensor Simulator Kanäle getestet>>I.O Sensorkassette Lot 527 eingebaut, Startphase QC und Kalibration. | Sensor Simulator, Sensorkassette und Lösungspack aus der GS geholt. Mit dem Sensor Simulator Kanäle getestet>>I.O Sensorkassette Lot 527 eingebaut, Startphase QC und Kalibration. | SWO-149808 | No | Sensor simulator, sensor cassette and solution pack from the GS. Tested with the Sensor Simulator channels>>I.O Sensor cassette Lot 527 built-in, start phase QC and calibration. | Sensor simulator, sensor cassette and solution pack from the GS. Tested with the Sensor Simulator channels>>I.O Sensor cassette Lot 527 built-in, start phase QC and calibration. | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX | Sustituida célula de hematrocrito y comprobado buen funcionamiento de la misma. | B | D1B6F1F5081E0345 | No | Replaced heametrit cell and proven good functioning of it. |  | Defect | HW | Aspiration Module | Needle Wash Unit
ABL825 FLEX |  | B | D1B828C37E30609B | Yes |  |  | Re classified RMED | Document dispatch of FSE+electrode
ABL825 FLEX | Analyzer was leaking out of front Gasket Noticed the probe was totally bent. Customer had tried to clean measuring chamber, replaced Pco2 membrane. | Replaced the Inlet probe Performed several calibration | D1B6C9580813D2A3 | No | Analyzer was leaking out of front Gasket Noticed the probe was totally bent. Customer had tried to clean measuring chamber, replaced Pco2 membrane. | Replaced the Inlet probe Performed several calibration | Defect | HW | Inlet module | Inlet Gasket
ABL825 FLEX | Messkammer und Schlauchsystem geprüft, pCo2 neu membraniert, Funktionskontrolle ok | Messkammer und Schlauchsystem geprüft, pCo2 neu membraniert, Funktionskontrolle ok | SWO-150185 | No | Measuring chamber and hose system tested, pCo2 re-membraneed, function control ok | Measuring chamber and hose system tested, pCo2 re-membraneed, function control ok | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL820 FLEX | confirmed issue with analyzer backed up all data | re installed 6.16 to correct the software issue. | D1B6B14D083D7F15 | No | confirmed issue with analyzer backed up all data | re installed 6.16 to correct the software issue. | Defect | SW | Radiometer SW | 6.16
ABL825 FLEX | Fehlersuche: ABL steht im eigenen Saft in einer großen Pfütze aus klarer Lösung, alle Messplätze untersucht, unterer Dichtring der Glu Elektrode beim Membranieren verschoben, dadurch Undichtigkeit und Probenabbrüche. Elektrode neu membraniert, alle anderen überprüft und pH gereinigt. Weiteres Verfahren im Rahmen der heutigen HJW. | Fehlersuche: ABL steht im eigenen Saft in einer großen Pfütze aus klarer Lösung; alle Messplätze untersucht, unterer Dichtring der Glu Elektrode beim Membranieren verschoben, dadurch Undichtigkeit und Probenabbrüche. Elektrode neu membraniert, alle anderen überprüft und pH gereinigt. Weiteres Verfahren im Rahmen der heutigen HJW. | SWO-149875 | No | Troubleshooting: ABL stands in its own juice in a large puddle of clear solution; examined all measuring stations, lower sealing ring of the glu electrode shifted during membraneing, thereby leakage and sample interruptions. Electrode re-membraneed, all others checked and pH cleaned. Further proceedings under today's HJW. | Troubleshooting: ABL stands in its own juice in a large puddle of clear solution; examined all measuring stations, lower sealing ring of the glu electrode shifted during membraneing, thereby leakage and sample interruptions. Electrode re-membraneed, all others checked and pH cleaned. Further proceedings under today's HJW. | Assisted Replacement | Electrode | Electrode | Glu | Sealing ring
ABL825 FLEX | Gerät neugestartet, Abgleich überprüft, Leihweise ein Ersatz Oxi Modul ins Gerät eingebaut. Kundeneigenes Oxi Modul zu Instandsetzung mitgenommen. | B | SWO-149858 | No | Device restarted, matching checked, loaning a replacement Oxi module built into the device. Customer-owned Oxi module taken for repair. |  | Revisit required | HW | Oxi module | Oxi module
ABL825 FLEX | Einlasssonde krum, erneuert, auch Einlassdichtung, da verschmutzt. Überprüfung des FlexQ Moduls im Service, alle Funktionen in Ordnung, auch keine Teile die Bewegung behindern. Mehrmals Proteinentfernung durchgeführt, kalibriert und tHb Kalibration mit eigenem Kalibrator durchgeführt. Werte wieder zielwertig. Kalibration und Qualitätskontrolle in Ordnung. | Einlasssonde krum, erneuert, auch Einlassdichtung, da verschmutzt. Überprüfung des FlexQ Moduls im Service, alle Funktionen in Ordnung, auch keine Teile die Bewegung behindern. Mehrmals Proteinentfernung durchgeführt, kalibriert und tHb Kalibration mit eigenem Kalibrator durchgeführt. Werte wieder zielwertig. Kalibration und Qualitätskontrolle in Ordnung. | SWO-149920 | No | Inlet probe crooked, renewed, also inlet seal, as dirty. Checking the FlexQ module in service, all functions in order, even no parts hinder the movement. Several times protein removal performed, calibrated and tHb calibration performed with its own calibrator. values again targeted. Calibration and quality control in order. | Inlet probe crooked, renewed, also inlet seal, as dirty. Checking the FlexQ module in service, all functions in order, even no parts hinder the movement. Several times protein removal performed, calibrated and tHb calibration performed with its own calibrator. values again targeted. Calibration and quality control in order. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Messkammer geprüft, pCO2 und pO2 Elektroden neu membraniert, Funktionskontrolle ok | Messkammer geprüft, pCO2 und pO2 Elektroden neu membraniert, Funktionskontrolle ok | SWO-150181 | No | Measuring chamber tested, pCO2 and pO2 electrodes re-membraneed, function control ok | Measuring chamber tested, pCO2 and pO2 electrodes re-membraneed, function control ok | Assisted Replacement | Electrode | Membrane | pCO2
AQT90 FLEX | After opening the machine, taken the optical unitout to access the process wheel, found that the cups were stuck in the process wheel, which is causing the process wheel obstruction. So removed those cups, and reconnected the machine. After restarting and after successsful initialization run 4 samples, results were fine. But after some time, the machine started giving the same error. So removed the cups in process wheel which were stuck, removed the solution pack after restarting the machine, found that the solution pack waste cup door was not functioning properly, so opened the solution pack top cover, fitted the waste cup door properly, after restarting the machine, done TropT new lot calibration, machine is working fine. | After opening the machine, taken the optical unitout to access the process wheel, found that the cups were stuck in the process wheel, which is causing the process wheel obstruction. So removed those cups, and reconnected the machine. After restarting and after successsful initialization run 4 samples, results were fine. But after some time, the machine started giving the same error. So removed the cups in process wheel which were stuck, removed the solution pack after restarting the machine, found that the solution pack waste cup door was not functioning properly, so opened the solution pack top cover, fitted the waste cup door properly, after restarting the machine, done TropT new lot calibration, machine is working fine. | SWO-151177 | No | After opening the machine, taken the optical unitout to access the process wheel, found that the cups were stuck in the process wheel, which is causing the process wheel obstruction. So removed those cups, and reconnected the machine. After restarting and after successsful initialization run 4 samples, results were fine. But after some time, the machine started giving the same error. So removed the cups in process wheel which were stuck, removed the solution pack after restarting the machine, found that the solution pack waste cup door was not functioning properly, so opened the solution pack top cover, fitted the waste cup door properly, after restarting the machine, done TropT new lot calibration, machine is working fine. | After opening the machine, taken the optical unitout to access the process wheel, found that the cups were stuck in the process wheel, which is causing the process wheel obstruction. So removed those cups, and reconnected the machine. After restarting and after successsful initialization run 4 samples, results were fine. But after some time, the machine started giving the same error. So removed the cups in process wheel which were stuck, removed the solution pack after restarting the machine, found that the solution pack waste cup door was not functioning properly, so opened the solution pack top cover, fitted the waste cup door properly, after restarting the machine, done TropT new lot calibration, machine is working fine. | Defect | Consumables | SP | Cup Stack
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-151224 | No | Replace the test card | Replace the test card
ABL90 FLEX | Papierrolle gewechselt | Papierrolle gewechselt | SWO-150204 | No | Paper roll changed | Paper roll changed | User error | HW | Printer module | Paper replacement
ABL90 FLEX | Remplacement de la carte électronique 'sensor interface PCB' La calibration et les 3 CQ sont tous dans les tolérances. | Remplacement de la carte électronique 'sensor interface PCB' La calibration et les 3 CQ sont tous dans les tolérances. | SWO-150273 | No | Replacing the electronic card 'sensor interface PCB' The calibration and the 3 CQ are all in tolerances. | Replacing the electronic card 'sensor interface PCB' The calibration and the 3 CQ are all in tolerances. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-150064 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | pO2 | 375
ABL90 FLEX | 更换测试卡，试剂包正常。 | 更换测试卡，试剂包正常。 | SWO-150573 | No | Replace the test card and the reagent package is OK. | Replace the test card and the reagent package is OK.
ABL90 FLEX | 更换测试卡试剂包正常。 | 更换测试卡试剂包正常。 | SWO-149950 | No | Replacing the test card kit is OK. | Replacing the test card kit is OK.
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-149973 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | FT | 58
ABL90 FLEX | 更换测试卡，试剂包正常 | 更换测试卡，试剂包正常 | SWO-150909 | No | Replace the test card, the reagent package is OK | Replace the test card, the reagent package is OK | Assisted Replacement | SC | SC |  | 375
AQT90 FLEX | 更换903-364，903-236各一个，重新定标正常。(by zhenjie) | 更换903-364，903-236各一个，重新定标正常。(by zhenjie) | SWO-149680 | Yes | Replace the 903-364, 903-236 each one, re-calibration normal. (by zhenjie) | Replace the 903-364, 903-236 each one, re-calibration normal. (by zhenjie) | Re classified RMED | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换902-924一个正常。(by zhenjie) | 更换902-924一个正常。(by zhenjie) | SWO-149740 | No | Replace the 902-924 one normal. (by zhenjie) | Replace the 902-924 one normal. (by zhenjie) | Defect | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX | Raplace | Raplace | SWO-150127 | No | Raplace | Raplace | Defect | HW | Process Module | PCB, Cup dryer
ABL820 FLEX | Viewed logs - Unit rebooted on 11 11 and 11 All restarts were in the AM between 8am-10am. Gray ethernet cable is pinned between unit and wall. | Downloaded event logs for CQC case. Ran tests to verify proper calibrations. | D1B7F1907E00A7E2 | No | Viewed logs - Unit rebooted on 11 11 and 11 All restarts were in the AM between 8am-10am. Gray ethernet cable is pinned between unit and wall. | Downloaded event logs for CQC case. Ran tests to verify proper calibrations. | Unknown | SW | Random restart | Random restart
ABL817 FLEX | Verfied. | Found some debris in the bottom of the Crea B electrode. Also, re-membraned the Cl- and the Ref. electrodes. All per the service manual. Extended calibrations needed to get all parameters in. Customer has on order new dummy electrodes and a new Crea A electrode. | D1B805E67E23F8F7 | No | Verfied. | Found some debris in the bottom of the Crea B electrode. Also, re-membraned the Cl- and the Ref. electrodes. All per the service manual. Extended calibrations needed to get all parameters in. Customer has on order new dummy electrodes and a new Crea A electrode. | Cleaned | Fluidic pathway | Met II | Crea
ABL825 FLEX | CONFIRMED VIA PHONE WITH CUSTOMER. | SEND OUT NEW LAC ELECTRODE AND A BOX OF MEMBRANE FOR TROUBLESHOOTING. WILL MONITOR STATUS WITH TODD. | D1B83BB77E15ED56 | Yes | CONFIRMED VIA PHONE WITH CUSTOMER. | SEND OUT NEW LAC ELECTRODE AND A BOX OF MEMBRANE FOR TROUBLESHOOTING. WILL MONITOR STATUS WITH TODD. | Re classified RMED | Phone support
ABL800 BASIC | Performed the protein removal, found calibration results are OK, performed the QC found results are within the range, and explained the issue with sample error. It’s under Observation. | Performed the protein removal, found calibration results are OK, performed the QC found results are within the range, and explained the issue with sample error. It’s under Observation. | SWO-152816 | No | Performed the protein removal, found calibration results are OK, performed the QC found results are within the range, and explained the issue with sample error. It’s under Observation. | Performed the protein removal, found calibration results are OK, performed the QC found results are within the range, and explained the issue with sample error. It’s under Observation. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | FOUND CLOT LODGED IN WASTE NIPPLE. | CLEANED AND DECONTAMINATED SYSTEM. BLEW OUT WASTE LINES. | D1BAAC2C6A14221E | No | FOUND CLOT LODGED IN WASTE NIPPLE. | CLEANED AND DECONTAMINATED SYSTEM. BLEW OUT WASTE LINES. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | Removed the cups in process wheel which were stuck, removed the solution pack after restarting the machine, found that the solution pack waste cup door was not functioning properly, so opened the solution pack top cover, fitted the waste cup door properly, after restarting the machine run 2 samples, machine is working fine. | Removed the cups in process wheel which were stuck, removed the solution pack after restarting the machine, found that the solution pack waste cup door was not functioning properly, so opened the solution pack top cover, fitted the waste cup door properly, after restarting the machine run 2 samples, machine is working fine. | SWO-151178 | No | Removed the cups in process wheel which were stuck, removed the solution pack after restarting the machine, found that the solution pack waste cup door was not functioning properly, so opened the solution pack top cover, fitted the waste cup door properly, after restarting the machine run 2 samples, machine is working fine. | Removed the cups in process wheel which were stuck, removed the solution pack after restarting the machine, found that the solution pack waste cup door was not functioning properly, so opened the solution pack top cover, fitted the waste cup door properly, after restarting the machine run 2 samples, machine is working fine. | Defect | Consumables | SP | Cup Stack
AQT90 FLEX |  | B | D1B7F0227E5D2A67 | No |  |  | Defect | Consumables | LQC Kit | Myo
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-150328 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX | Analysator Überprüfung - Scannerglas gereinigt - Anwender hatte Analysator nochmals neu gestartet - Funktionsüberprüfung ohne Beanstandung! - Kunde meldet sich bei erneuten Problemen ! | Analysator Überprüfung - Scannerglas gereinigt - Anwender hatte Analysator nochmals neu gestartet - Funktionsüberprüfung ohne Beanstandung! - Kunde meldet sich bei erneuten Problemen ! | SWO-150384 | No | Analyzer verification - scanner glass cleaned - user had restarted analyzer again - function check without complaint! - Customer reports in case of new problems ! | Analyzer verification - scanner glass cleaned - user had restarted analyzer again - function check without complaint! - Customer reports in case of new problems ! | Cleaned | HW | Basic Unit | Barcode reader, display
ABL90 FLEX | Es wurde bereits mehrfach das LP ern. Einlassdichtung ern. - verschmutzt, Nadel überprüft - OK, Gerät kalibriert, QCs durchgef. Sollte der Fehler erneut auftreten erneuert der Kunde die Messkassette. | Frau Richter möchte nichts mehr machen und wünscht einen Techniker. Gerät zur Zeit in der Startphase. | SWO-150502 | No | The LP has already been re-energised several times. Inlet seal. - dirty, needle checked - OK, device calibrated, QCs mixed. If the error occurs again, the customer renews the measuring cassette. | Ms. Richter doesn't want to do anything anymore and wants a technician. device at the time of launch. | Defect | HW | Inlet module | Inlet Gasket
ABL90 | INSPECTED DATA LOGS AND SYSTEM FOR LEAKS OR FAILURES | REPLACED ALL TUBINGS, SEALS, AND GASKETS. PINCH VALVE SUSPECTED AND REPLACED. | D1BAA4F36A392219 | No | INSPECTED DATA LOGS AND SYSTEM FOR LEAKS OR FAILURES | REPLACED ALL TUBINGS, SEALS, AND GASKETS. PINCH VALVE SUSPECTED AND REPLACED. | Defect | HW | Oxi module | Pinch valve
ABL90 | checked power in service, power will not transfer from battery back to mains | replaced power supply | D1B92822F41E2F42 | No | checked power in service, power will not transfer from battery back to mains | replaced power supply | Defect | HW | Power Supply | Power Supply
ABL90 | battery test mains will not transfer back to main from battery | replaced power supply | D1B926CCF44E2F3D | No | battery test mains will not transfer back to main from battery | replaced power supply | Defect | HW | Power Supply | Power Supply
ABL90 FLEX PLUS | Inlet position check and calibration done | Inlet position check and calibration done | SWO-156975 | No | Inlet position check and calibration done | Inlet position check and calibration done | Adjustment | HW | Auto Inlet module | Inlet Position Cal.
ABL90 FLEX PLUS | Via het ziekenhuis netwerk zijn de instrumenten via een NAT router gekoppeld die beide analysers als één aan Radiance laat zien en dat geeft problemen. Dat zal de ziekenhuis ICT moeten oplossen. | B | SWO-150533 | No | Via the hospital network, the instruments are connected via a NAT router which shows both analyzers as one to Radiance and that gives problems. That will have to solve the hospital ICT. |  | Setup | His/lis | His/lis | His/lis
TCM CombiM | 更换431-018一个正常。(by zhenjie) | 更换431-018一个正常。(by zhenjie) | SWO-150101 | No | Replace 431-018 one normal. (by zhenjie) | Replace 431-018 one normal. (by zhenjie) | Defect | HW | Battery Module | Battery Module
TCM CombiM | 更换431-018一个正常。(by zhenjie) | 更换431-018一个正常。(by zhenjie) | SWO-152925 | No | Replace 431-018 one normal. (by zhenjie) | Replace 431-018 one normal. (by zhenjie) | Defect | HW | Battery Module | Battery Module
ABL90 FLEX | 更换903-183一个正常。(by zhenjie) | 更换903-183一个正常。(by zhenjie) | SWO-149764 | No | Replace 903-183 one normal. (by zhenjie) | Replace 903-183 one normal. (by zhenjie) | Defect | HW | Basic Unit | Sensor Interface PCB | 375
ABL800 BASIC | need the replacement of printer head unit. | need the replacement of printer head unit. | SWO-150598 | No | need the replacement of printer head unit. | need the replacement of printer head unit. | Defect | HW | Printer module | Printer head
ABL835 FLEX | Startphase QC und Kalibration. | Startphase QC und Kalibration. | SWO-150639 | No | Start phase QC and calibration. | Start phase QC and calibration. | Defect | HW | pH/BG Module | Valve, air
ABL800 BASIC | done the decontamination, now working fine also trained the staff why there is difference in electrolyte values from Lab | done the decontamination, now working fine also trained the staff why there is difference in electrolyte values from Lab | SWO-151141 | No | done the decontamination, now working fine also trained the staff why there is difference in electrolyte values from Lab | done the decontamination, now working fine also trained the staff why there is difference in electrolyte values from Lab | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber | Clot
ABL835 FLEX | Flachbandkabel richtig eingebaut, mehrfach getestet durch die erforderliche Elektrodenwartung, Einlassklappen im Betriebsmittelmodus angehoben und geprüft ob das Gerät die Neustart-Taste betätigen lässt, wenn zuvor auch ein Betriebsmittel eingescannt wurde --> Neustart-Taste lässt sich bedienen, Dichtring erneuert, System und Funktion i.O. | Flachbandkabel richtig eingebaut, mehrfach getestet durch die erforderliche Elektrodenwartung, Einlassklappen im Betriebsmittelmodus angehoben und geprüft ob das Gerät die Neustart-Taste betätigen lässt, wenn zuvor auch ein Betriebsmittel eingescannt wurde --> Neustart-Taste lässt sich bedienen, Dichtring erneuert, System und Funktion i.O. | SWO-150425 | No | Ribbon cable properly installed, tested several times by the required electrode maintenance, raised inlet flaps in operating mode and checked whether the device has the restart button pressed if an equipment has been scanned beforehand --> Restart button can be operated, sealing ring renewed, system and function in the o. | Ribbon cable properly installed, tested several times by the required electrode maintenance, raised inlet flaps in operating mode and checked whether the device has the restart button pressed if an equipment has been scanned beforehand --> Restart button can be operated, sealing ring renewed, system and function in the o. | Defect | HW | Inlet module | Cable
ABL827 FLEX |  | B | D1B95BF4F4626234 | Yes |  |  | Re classified RMED | Training
ABL815 FLEX |  | B | D1B9763CF41233FE | No |  |  | Defect | HW | DMS_PC | CPU | Freeze
AQT90 FLEX |  | B | D1BA5DC06A3D4CAB | No |  |  | Defect | Consumables | Test Kit | NT-proBNP | Cartridge/Cartridge 24 (Broken)
ABL80 FLEX BASIC | 更换903-001，902-985各一个正常。(by zhenjie) | 更换903-001，902-985各一个正常。(by zhenjie) | SWO-149970 | No | Replace 903-001, 902-985 one for each normal. (by zhenjie) | Replace 903-001, 902-985 one for each normal. (by zhenjie)
ABL90 FLEX | Lösungspack erneuert. Gerät Kalibriert, danach Qualitätskontrolle im Bereich. Kassette und Lösungspack reklamiert. Gerät überprüft. | Lösungspack erneuert. Gerät Kalibriert, danach Qualitätskontrolle im Bereich. Kassette und Lösungspack reklamiert. Gerät überprüft. | SWO-150966 | No | Solution pack renewed. Device calibrated, then quality control in the area. cassette and solution pack. device checked. | Solution pack renewed. Device calibrated, then quality control in the area. cassette and solution pack. device checked. | Assisted Replacement | SP | SP | pO2
ABL90 FLEX | LP & SC installiert-> Meldung Kal3 Lösung inhomogen -> gleichen LP erneut installiert -> Einlassdichtungswechsel wird empfohlen , alle Programme im Serviceprog. Fluid Transport durchgeführt -> Fehlermeldung in Solution Pack Check bleibt trotz mehrfacher Durchführung bestehen -> Startphase abgewartet -> Gerät nach 2h grün & messbereit | LP & SC installiert-> Meldung Kal3 Lösung inhomogen -> gleichen LP erneut installiert -> Einlassdichtungswechsel wird empfohlen , alle Programme im Serviceprog. Fluid Transport durchgeführt -> Fehlermeldung in Solution Pack Check bleibt trotz mehrfacher Durchführung bestehen -> Startphase abgewartet -> Gerät nach 2h grün & messbereit | SWO-150726 | No | LP & SC installed-> Message Kal3 solution inhomogeneous -> same LP reinstalled -> Inlet seal change is recommended, all programs in the service program. Fluid transport carried out -> Error message in Solution Pack Check remains despite multiple executions -> Start phase awaited -> Device after 2h green & ready to measure | LP & SC installed-> Message Kal3 solution inhomogeneous -> same LP reinstalled -> Inlet seal change is recommended, all programs in the service program. Fluid transport carried out -> Error message in Solution Pack Check remains despite multiple executions -> Start phase awaited -> Device after 2h green & ready to measure | Unknown | FT | FT | FT
ABL90 FLEX | Alle Serviceprog. Im Prog. Fluidtransport mehrfach durchgeführt bis keine Fehlermeldung mehr vorhanden waren -> danach besteht Gerät Startphase und ist nach 2h messbereit Gerät meldet 329 QC Verfallsdatum überschritten Los 613479 von LP Lot XT-03 Haltbar bis 01.02.2017 | Datalogs & Servicedump erstellt und per Mail an TS geschickt | SWO-150905 | No | All service prog. In the prog. Fluid transport carried out several times until no error message was present -> after that device is starting phase and is ready for measurement after 2h Device reports 329 QC expiry date exceeded lot 613479 of LP Lot XT-03 Durable until 01.02.2017 | Datalogs & Servicedump created and sent by mail to TS | Cleaned | Fluidic pathway | Fluidic pathway | FT
ABL90 FLEX | Gerät bei eintreffen grün und messbereit, Kunde hat LP & SC ersetzt, nach längerer Startphase war Gerät dann wieder messbereit Roma.ini Änderung Radiance -> Aqrue, CCF ausgedruckt und an TS geschickt, | Gerät bei eintreffen grün und messbereit, Kunde hat LP & SC ersetzt, nach längerer Startphase war Gerät dann wieder messbereit Roma.ini Änderung Radiance -> Aqrue, CCF ausgedruckt und an TS geschickt, | SWO-150906 | No | Device when green and ready to measure, customer has replaced LP & SC, after a long start phase, the device was ready for measurement again Roma.ini Change Radiance -> Aqrue, CCF printed and sent to TS, | Device when green and ready to measure, customer has replaced LP & SC, after a long start phase, the device was ready for measurement again Roma.ini Change Radiance -> Aqrue, CCF printed and sent to TS, | Assisted Replacement | SP | SP | FT
ABL90 FLEX PLUS |  | B | D1BA61BE6A12D728 | No |  |  | Assisted Replacement | SC | SC | Reference
ABL90 FLEX | 更换903-183后修复(by zhenjie) | 更换903-183后修复(by zhenjie) | SWO-150356 | No | Repair after replacement 903-183 (by zhenjie) | Repair after replacement 903-183 (by zhenjie) | Defect | HW | Basic Unit | Sensor Interface PCB | 375
AQT90 FLEX | 更换903-364一个正常。(by zhenjie) | 更换903-364一个正常。(by zhenjie) | SWO-150358 | No | Replace 903-364 one normal. (by zhenjie) | Replace 903-364 one normal. (by zhenjie) | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换580-098后修复(by zhenjie) | 更换580-098后修复(by zhenjie) | SWO-150349 | No | Repair after replacement 580-098 (by zhenjie) | Repair after replacement 580-098 (by zhenjie) | Defect | HW | Process Module | Stepper motor, process module | 5784
AQT90 FLEX | 更换903-127，636-538各一个正常。(by zhenjie) | 更换903-127，636-538各一个正常。(by zhenjie) | SWO-150895 | No | Replace 903-127, 636-538 each normal. (by zhenjie) | Replace 903-127, 636-538 each normal. (by zhenjie) | Defect | HW | Storage Module | PCB, Storage wheel
AQT90 FLEX | Reiniciado el equipo. Comprobado buen funcionamiento del equipo.. | B | D1BA8EA06A22F949 | No | Restart the computer. Proven proper equipment operation. |  | Defect | SW | Analyzer Control Module | Off/On
AQT90 FLEX | 更换903-210后修复(by zhenjie) | 更换903-210后修复(by zhenjie) | SWO-150319 | No | Repair after replacement 903-210 (by zhenjie) | Repair after replacement 903-210 (by zhenjie) | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL800 BASIC | Gerät geprüft. AC Modul ausgebaut und überprüft. Flachbandkabel erneuert. Startphase abgewartet. Kal und QC durchgeführt. | FSE-Einsatz erforderlich | SWO-150979 | No | device checked. AC module removed and checked. ribbon cable is renewed. start-up phase. Kal and QC. | FSE deployment required | Defect | HW | Auto Check Module | FPC Cable
ABL825 FLEX |  | B | D1BAE3266A4E9AD8 | No |  |  | Assisted Replacement | Electrode | Electrode | Sealing Ring
ABL815 FLEX | Gas 2 - Ventil ausgetauscht 2-Pkt.-Kalibrierung | Gas 2 - Ventil ausgetauscht 2-Pkt.-Kalibrierung | SWO-150782 | No | Gas 2 - Valve replaced 2-pk.-calibration | Gas 2 - Valve replaced 2-pk.-calibration | Defect | HW | Basic Unit | Sirai Valve 2
ABL825 FLEX | Customer tried new cl and ref electrode and membranes. | None. Ran ABL90 side by side to compare.Ran 10 samples 1-2 mmol lower on ABL90 but still both higher than Roche | D1BB9CB9E017BED1 | No | Customer tried new cl and ref electrode and membranes. | None. Ran ABL90 side by side to compare. Ran 10 samples 1-2 mmol lower on ABL90 but still both higher than Roche | Unknown | Electrode | Electrode | Cl
ABL827 FLEX | Crea out on cals. Replaced membranes and then all bottles. | Calibrations finally in after membranes 2nd time. | D1BAC68A6A63E80C | No | Crea out on cals. Replaced membranes and then all bottles. | Calibrations finally in after membranes 2nd time. | Assisted Replacement | Electrode | Membrane | Crea
ABL825 FLEX | Daten gesichert und TSP weitergeleitet | B | SWO-150752 | No | Data backed up and TSP forwarded |  | Revisit required | HW | Oxi module | tHb cal | Email to Boris
ABL825 FLEX | Key eingespielt. | Key eingespielt. | SWO-151319 | No | Key. | Key. | Setup | SW | Option key | Option key | tHb
TCM-COMBIM 84 | None CPU, Communications board | Replaced Power supply. Boxed and shipped to Brea CA for depot repair. | D1BAA4556A581BC5 | No | None CPU, Communications board | Replaced Power supply. Boxed and shipped to Brea CA for depot repair. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | LP erneuert | LP erneuert | SWO-150996 | No | LP renewed | LP renewed | Assisted Replacement | SP | SP | Unknown
ABL90 FLEX | Flow Selector gereinigt, LP eingelegt, Gerät kalibriert, QCs durchgef. | Flow Selector gereinigt, LP eingelegt, Gerät kalibriert, QCs durchgef. | SWO-151314 | No | Flow Selector cleaned, LP inserted, device calibrated, QCs worked through. | Flow Selector cleaned, LP inserted, device calibrated, QCs worked through. | Cleaned | HW | Basic Unit | Flow selector
ABL90 |  | B | D1BBDDB5E05785AD | No |  |  | Assisted Replacement | SP | SP | Pump Tube
ABL90 FLEX PLUS | Allgemeine Kontrolle ok. Servicedump und Dataexport rausgezogen ok. Hemolyzer ersetzt ok. System i.o. Kal. + QC + Thb i.o. | B | SWO-150989 | No | General control ok. Service dump and data export pulled out ok. Hemolyzer replaces ok. System i.o. Kal. + QC + Thb i.o. |  | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX |  | B | D1D2A2202C5BBFD6 | No |  |  | Defect | HW | Oxi module | Oxi module
ABL800 FLEX | Replace the membrane & calibration solution. Clean the chamber. Ca chamber leakage causing the calibration error. need to replace the chamber. | Replace the membrane & calibration solution. Clean the chamber. Ca chamber leakage causing the calibration error. need to replace the chamber. | SWO-151128 | No | Replace the membrane & calibration solution. Clean the chamber. Ca chamber leakage causing the calibration error. need to replace the chamber. | Replace the membrane & calibration solution. Clean the chamber. Ca chamber leakage causing the calibration error. need to replace the chamber. | Assisted Replacement | Electrode | Membrane | Ca
ABL825 FLEX | tried to scan 4 botltles | Looked at calibration logs. Noticed several rarameters not installed. Checked option keys. Original option keys appears corrupted (does not contain a blood gas key). Contacted mark Ziegler to have original keys resent. | D1BB1EFBE0411EF3 | No | tried to scan 4 botltles | Looked at calibration logs. Noticed several rarameters not installed. Checked option keys. Original option keys appears corrupted (does not contain a blood gas key). Contacted mark Ziegler to have original keys resent. | Defect | SW | Analyzer control PCB | Key | LEH Mail
ABL825 FLEX | Tried several rinse bottles. | Looked at calibration logs. Noticed several rarameters not installed. Checked option keys. Original option keys appears corrupted (does not contain a blood gas key). Contacted mark Ziegler to have original keys resent. | D1BB20BDE05B1EF6 | No | Tried several rinse bottles. | Looked at calibration logs. Noticed several rarameters not installed. Checked option keys. Original option keys appears corrupted (does not contain a blood gas key). Contacted mark Ziegler to have original keys resent. | Defect | SW | Analyzer control PCB | Key | LEH Mail
ABL825 FLEX |  | B | D1BBAAA6E0197FEA | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Monitor ersetzt | Monitor ersetzt | SWO-151093 | No | Monitor replaced | Monitor replaced | Defect | HW | DMS_PC | Touch screen
ABL810 FLEX | Anbindungsdaten laut Herrn Nobel eingegeben. Anbindung mit OSM hergestellt. Datensendung eingestellt. QC nachgesendet. Telefonische Rücksprache mit Herrn Nobel. Übertragung ohne Beanstandung Installation der finalen Keys. Weiter Einweisungschritte erläutert | Der Kunde bittet um einen zeitnahen Rückruf. um abzuklären, wann die Netzwerkanbindung abgeschlossen wird, bzw. um noch offene Fragen (Rücknahme ABL5) zu besprechen. | SWO-151022 | No | connection data according to Mr. Nobel. connection with OSM. data transmission. QC. Telephone consultation with Mr Nobel. Transmission without complaint Installation of the final keys. Further instruction steps explained | The customer asks for a timely recall. to clarify when the network connection is completed or to discuss outstanding issues (withdrawal ABL5). | Setup | SW | His/lis | His/lis
ABL90 FLEX | Pendant inspection du joint d'intro et l'aiguille j'ai constaté que le module d'intro (mécanisme qui pousse le joint contre le pack de solution) ne fonctionne pas bon. Le resort interne qui doit pousser le joint est pas assez fort : à remplacer. Pour l'instant l'appareil fonctionne (les test en service sont OK) mais cette problème peux revenir! Calibrations et CQ sont OK | Pendant inspection du joint d'intro et l'aiguille j'ai constaté que le module d'intro (mécanisme qui pousse le joint contre le pack de solution) ne fonctionne pas bon. Le resort interne qui doit pousser le joint est pas assez fort : à remplacer. Pour l'instant l'appareil fonctionne (les test en service sont OK) mais cette problème peux revenir! Calibrations et CQ sont OK | SWO-151309 | No | During inspection of the intro seal and needle I found that the intro module (mechanism that pushes the joint against the solution pack) is not working well. The internal resort that must push the joint is not strong enough: to replace. For now the device is working (the in-service tests are OK) but this problem can come back! Calibrations and CQ are OK | During inspection of the intro seal and needle I found that the intro module (mechanism that pushes the joint against the solution pack) is not working well. The internal resort that must push the joint is not strong enough: to replace. For now the device is working (the in-service tests are OK) but this problem can come back! Calibrations and CQ are OK | Defect | HW | Inlet module | Inlet module | Spring
ABL90 FLEX | 更换903-644后修复(by zhenjie) | 更换903-644后修复(by zhenjie) | SWO-150899 | No | Repair after replacement 903-644 (by zhenjie) | Repair after replacement 903-644 (by zhenjie) | Defect | HW | Basic Unit | Wet sec PCB
ABL90 FLEX PLUS | Checked service mode for any causes and found nothing investigated tubings and found nothing. Done sesnsor simulation test all ok. Could not find cause. | B | SWO-151755 | No | Checked service mode for any causes and found nothing investigated tubings and found nothing. Done sesnsor simulation test all ok. Could not find cause. |  | Revisit required | Revisit required | Revisit required | Revisit required | wetability?
ABL90 FLEX PLUS | Stripped analyzer and replaced tubings(yearly service) restarted analyzer - passed calibrations and QC's on investigation looks like the waste line tubing had small pin hole in it at the solution pack connector side. | Stripped analyzer and replaced tubings(yearly service) restarted analyzer - passed calibrations and QC's on investigation looks like the waste line tubing had small pin hole in it at the solution pack connector side. | SWO-151960 | No | Stripped analyzer and replaced tubings(yearly service) restarted analyzer - passed calibrations and QC's on investigation looks like the waste line tubing had small pin hole in it at the solution pack connector side. | Stripped analyzer and replaced tubings(yearly service) restarted analyzer - passed calibrations and QC's on investigation looks like the waste line tubing had small pin hole in it at the solution pack connector side. | Defect | HW | PM Kit | PM Kit
AQT90 FLEX | 更换903-364后修复---jack.chu | 更换903-364后修复---jack.chu | SWO-151137 | Yes | Repair ---jack.chu after replacing 903-364 | Repair ---jack.chu after replacing 903-364 | Re classified RMED | HW | Process Module | Optical Unit | 124
ABL80 FLEX BASIC | 更换902-972一个正常。---jack.chu | B | SWO-151235 | No | Replace 902-972 one normal. ---jack.chu
ABL80 | checked data logs for errors, checked QC log, Red qc upper range is 69 instrument reading 71-74mmHg | checked flow of solutions and samples, performed Set0 adjustment, increased upper range of control | D1C00208421EA27B | No | checked data logs for errors, checked QC log, Red qc upper range is 69 instrument reading 71-74mmHg | checked flow of solutions and samples, performed Set0 adjustment, increased upper range of control
ABL90 FLEX | Lp Wechsel, Wechsel von ELD und Sonde, Neustart | Lp Wechsel, Wechsel von ELD und Sonde, Neustart | SWO-151373 | No | Lp Change, change of ELD and probe, restart | Lp Change, change of ELD and probe, restart | User error | SP installation | SP installation | SP installation
ABL90 |  | B | D1BF811A4230FA71 | Yes
ABL90 FLEX | Überprüfung Analysator - Clot Detection Status fehlerhaft - Analysator mehrfach gespült - Service Flüssigkeitstransport durchgeführt - QC Verletzungen weiterhin vorhanden - Erneuerung Lösungspack - Funktionsüberprüfung etc. ohne Beanstandung! | Überprüfung Analysator - Clot Detection Status fehlerhaft - Analysator mehrfach gespült - Service Flüssigkeitstransport durchgeführt - QC Verletzungen weiterhin vorhanden - Erneuerung Lösungspack - Funktionsüberprüfung etc. ohne Beanstandung! | SWO-151516 | No | Review Analyzer - Clot Detection Status Faulty - Analyzer Multiple Rinsed - Service Liquid Transport Performed - QC Injuries Still Present - Renewal Solution Pack - Functional Check etc. without complaint! | Review Analyzer - Clot Detection Status Faulty - Analyzer Multiple Rinsed - Service Liquid Transport Performed - QC Injuries Still Present - Renewal Solution Pack - Functional Check etc. without complaint! | Assisted Replacement | SP | SP | QC measurement
ABL90 FLEX | Analyser samples of 50 per day, found that failed to aspirate sample error had preceded both previous changes and intervention and would lead to advised action of replace SP and usage of activities | Analyser samples of 50 per day, found that failed to aspirate sample error had preceded both previous changes and intervention and would lead to advised action of replace SP and usage of activities | SWO-151470 | No | Analyser samples of 50 per day, found that failed to aspirate sample error had preceded both previous changes and intervention and would lead to advised action of replace SP and usage of activities | Analyser samples of 50 per day, found that failed to aspirate sample error had preceded both previous changes and intervention and would lead to advised action of replace SP and usage of activities | User error | Sample Handling | Sample Handling | Sample Handling
ABL90 FLEX | Serviceprog. Fluidtransport mehrfach durchgeführt, Gerät besteht Startphase, Uhrzeit korrigiert | Serviceprog. Fluidtransport mehrfach durchgeführt, Gerät besteht Startphase, Uhrzeit korrigiert | SWO-152150 | No | Serviceprog. Fluid transport carried out several times, device passes start phase, time corrected | Serviceprog. Fluid transport carried out several times, device passes start phase, time corrected | Cleaned | Fluidic pathway | Fluidic pathway | FT
ABL90 |  | B | D1BF8EC5425BBF04 | No |  |  | Assisted Replacement | SP | SP | SP
ABL90 FLEX PLUS |  | B | D1BFA43042136179 | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Analyser now passing Cals and QCs, inspected pins just in case, all ok. | Analyser now passing Cals and QCs, inspected pins just in case, all ok. | SWO-151568 | No | Analyser now passing Cals and QCs, inspected pins just in case, all ok. | Analyser now passing Cals and QCs, inspected pins just in case, all ok. | No fault found | No fault found | No fault found | No fault found | Ca++
ABL90 FLEX PLUS | Stripped down analyser to reveal the actuator connecting rod had become unscrewed. Reattached con rod and reassembled. Tested all ok | Stripped down analyser to reveal the actuator connecting rod had become unscrewed. Reattached con rod and reassembled. Tested all ok | SWO-151571 | No | Stripped down analyser to reveal the actuator connecting rod had become unscrewed. Reattached con rod and reassembled. Tested all ok | Stripped down analyser to reveal the actuator connecting rod had become unscrewed. Reattached con rod and reassembled. Tested all ok | Adjustment | HW | Basic Unit | Actuator rod
ABL90 FLEX | 更换测试卡修复 | 更换测试卡修复 | SWO-151384 | No | Replacement test card repair | Replacement test card repair
ABL90 FLEX |  | B | SWO-151214 | No |  |  | Assisted Replacement | SC | SC | Temperature
AQT90 FLEX |  | B | SWO-151250 | No |  |  | Defect | HW | Incubator module | Punch up Unit | 5297
AQT90 FLEX |  | B | SWO-151225 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | 1218
AQT90 FLEX | 更换备用件636-541修复---jack.chu | 更换备用件636-541修复---jack.chu | SWO-151278 | No | Replace spare parts 636-541 repair ---jack.chu | Replace spare parts 636-541 repair ---jack.chu | Defect | HW | Analyzer Control Module | Cable | 5784
ABL800 BASIC | AC Modul ausgebaut und erneuert. Startphase abgewartet. Analysator startet mit Fehler Spectrometer AD Converter Failed ! - Analysator ausgeschaltet und nach 20 Min. neugestartet. Kal und QC durchgeführt. | AC Modul ausgebaut und erneuert. Startphase abgewartet. Analysator startet mit Fehler Spectrometer AD Converter Failed ! - Analysator ausgeschaltet und nach 20 Min. neugestartet. Kal und QC durchgeführt. | SWO-151612 | No | AC module expanded and renewed. start-up phase. Analyzer starts with error Spectrometer AD Converter Failed ! - Analyzer turned off and restarted after 20 minutes. Kal and QC. | AC module expanded and renewed. start-up phase. Analyzer starts with error Spectrometer AD Converter Failed ! - Analyzer turned off and restarted after 20 minutes. Kal and QC. | Defect | HW | Auto Check Module | AutoCheck Module | Misaligned Barcode reader
ABL810 FLEX | Ablauf wird nicht befördert. Fließwege gespült, Einlassnadel ist verbogen - erneuert, Ablauf Rollenkranz blutig - gereinigt. Die Rollen sind gängig dennoch kommen hin und wieder Flüssigkeitsreste aus dem Einlass heraus - Rollen blockieren zwischendurch dadurch wird die Flüssigkeit nicht ordentlich transportiert. Rollenkranz gereinigt - Situation wird beobachtet - neuer Rollenkranz wird bestellt. QC's durchgeführt. | Ablauf wird nicht befördert. Fließwege gespült, Einlassnadel ist verbogen - erneuert, Ablauf Rollenkranz blutig - gereinigt. Die Rollen sind gängig dennoch kommen hin und wieder Flüssigkeitsreste aus dem Einlass heraus - Rollen blockieren zwischendurch dadurch wird die Flüssigkeit nicht ordentlich transportiert. Rollenkranz gereinigt - Situation wird beobachtet - neuer Rollenkranz wird bestellt. QC's durchgeführt. | SWO-151616 | No | Expiration is not transported. Flow paths rinsed, inlet needle is bent - renewed, drain roller wreath bloody - cleaned. The rollers are common but occasionally liquid residues come out of the inlet - rollers block in between, so the liquid is not transported properly. Roller wreath cleaned - situation is observed - new roller wreath is ordered. QC's. | Expiration is not transported. Flow paths rinsed, inlet needle is bent - renewed, drain roller wreath bloody - cleaned. The rollers are common but occasionally liquid residues come out of the inlet - rollers block in between, so the liquid is not transported properly. Roller wreath cleaned - situation is observed - new roller wreath is ordered. QC's. | Defect | HW | Inlet module | Probe bend
ABL810 FLEX | Tausch der pCO2 Elektrode. Messkammer gereinigt und Elektrodenkontakte gebürstet. Kalibrations Werte wieder in Ordnung. Empfindlichkeit liegt bei 94%. Kalibrationen und Qualitätskontrolle in Ordnung. | Tausch der pCO2 Elektrode. Messkammer gereinigt und Elektrodenkontakte gebürstet. Kalibrations Werte wieder in Ordnung. Empfindlichkeit liegt bei 94%. Kalibrationen und Qualitätskontrolle in Ordnung. | SWO-151563 | No | Replacement of the pCO2 electrode. measuring chamber cleaned and electrode contacts brushed. Calibration values are fine again. Sensitivity is 94%. Calibrations and quality control in order. | Replacement of the pCO2 electrode. measuring chamber cleaned and electrode contacts brushed. Calibration values are fine again. Sensitivity is 94%. Calibrations and quality control in order. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Looked at system messages.0247, Spectrometer AD converter failed, 0258, Neon lamp intesity to low, 0316, neon number to low, 0248 Spectrometer missing trigger. | rebooted analyzer, did complete oximetry. | D1BF86134243770E | No | Looked at system messages.0247, Spectrometer AD converter failed, 0258, Neon lamp intesity to low, 0316, neon number to low, 0248 Spectrometer missing trigger. | rebooted analyzer, did complete oximetry. | Adjustment | HW | Oxi module | Off/On | Missing trigger
ABL800 BASIC | ｐH電極、交換。 2ﾎﾟｲﾝﾄｷｬﾘﾌﾞﾚｰｼｮﾝ実施、OK。 QCﾁｪｯｸ、OK。 | B | SWO-151586 | No | pH electrodes, replacement. Two-point calibration implementation, OK. QC check, OK. |  | Assisted Replacement | Electrode | Electrode | pH
ABL815 FLEX | Demokey eingelesen. | Demokey eingelesen. | SWO-151318 | No | Demokey. | Demokey. | Setup | SW | Demo key exp | Demo key exp
ABL800 BASIC | Dismantled the touch screen and cleaned it I.P.A. Assembled the display unit and re-installed. Checked and found no options opening automatically. Advised the Customer to use only Ethanol or Isopropyl alcohol to clean the display. | Dismantled the touch screen and cleaned it I.P.A. Assembled the display unit and re-installed. Checked and found no options opening automatically. Advised the Customer to use only Ethanol or Isopropyl alcohol to clean the display. | SWO-152734 | No | Dismantled the touch screen and cleaned it I.P.A. Assembled the display unit and re-installed. Checked and found no options opening automatically. Advised the Customer to use only Ethanol or Isopropyl alcohol to clean the display. | Dismantled the touch screen and cleaned it I.P.A. Assembled the display unit and re-installed. Checked and found no options opening automatically. Advised the Customer to use only Ethanol or Isopropyl alcohol to clean the display. | Cleaned | HW | DMS_PC | Touch screen
ABL825 FLEX | checked data logs for errors, checked QC logs, checked hemolyzer and lamp unit | performed lamp adjust, performed tHb cal in service, QC failing | D1C00DB04205C581 | No | checked data logs for errors, checked QC logs, checked hemolyzer and lamp unit | performed lamp adjust, performed tHb cal in service, QC failing | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX |  | B | D1BF8A98420A1763 | No |  |  | Defect | Consumables | SP | Cup Stack | Blocked
ABL80 CO-OX | Many times, the sensor will fail the initial startup initialization and often take an hour or two to start working. | Replaced the Analyzer Control Board. Performed the zero alignment, tHb Calibration and a System Check. | D1C0BA3CB8499F30 | No | Many times, the sensor will fail the initial startup initialization and often take an hour or two to start working. | Replaced the Analyzer Control Board. Performed the zero alignment, tHb Calibration and a System Check.
ABL90 FLEX | Touch Screen erneuert Kalibration und QC durchgeführt | Touch Screen erneuert Kalibration und QC durchgeführt | SWO-151812 | No | Touch Screen renews calibration and QC performed | Touch Screen renews calibration and QC performed | Defect | HW | DMS_PC | Touch screen
ABL90 |  | B | D1C0ECD5B8633524 | No |  |  | Adjustment | SW | SW | Off/On
ABL90 FLEX PLUS |  | B | D1C10347B83904FE | No |  |  | Assisted Replacement | SC | SC | Glu/Lac
ABL90 FLEX PLUS |  | B | D1C0DDAAB8000336 | No |  |  | Cleaned | Fluidic pathway | Oxi module | Hemolyzer - clot
ABL90 FLEX PLUS |  | B | D1C1E53B2E63A47F | No |  |  | Assisted Replacement | SP | SP | Leaking
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-151618 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Glu
AQT90 FLEX | 更换903-364一个正常。---jack.chu | 更换903-364一个正常。---jack.chu | SWO-151240 | Yes | Replace 903-364 one normal. ---jack.chu | Replace 903-364 one normal. ---jack.chu | Re classified RMED | HW | Process Module | Optical Unit | 131
ABL835 FLEX | Had customer replace pump tubing | Customer did not have inlet tubing. Replaced inlet tubing. | D1C0FC38B85CDCC7 | Yes | Had customer replace pump tubing | Customer did not have inlet tubing. Replaced inlet tubing. | Defect | HW | Inlet module | Inlet Tubing
ABL835 FLEX | Einweisung der Personal vom Kreißsaal. Teilnehmer sind in einem Einweisungsprotokoll aufgelistet. | Einweisung der Personal vom Kreißsaal. Teilnehmer sind in einem Einweisungsprotokoll aufgelistet. | SWO-151862 | Yes | Instruction of the staff from the Kreißsaal. Participants are listed in a briefing log. | Instruction of the staff from the Kreißsaal. Participants are listed in a briefing log. | Re classified RMED | Training
AQT90 FLEX | refix the slider with inlet door and observed for 1.5 hour now it is working fine. | refix the slider with inlet door and observed for 1.5 hour now it is working fine. | SWO-152770 | No | refix the slider with inlet door and observed for 1.5 hour now it is working fine. | refix the slider with inlet door and observed for 1.5 hour now it is working fine. | Adjustment | HW | Cover and Chassis | Inlet Door
ABL80 FLEX Basic version |  | B | D1C206DD2E31192C | No
ABL90 |  | B | D1C20CD12E3F1CC6 | Yes
ABL90 FLEX | Gerät überprüft, Einstellungen Windowsebene kontrolliert | Gerät überprüft, Einstellungen Windowsebene kontrolliert | SWO-152129 | No | Device checked, windows level settings controlled | Device checked, windows level settings controlled | Setup | SW | Time setting | Time setting
ABL90 | user concerned the cassette doesnot respond to change either | rebooted instrument, unexpected problem error occurred, rebooted exercised all functions | D1C21DAB2E4C42B0 | No | user concerned the cassette doesnot respond to change either | rebooted instrument, unexpected problem error occurred, rebooted exercised all functions | Adjustment | SW | SW | Off/On
ABL90 FLEX | Fehlersuche. LP getauscht --> ohne Erfolg. SC erneuert. | Fehlersuche. LP getauscht --> ohne Erfolg. SC erneuert. | SWO-152156 | No | Troubleshooting. LP exchanged --> to no avail. SC renewed. | Troubleshooting. LP exchanged --> to no avail. SC renewed. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | Datalogs & Servicedump erstellt und per Mail an TS geschickt | Datalogs & Servicedump erstellt und per Mail an TS geschickt | SWO-152144 | No | Datalogs & Servicedump created and sent by mail to TS | Datalogs & Servicedump created and sent by mail to TS | Revisit required | Fluidic pathway | Fluidic pathway | FT
ABL90 FLEX PLUS | Data taken for tech line. | customer can resolve by logging in as super user and closing inlet in service. escalated to dave, as may be motor bushing? | SWO-152115 | No | Data taken for tech line. | customer can resolve by logging in as super user and closing inlet in service. escalated to dave, as may be motor bushing? | Adjustment | HW | Auto Inlet module | Inlet Position Cal. | Techline
ABL90 FLEX | 更换试剂包修复 | 更换试剂包修复 | SWO-152029 | No | Replacement reagent kit repair | Replacement reagent kit repair
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-151991 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Lac | 378
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-152228 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Cl
ABL800 BASIC | Replaced the oxi-module, after restarting the machine. Done the complete oximetry test, run protein removal, replaced the Lac and Ca membrane. After a successful 2 point calibration, run 3 samples, machine is working fine. | Replaced the oxi-module, after restarting the machine. Done the complete oximetry test, run protein removal, replaced the Lac and Ca membrane. After a successful 2 point calibration, run 3 samples, machine is working fine. | SWO-152784 | No | Replaced the oxi-module, after restarting the machine. Done the complete oximetry test, run protein removal, replaced the Lac and Ca membrane. After a successful 2 point calibration, run 3 samples, machine is working fine. | Replaced the oxi-module, after restarting the machine. Done the complete oximetry test, run protein removal, replaced the Lac and Ca membrane. After a successful 2 point calibration, run 3 samples, machine is working fine. | Defect | HW | Oxi module | Oxi module | Lamp Unit should have been replaced
ABL825 FLEX | AC Modul demontiert und gereinigt, Funktionskontrolle ok | AC Modul demontiert und gereinigt, Funktionskontrolle ok | SWO-152265 | No | AC module disassembled and cleaned, function control ok | AC module disassembled and cleaned, function control ok | Cleaned | HW | Auto Check Module | Barcode reader, AC
ABL800 BASIC | Manually done up down movement of the inlet probe,after that run the inlet calibration.after that instrument was in the Ready status,there after turn off the instrument for 2 hours. After switched on,run the 2point cal and oximetry parameters became green. | Manually done up down movement of the inlet probe,after that run the inlet calibration.after that instrument was in the Ready status,there after turn off the instrument for 2 hours. After switched on,run the 2point cal and oximetry parameters became green. | SWO-152760 | No | Manually done up down movement of the inlet probe,after that run the inlet calibration.after that instrument was in the Ready status,there after turn off the instrument for 2 hours. After switched on,run the 2point cal and oximetry parameters became green. | Manually done up down movement of the inlet probe,after that run the inlet calibration.after that instrument was in the Ready status,there after turn off the instrument for 2 hours. After switched on,run the 2point cal and oximetry parameters became green. | Adjustment | HW | Inlet module | Inlet module | Revisit required??
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-152520 | No | Replacement test card repair | Replacement test card repair
ABL80 FLEX BASIC | Checked and performed zeroing of all electrode channels, flushed sensor cassette with normal saline, performed 2 point calibration successfully. Suggested to repeat sample measurement to the customer. | Checked and performed zeroing of all electrode channels, flushed sensor cassette with normal saline, performed 2 point calibration successfully. Suggested to repeat sample measurement to the customer. | SWO-152771 | No | Checked and performed zeroing of all electrode channels, flushed sensor cassette with normal saline, performed 2 point calibration successfully. Suggested to repeat sample measurement to the customer. | Checked and performed zeroing of all electrode channels, flushed sensor cassette with normal saline, performed 2 point calibration successfully. Suggested to repeat sample measurement to the customer.
ABL90 FLEX | LP & SC ersetzt, gerät fordert zur LP wechsel, Einlassdichtungshalter wechsle & zum Durchspülen auf, alle Serviceprog. im Fluidtransport Serviceprogramm mehrfach durchgeführt, Uhrzeit korrgiert, Konfig. eingespielt, an AQURE angebunden und Test QC & Messung durchgeführt & gesendet | LP & SC ersetzt, gerät fordert zur LP wechsel, Einlassdichtungshalter wechsle & zum Durchspülen auf, alle Serviceprog. im Fluidtransport Serviceprogramm mehrfach durchgeführt, Uhrzeit korrgiert, Konfig. eingespielt, an AQURE angebunden und Test QC & Messung durchgeführt & gesendet | SWO-152430 | No | LP & SC replaces, comes to the LP change, inlet seal holder changes & rinses, all service prog. performed several times in the Fluid Transport service program, time correlated, config. recorded, connected to AQURE and test QC & measurement performed & sent | LP & SC replaces, comes to the LP change, inlet seal holder changes & rinses, all service prog. performed several times in the Fluid Transport service program, time correlated, config. recorded, connected to AQURE and test QC & measurement performed & sent | Setup | SW | SW | Time setting
ABL90 FLEX PLUS |  | B | D1C35242A43FC7CD | No |  |  | Defect | HW | DMS_PC | CF card
ABL90 FLEX PLUS |  | B | D1C3CEC5A4628787 | Yes |  |  | Re classified RMED | Setup New install
ABL90 FLEX | 更换测试卡修复 | 更换测试卡修复 | SWO-152329 | No | Replacement test card repair | Replacement test card repair
ABL90 FLEX | 更换测试卡修复 | 更换测试卡修复 | SWO-152288 | No | Replacement test card repair | Replacement test card repair
AQT90 FLEX |  | B | SWO-152279 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX |  | B | D1C39566A431F967 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL827 FLEX | found debris in module II air valve | installed pm kit changing all tubing. cleaned valves, ran thb cal, remembraned crea | D1C359AAA446D07C | No | found debris in module II air valve | installed pm kit changing all tubing. cleaned valves, ran thb cal, remembraned crea | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL827 FLEX | on 9 30 i sent lac and cl electrode to customer, 12 13 replaced the reference electrode | changed electrodes | D1C3599FA42CD07B | No | on 9 30 i sent lac and cl electrode to customer, 12 13 replaced the reference electrode | changed electrodes | Assisted Replacement | Electrode | Electrode | Reference
TCM-COMBIM 84 | Verify issue. | Reconfigure TCM. | D1C345FEA45886C7 | No | Verify issue. | Reconfigure TCM. | Defect | SW | Communication
ABL80 FLEX BASIC |  | B | SWO-152534 | No
ABL80 FLEX BASIC | 更换试剂包修复 | 更换试剂包修复 | SWO-152852 | No | Replacement reagent kit repair | Replacement reagent kit repair
ABL80 FLEX BASIC |  | B | SWO-152523 | No
ABL80 FLEX CO-OX | Fehlersuche, Reinigung des Gerätes, Neuinstallation Lösungspack, Funktionskontrolle | B | SWO-152655 | No | Troubleshooting, cleaning the device, reinstallation solution pack, function control
ABL90 FLEX | LP Wechsel | LP Wechsel | SWO-152640 | No | LP Change | LP Change | User error | SP | SP | Empty
ABL90 FLEX | Kalibration durchgeführt | Kalibration durchgeführt | SWO-152643 | No | Calibration carried out | Calibration carried out | Assisted Replacement | SC | SC | SC
ABL90 | Analyzer is not switching | Replaced Power supply. | D1C60833900B1883 | No | Analyzer is not switching | Replaced Power supply. | Defect | HW | Power Supply | Power Supply
ABL90 | Analyzer not switching | Replaced Power Supply | D1C609B490511887 | No | Analyzer not switching | Replaced Power Supply | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | ABL mittels dem Sensor Simulator überprüft, Test in allen Punkten bestanden, Sensorkassette neu initialisiert, Kal. und QC bestanden, System- und Funktionskontrolle in Ordnung | ABL mittels dem Sensor Simulator überprüft, Test in allen Punkten bestanden, Sensorkassette neu initialisiert, Kal. und QC bestanden, System- und Funktionskontrolle in Ordnung | SWO-152738 | No | ABL checked by means of the sensor simulator, passed test in all points, sensor cartridge reinitialized, Kal. and QC, system and function control in order | ABL checked by means of the sensor simulator, passed test in all points, sensor cartridge reinitialized, Kal. and QC, system and function control in order | Unknown | SC | SC | SC
ABL90 FLEX PLUS | Fitted new power supply and test ok. Battery P0128N035 | B | SWO-152665 | No | Fitted new power supply and test ok. Battery P0128N035 |  | Defect | HW | Power Supply | Power Supply
ABL90 FLEX PLUS | Fitted new power supply unit and test ok. Switches from battery to mains ok. Battery P0130N021 | B | SWO-152661 | No | Fitted new power supply unit and test ok. Switches from battery to mains ok. Battery P0130N021 |  | Defect | HW | Power Supply | Power Supply
ABL90 FLEX PLUS | IP Adresse eingestellt. Das Gerät in AQURE konfiguriert. In Betrieb genommen. | IP Adresse eingestellt. Das Gerät in AQURE konfiguriert. In Betrieb genommen. | SWO-152729 | No | IP address set. The device is configured in AQURE. Put into operation. | IP address set. The device is configured in AQURE. Put into operation. | Setup | His/lis | His/lis | His/lis
ABL90 FLEX | 更换测试卡修复 | 更换测试卡修复 | SWO-152846 | No | Replacement test card repair | Replacement test card repair
AQT90 FLEX |  | B | SWO-152580 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364一个正常。---jack.chu | 更换903-364一个正常。---jack.chu | SWO-152547 | Yes | Replace 903-364 one normal. ---jack.chu | Replace 903-364 one normal. ---jack.chu | Re classified RMED | HW | Process Module | Optical Unit | 131
ABL827 FLEX | Allgemeine Kontrolle ok. Downgrade von 6.16 auf 6.13 Windows und DMS neu installiert ok. Gerät manuell neu konfiguriert ok. System i.o. Kal. + QC + Thb i.o. | Allgemeine Kontrolle ok. Downgrade von 6.16 auf 6.13 Windows und DMS neu installiert ok. Gerät manuell neu konfiguriert ok. System i.o. Kal. + QC + Thb i.o. | SWO-152726 | No | General control ok. Downgrade from 6.16 to 6.13 Windows and DMS reinstalled ok. Device manually reconfigured ok. System i.o. Kal. + QC + Thb i.o. | General control ok. Downgrade from 6.16 to 6.13 Windows and DMS reinstalled ok. Device manually reconfigured ok. System i.o. Kal. + QC + Thb i.o. | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | Cleared used cups from process wheel, it was found that entry on reagent pack for discarded cup receptacle too hard, pushed the cups manually and later it was working normal, run sample measurement and error 1029: Optical sensitivity out of range was reported, cleaned optical unit and process wheel. | Cleared used cups from process wheel, it was found that entry on reagent pack for discarded cup receptacle too hard, pushed the cups manually and later it was working normal, run sample measurement and error 1029: Optical sensitivity out of range was reported, cleaned optical unit and process wheel. | SWO-152772 | No | Cleared used cups from process wheel, it was found that entry on reagent pack for discarded cup receptacle too hard, pushed the cups manually and later it was working normal, run sample measurement and error 1029: Optical sensitivity out of range was reported, cleaned optical unit and process wheel. | Cleared used cups from process wheel, it was found that entry on reagent pack for discarded cup receptacle too hard, pushed the cups manually and later it was working normal, run sample measurement and error 1029: Optical sensitivity out of range was reported, cleaned optical unit and process wheel. | Defect | Consumables | SP | SP
AQT90 FLEX | 更换903-364修复---jack.chu | 更换903-364修复---jack.chu | SWO-151634 | Yes | Replace 903-364 Repair ---jack.chu | Replace 903-364 Repair ---jack.chu | Re classified RMED | HW | Process Module | Optical Unit | 131
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-153065 | No | Replacement test card repair | Replacement test card repair
ABL80 FLEX BASIC | 连续更换了两张测试卡后解决 | 连续更换了两张测试卡后解决 | SWO-152904 | No | Resolved after two consecutive test cards replaced | Resolved after two consecutive test cards replaced
ABL80 FLEX BASIC | 更换测试卡修复 | 作为AX操作建立，无实质性性作用 | SWO-152837 | No | Replacement test card repair | Established as AX operation, withno substantive effect
ABL90 FLEX |  | B | SWO-152833 | No |  |  | Defect | HW | Oxi module | Oxi module | 89
ABL825 FLEX | Calibrations all working fine. Worked with customer on the phone and was unable to resolve. | Showed up to fix problem. Machine had run a qc and passed. Performed 3 more qc's to make sure problem was resolved. Troubleshooting that we did over the phone seemed to resolve all issues. | D1C6F52A060E133D | Yes | Calibrations all working fine. Worked with customer on the phone and was unable to resolve. | Showed up to fix problem. Machine had run a qc and passed. Performed 3 more qc's to make sure problem was resolved. Troubleshooting that we did over the phone seemed to resolve all issues. | Re classified RMED | courtesy
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-153291 | No | Replacement test card repair | Replacement test card repair
ABL90 FLEX | SC und LP getauscht | SC und LP getauscht | SWO-153002 | No | SC and LP exchanged | SC and LP exchanged | Assisted Replacement | SP | SP | SP
ABL90 FLEX | Testprogramme ausgeführt, Fehlersuche, undicht am Einlaß - Nadel sitzt nicht korrekt in Halterung -verbogen - ersetzt, erneut Testprogramme ausgeführt, Startphase abgewartet, QC, ABL90 Box wurde nicht mit ausgeliefert - ein Exemplar hinterlegt, LP als Reserve hinterlassen | B | SWO-153254 | No | Test programs executed, troubleshooting, leaking at the inlet - needle does not sit correctly in bracket - bent - replaced, re-run test programs, start phase waited, QC, ABL90 box was not delivered - a copy deposited, LP left as reserve |  | Defect | HW | Inlet module | Probe bend
ABL90 FLEX | Einlassdichtung erneuert, Gerät neu gestartet, Kaltstartphase abgewartet, Kal und QCs durchgeführt, Sensorkassette erneuert. | Einlassdichtung erneuert, Gerät neu gestartet, Kaltstartphase abgewartet, Kal und QCs durchgeführt, Sensorkassette erneuert. | SWO-153558 | No | Inlet seal renewed, device restarted, cold start phase awaited, Kal and QCs performed, sensor cassette renewed. | Inlet seal renewed, device restarted, cold start phase awaited, Kal and QCs performed, sensor cassette renewed. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D1C87E8B7C3B30C9 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | LIMPIEZA Y CALIBRACION DE Htb. CALIBRADO Y VERIFICADO CON Q.C. | B | D1C8A1457C2C785A | No | CLEANING AND CALIBRATION OF HTb. CALIBRATED AND VERIFIED WITH Q.C. |  | Adjustment | HW | Oxi module | tHb Cal
ABL800 FLEX | Replace the chamber. After replacement calibrate system multiple time. calibration pass qc pass. | Replace the chamber. After replacement calibrate system multiple time. calibration pass qc pass. | SWO-153493 | No | Replace the chamber. After replacement calibrate system multiple time. calibration pass qc pass. | Replace the chamber. After replacement calibrate system multiple time. calibration pass qc pass. | Defect | HW | El/Met Module | Measuring chamber
ABL825 FLEX | Konfigurationseinstellung angepasst und Timeserver konfiguriert | Konfigurationseinstellung angepasst und Timeserver konfiguriert | SWO-153553 | No | Configuration setting adjusted and time server configured | Configuration setting adjusted and time server configured | Setup | SW | DMS_PC | Time setting
ABL835 FLEX |  | B | D1C862257C5EACC5 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX |  | B | D1C85A1D7C3925F8 | No |  |  | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX |  | B | D1C85BAC7C2E25FC | No |  |  | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX | AirVentil-Membran durchstanzt | AirVentil-Membran durchstanzt | SWO-153034 | No | AirValve membrane punched through | AirValve membrane punched through | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX |  | B | D1C849B87C3D40B1 | No |  |  | Revisit required | HW | FlexQ Module | PCB, Auto inlet control
ABL800 BASIC | clean the gasket by flushing it,Run a capillary sample ,got the satisfactory result,handover it the instrument to the customer in good working condition. | clean the gasket by flushing it,Run a capillary sample ,got the satisfactory result,handover it the instrument to the customer in good working condition. | SWO-154416 | No | clean the gasket by flushing it,Run a capillary sample ,got the satisfactory result,handover it the instrument to the customer in good working condition. | clean the gasket by flushing it,Run a capillary sample ,got the satisfactory result,handover it the instrument to the customer in good working condition. | User error | HW | Inlet module | Capillary broken
AQT90 FLEX | Cleaned the process wheel and reference cups performed system cleaning .Still showing the same error. Has to change the optical unit. part number:903-364 | B | SWO-153255 | No | Cleaned the process wheel and reference cups performed system cleaning . Still showing the same error. Has to change the optical unit. part number:903-364 |  | Revisit required | HW | Process Module | Optical Unit
ABL80 FLEX BASIC | 更换试剂包修复 | 更换试剂包修复 | SWO-153692 | No | Replacement reagent kit repair | Replacement reagent kit repair
ABL80 FLEX BASIC | 测试卡漏液，引起测试卡引脚短路，机器出现死机状况，更换1张测试卡和备用件902-972修复。by jesson | 测试卡漏液，引起测试卡引脚短路，机器出现死机状况，更换1张测试卡和备用件902-972修复。by jesson | SWO-153248 | No | Test card leakage, causing the test card pin short circuit, the machine in a dead condition, replace a test card and spare parts 902-972 repair. by jesson | Test card leakage, causing the test card pin short circuit, the machine in a dead condition, replace a test card and spare parts 902-972 repair. by jesson
ABL80 FLEX BASIC | 更换新试剂包后解决 | 更换新试剂包后解决 | SWO-153542 | No | Resolve after replacing a new reagent kit | Resolve after replacing a new reagent kit
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-153880 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | A megjavított műszer visszaszállítása és beüzemelése. Kalibrálás, QC mérés | Készülék az előző napi szerzív után ugyanzat a hibajelenséget produkálta, ezért a helyszinen nem javítható. A helyszinen cserekészülék üzemel a javítás ideéig. Fogyóanyagok cseréje, Kalibrálás, QC mérés | SWO-153388 | Yes | Return and commissioning of the repaired instrument. Calibration, QC measurement | After the previous day's contract, the device produced the same problem, so it cannot be repaired on the site. There is a replacement for the site. Replacement of consumables, calibration, QC measurement | Re classified RMED | Re install analyzer
ABL90 | Customer recently replaced the sensor cassette (SC) and the pCO2 and Cl- parameters failed for several hours. Eventually the Cl- calibration came into range but the pCO2 remains out. | Instructed customer to replace the SC again due to the length of time taken for the cassette to properly condition and calibrate. Customer insisted that a service visit be made. I replaced the SC and allowed it to condition. I also ran blood provided by the lab through the analyzer to assist with conditioning. All calibrations and QC's have passed successfully. All parameters are green. A credit claim has been submitted on the customer's behalf. Analyzer is within acceptable operating specifications. | D1C9BED0F2299FF5 | No | Customer recently replaced the sensor cassette (SC) and the pCO2 and Cl- parameters failed for several hours. Eventually the Cl- calibration came into range but the pCO2 remains out. | Instructed customer to replace the SC again due to the length of time taken for the cassette to properly condition and calibrate. Customer insisted that a service visit be made. I replaced the SC and allowed it to condition. I also ran blood provided by the lab through the analyzer to assist with conditioning. All calibrations and QC's have passed successfully. All parameters are green. A credit claim has been submitted on the customer's behalf. Analyzer is within acceptable operating specifications. | Assisted Replacement | SC | SC | pCO2
ABL90 |  | B | D1C9B54FF2256199 | Yes
ABL90 FLEX | Flush work , back Flush work. Gasket modul clean. Service mode - Pump calibration, cult detector, solution pack check => Fail Solution pack replace => OK | Flush work , back Flush work. Gasket modul clean. Service mode - Pump calibration, cult detector, solution pack check => Fail Solution pack replace => OK | SWO-153220 | No | Flush work , back Flush work. Gasket modul clean. Service mode - Pump calibration, cult detector, solution pack check => Fail Solution pack replace => OK | Flush work , back Flush work. Gasket modul clean. Service mode - Pump calibration, cult detector, solution pack check => Fail Solution pack replace => OK | Assisted Replacement | SP | SP | SP
ABL90 FLEX | Kunde hat bereits 2x das Lösungspack erneuert und mehrmals initialisiert aber ohne Erfolg. Spülung des Gerätes nicht möglich - Gegendruck in den Fließwegen. Die Einlassnadel ist verstopft - Reinigung nicht möglich - erneuert. Lösungspack erneut initialisiert. Sensorkassette benötigt eine längere Zeit um wieder einzulaufen - pCO2 Empfindlichkeit auf 44%. | Kunde hat bereits 2x das Lösungspack erneuert und mehrmals initialisiert aber ohne Erfolg. Spülung des Gerätes nicht möglich - Gegendruck in den Fließwegen. Die Einlassnadel ist verstopft - Reinigung nicht möglich - erneuert. Lösungspack erneut initialisiert. Sensorkassette benötigt eine längere Zeit um wieder einzulaufen - pCO2 Empfindlichkeit auf 44%. | SWO-153523 | No | Customer has already renewed the solution pack 2x and initialized it several times but without success. Flushing of the device is not possible - back pressure in the flow paths. The inlet needle is clogged - cleaning not possible - renewed. solution pack reinitialized. Sensor cartridge takes a longer time to re-enter - pCO2 sensitivity to 44%. | Customer has already renewed the solution pack 2x and initialized it several times but without success. Flushing of the device is not possible - back pressure in the flow paths. The inlet needle is clogged - cleaning not possible - renewed. solution pack reinitialized. Sensor cartridge takes a longer time to re-enter - pCO2 sensitivity to 44%. | Replaced | HW | Inlet module | Probe | Clot
ABL90 FLEX |  | Sie möchte beim nächsten Besuch von einem Techniker einen zusätzlichen Managerkennwort eingerichtet bekommen. 1-6 ist nicht vorhanden. | SWO-153360 | Yes |  | She wants to get an additional manager password set up by a technician the next time she visits. 1-6 does not exist. | Re classified RMED | Telephone support
ABL90 FLEX | Replaced 3 bent contact probes and performed the alignment procedure to make sure they lined up with the sensor cassette correctly. | Replaced 3 bent contact probes and performed the alignment procedure to make sure they lined up with the sensor cassette correctly. | SWO-153565 | No | Replaced 3 bent contact probes and performed the alignment procedure to make sure they lined up with the sensor cassette correctly. | Replaced 3 bent contact probes and performed the alignment procedure to make sure they lined up with the sensor cassette correctly. | Defect | HW | Sensor bay unit | Contact Probe
ABL90 | Removed the connector gasket and twisted and stretch the connector to make it more pliable. After reinserting the gasket, the analyzer passed calibrations and internal QC, but the Level 1 external QC measures 5 to 35 mmHg below the pO2 range. | Performed annual PM. Replaced the pinch valve tubing, connector tubing and inlet gasket. Flushed with hot water. | D1CAAE056836FEF6 | No | Removed the connector gasket and twisted and stretch the connector to make it more pliable. After reinserting the gasket, the analyzer passed calibrations and internal QC, but the Level 1 external QC measures 5 to 35 mmHg below the pO2 range. | Performed annual PM. Replaced the pinch valve tubing, connector tubing and inlet gasket. Flushed with hot water. | Defect | HW | PM Kit | PM Kit
ABL90 | The pO2 keeps failing calibrations and QC measurements due to inhomogeneous errors. Removed the connector gasket and twisted and stretch the connector to make it more pliable. After reinserting the gasket, the analyzer passed calibrations but failed internal QC measurements. Analyzer had message to replace the sensor cassette. | Replaced the pinch valve tubing, connector tubing and inlet gasket. Flushed with hot water. | D1CABFB16826FEFF | No | The pO2 keeps failing calibrations and QC measurements due to inhomogeneous errors. Removed the connector gasket and twisted and stretch the connector to make it more pliable. After reinserting the gasket, the analyzer passed calibrations but failed internal QC measurements. Analyzer had message to replace the sensor cassette. | Replaced the pinch valve tubing, connector tubing and inlet gasket. Flushed with hot water. | Defect | HW | PM Kit | PM Kit
ABL90 | The pO2 keeps failing calibrations and QC measurements due to inhomogeneous errors. Removed the connector gasket and twisted and stretch the connector to make it more pliable. After reinserting the gasket, the analyzer passed calibrations but failed internal QC measurements. Analyzer had message to replace the sensor cassette. | Replaced the pinch valve tubing, connector tubing and inlet gasket. Flushed with hot water. | D1CAB2C06826FEFC | Yes | The pO2 keeps failing calibrations and QC measurements due to inhomogeneous errors. Removed the connector gasket and twisted and stretch the connector to make it more pliable. After reinserting the gasket, the analyzer passed calibrations but failed internal QC measurements. Analyzer had message to replace the sensor cassette. | Replaced the pinch valve tubing, connector tubing and inlet gasket. Flushed with hot water. | Re classified RMED | Double entry
ABL90 |  | B | D1D4F296182F096A | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX | Gerät wurde neu aufgestellt, Lösungspack ersetzt, Einweisungen, Kalibration, QC Kontrolle | Gerät wurde neu aufgestellt, Lösungspack ersetzt, Einweisungen, Kalibration, QC Kontrolle | SWO-153427 | No | Device has been repositioned, solution pack replaced, briefings, calibration, QC control | Device has been repositioned, solution pack replaced, briefings, calibration, QC control | Assisted Replacement | SP | SP | pO2
ABL90 FLEX PLUS |  | B | D1C9C06DF208B1B1 | Yes |  |  | Re classified RMED | Collecting data logs
ABL90 FLEX | 更换新测试卡后修复 | 更换新测试卡后修复 | SWO-153341 | No | Repair after replacing a new test card | Repair after replacing a new test card
ABL825 FLEX | Reviewed datalogs and confirmed problem. | Removed the pCO2 and pH electrodes and checked the sealing ring. Installed 6.16 software. | D1C9CCE9F24B71A7 | Yes | Reviewed datalogs and confirmed problem. | Removed the pCO2 and pH electrodes and checked the sealing ring. Installed 6.16 software. | Re classified RMED | SWU
ABL810 FLEX | Tausch der pO2 Elektrode. Nach mehreren Kalibrationen und Qualitätskontroll Messung sind Werte schlechter und nicht besser. Rückbau. Autocheck Modul öffnet in unregelmäßigen Abständen die falsche Qualitätskontroll Ampullen. Dadurch Fehlmessungen und Mehverbrauch. Qualitätskontroll Level rot verbraucht, neue Lieferung ausgelöst. Autocheck Modul bestellt und pH BG Modul. Im pH BG Modul bei der pO2 Elektrode ist der Elektrodensicht defekt. Mehrere Ersatzteile per Nachtexpress bestellt. Kalibration und Qualitätskontrolle in Ordnung, außer pO2 bei Level rot. | Tausch des AutoCheck Moduls, siehe letzten Servicereport. Erneuerung der Friction Unit im Einlassmodul. LS Justage und Pumpenkal durchgeführt, mehrmals einkalibriert. Leichtes Nässen bei Einlassdichtung reduziert. Qualitätskontroll Ampullen werden richtig geöffnet und angesaugt. Qualitätskontrollen bei pCO2 sind zielwertiger, pO2 Qualitätskontroll Werte werden höher gemessen. Tausch des pH BG Moduls nicht erfolgreich, da Lösung am blauen Schlauch zwischen pH Modul und Einlassmodul entweicht und Qualitätskontrolle abgebrochen wird mit Meldung kein vorderes Luftsegment... Rückbau des pH Moduls, Einbau des eigenen Moduls. Ursache liegt im Einlassmodul. Keines auf dem Fahrzeug. Platine für Einlassmodul bestellt. Weiterhin Probleme mit Probenansaugung. Rückbau Friction Unit, Proben werden wieder richtig angesaugt, Steuerung Einlassmodul nicht in Ordnung, da Lösung bei Einlassdichtung herausfließt. Kalibration und Qualitätskontrolle in Ordnung. | SWO-153582 | No | Replacement of the pO2 electrode. After several calibrations and quality control measurement, values are worse and not better. Dismantling. Autocheck module opens the wrong quality control ampoules at irregular intervals. As a result, incorrect measurements and meh consumption. Quality control level red consumed, new delivery triggered. Autocheck module ordered and pH BG module. In the pH BG module at the pO2 electrode, the electrode vision is defective. Several spare parts ordered by Night Express. Calibration and quality control in order, except pO2 at level red. | Replacement of the AutoCheck module, see last service report. Renewal of the Friction Unit in the inlet module. LS adjustment and pump scale, calibrated several times. Slight wetting when inlet seal is reduced. Quality control ampoules are opened and sucked in correctly. Quality controls at pCO2 are more targeted, pO2 quality control values are measured higher. Replacement of the pH BG module not successful, because solution on the blue hose between pH module and inlet module escapes and quality control is aborted with message no front air segment... Dismantling of the pH module, installation of the own module. The cause is the inlet module. None on the vehicle. board for inlet module. Further problems with sample suction. Dismantling Friction Unit, samples are sucked in properly again, control inlet module is not in order, as solution flows out at inlet seal. Calibration and quality control in order. | Defect | HW | Auto Check Module | Barcode reader | Alignment
ABL825 FLEX | Schaden begutachtet, Schaden durch äußere Gewalteinwirkung verursacht => wird berechnet. Scanner zurzeit einsatzfähig ohne Frontscheibe. Ersatzteil bestellt, Einsatz folgt in Kürze. | FlexQ-, Inlet- und Autocheck-Module ausgebaut, BCR ersetzt. Alle Module installiert, justiert und BCR überprüft, OK. Neustart, Kalibrationen, Funktionstest. >Berechnung wegen äußerer Gewalteinwirkung< | SWO-153429 | No | Damage assessed, damage caused by external force => is calculated. Scanner currently operational without windshield. Spare part ordered, use to follow shortly. | FlexQ, Inlet and Autocheck modules removed, BCR replaced. All modules installed, adjusted and BCR checked, OK. Restart, calibrations, function test. >Calculation due to external force | Revisit required | HW | Basic Unit | Barcode window broken
ABL825 FLEX | FlexQ Modul ausgebaut und zerlegt, SamplerDetectorPCB ersetzt, Montage, Installation und Funktionstest. Endkontrolle: QC Messung | FlexQ Modul ausgebaut und zerlegt; SamplerDetectorPCB ersetzt, Montage, Installation und Funktionstest. Endkontrolle: QC Messung | SWO-153418 | No | FlexQ module removed and disassembled; SamplerDetectorPCB replaces, assembly, installation and function test. Final control: QC measurement | FlexQ module removed and disassembled; SamplerDetectorPCB replaces, assembly, installation and function test. Final control: QC measurement | Defect | HW | FlexQ Module | PCB, FlexQ
ABL825 FLEX | Erneut eine krumme und wieder nach oben abgebogene Einlassnadel ohne äußeren Eingriff! Nadel ersetzt und defekte zur Versendung an RDEU mitgenommen. Verschmutzte Einlassdichtung ebenfalls ersetzt. | Erneut eine krumme und wieder nach oben abgebogene Einlassnadel ohne äußeren Eingriff! Nadel ersetzt und defekte zur Versendung an RDEU mitgenommen. Verschmutzte Einlassdichtung ebenfalls ersetzt. | SWO-153422 | No | Once again a crooked and bent upwards inlet needle without external intervention! Needle replaced and defective ones taken to RDEU for dispatch. Contaminated inlet seal also replaced. | Once again a crooked and bent upwards inlet needle without external intervention! Needle replaced and defective ones taken to RDEU for dispatch. Contaminated inlet seal also replaced. | Defect | HW | Inlet module | Probe bend
ABL835 FLEX | Needs linearity run for IVW | Ran kit will read ever other ampoule correctly the othes have an error of out of reportable range. Ordered spectometer back odered tio MondayDec 19th | D1CC0623DE0EAC48 | No | Needs linearity run for IVW | Ran kit will read ever other ampoule correctly the othes have an error of out of reportable range. Ordered spectometer back odered tio MondayDec 19th | Revisit required | HW | Oxi module | Oxi module
ABL825 FLEX | 0579 errors are present along with A D convertor failed errors | Restart of the analyzer did not reconnect the A D convertor. New Oxi pcb is likely needed. FSE David Sedlock to come on site to complete repair with the necessary part(s) 12-14-2016. | D1CD6AB15459BE61 | No | 0579 errors are present along with A D convertor failed errors | Restart of the analyzer did not reconnect the A D convertor. New Oxi pcb is likely needed. FSE David Sedlock to come on site to complete repair with the necessary part(s) 12-14-2016. | Revisit required | HW | Oxi module | PCB, Oxi
ABL80 FLEX | Sensor Connector Board überprüft und mit Fiberglasbürste gereinigt. ABL startet normal. Im Fehlerdiagnosemenu Temp. überprüft und Nullpunkt eingestellt. Neue SC installiert, alles i.o QA Zyklus und System i.o. Funktionskontrolle i.o. | Sensor Connector Board überprüft und mit Fiberglasbürste gereinigt. ABL startet normal. Im Fehlerdiagnosemenu Temp. überprüft und Nullpunkt eingestellt. Neue SC installiert, alles i.o QA Zyklus und System i.o. Funktionskontrolle i.o. | SWO-153794 | No | Sensor Connector Board checked and cleaned with fiberglass brush. ABL starts normally. In the error diagnostic menu Temp. checked and zero point set. New SC installed, all i.o QA cycle and system i.o. function control i.o. | Sensor Connector Board checked and cleaned with fiberglass brush. ABL starts normally. In the error diagnostic menu Temp. checked and zero point set. New SC installed, all i.o QA cycle and system i.o. function control i.o.
ABL80 FLEX CO-OX | Problematik LP und SC Wechsel mit Personal besprochen (Wechselzyklus, mögliche Problem bei gleichzeitigem Wechsel von LP und SC, Reklamation) Problematik der tHb-Kal und ausgelaufenen LPs besprochen LP Anschlüsse gereinigt, SC gespült Fließwege ausgebaut, überprüft und gespült LS-Justgae, Zero-Kal und Servicetest durchgeführt => alles ohne Beanstandung tHb-Kal i.O. => Fcuv = 1,013 Anschließende Kal und QC Level 1 i.O. => Level 2 weiterhin mit pH und pCO2 Fehler LP zurückgesetzt und reinstalliert Station übern Hr. Löwa Bedienerhandbuch per PDF zur Verfügung stellen lassen Fehlerbehebung (Tabelle Kapitel 10) auf Station hinterlegt | Problematik LP und SC Wechsel mit Personal besprochen (Wechselzyklus, mögliche Problem bei gleichzeitigem Wechsel von LP und SC, Reklamation) Problematik der tHb-Kal und ausgelaufenen LPs besprochen LP Anschlüsse gereinigt, SC gespült Fließwege ausgebaut, überprüft und gespült LS-Justgae, Zero-Kal und Servicetest durchgeführt => alles ohne Beanstandung tHb-Kal i.O. => Fcuv = 1,013 Anschließende Kal und QC Level 1 i.O. => Level 2 weiterhin mit pH und pCO2 Fehler LP zurückgesetzt und reinstalliert Station übern Hr. Löwa Bedienerhandbuch per PDF zur Verfügung stellen lassen Fehlerbehebung (Tabelle Kapitel 10) auf Station hinterlegt | SWO-153750 | No | Problem LP and SC change with staff discussed (change cycle, possible problem with simultaneous change of LP and SC, complaint) The problem of tHb-Kal and expired LPs discussed LP connections cleaned, SC rinsed Flow paths removed, checked and rinsed LS-Justgae, Zero-Kal and service test performed => all without complaint tHb-Kal i.O. => Fcuv = 1,013 Subsequent Kal and QC Level 1 i.O. => Level 2 continues with pH and pCO2 errors LP resets and reinstalls Station via Mr. Löwa to make operator's manual available via PDF Troubleshooting (Table Chapter 10) stored on station | Problem LP and SC change with staff discussed (change cycle, possible problem with simultaneous change of LP and SC, complaint) The problem of tHb-Kal and expired LPs discussed LP connections cleaned, SC rinsed Flow paths removed, checked and rinsed LS-Justgae, Zero-Kal and service test performed => all without complaint tHb-Kal i.O. => Fcuv = 1,013 Subsequent Kal and QC Level 1 i.O. => Level 2 continues with pH and pCO2 errors LP resets and reinstalls Station via Mr. Löwa to make operator's manual available via PDF Troubleshooting (Table Chapter 10) stored on station
ABL90 FLEX PLUS |  | B | D1CADA91680ECFCB | Yes
ABL90 FLEX PLUS |  | B | D1CAD7A26842CFC8 | Yes
ABL90 FLEX PLUS | Laborleitung Frau Bauch berichtet das die Werte des Leihgerätes nun niedriger liegen Bsp. 13,6mg dl zu 14,99mg dl am olympus oder 17,4 zu 18,92 Da das Originalgerät zu hoch gemessen hatte liegt hier scheinbar ein technisches Problem vor das mit uns zu Verfügung stehenden Mitteln nicht weiter verifiziert werden kann. Aus diesem Grund wird das Gerät an TS geschickt und nicht wieder aufgestellt. | Laborleitung Frau Bauch berichtet das die Werte des Leihgerätes nun niedriger liegen Bsp. 13,6mg dl zu 14,99mg dl am olympus oder 17,4 zu 18,92 Da das Originalgerät zu hoch gemessen hatte liegt hier scheinbar ein technisches Problem vor das mit uns zu Verfügung stehenden Mitteln nicht weiter verifiziert werden kann. Aus diesem Grund wird das Gerät an TS geschickt und nicht wieder aufgestellt. | SWO-153780 | No | Laboratory manager Ms Bauch reports that the values of the loaner are now lower e.g. 13.6mg dl to 14.99mg dl at olympus or 17.4 to 18.92 Since the original device had been measured too high, there seems to be a technical problem here before the means at our disposal cannot be further verified. For this reason, the device is sent to TS and not repositioned. | Laboratory manager Ms Bauch reports that the values of the loaner are now lower e.g. 13.6mg dl to 14.99mg dl at olympus or 17.4 to 18.92 Since the original device had been measured too high, there seems to be a technical problem here before the means at our disposal cannot be further verified. For this reason, the device is sent to TS and not repositioned. | Unknown | Performance | Performance | ctBil
ABL90 FLEX PLUS | Replaced inlet probe. Temperature out of range upon arrival. Performed reboot to resolve problem. Calibration failed repeatedly due to PCo2 sensitivity. Sensor cassette replaced. QC B and C failed on number of oximetry values. Carried out complete oximetry test and thb calibration to resolve. | Replaced inlet probe. Temperature out of range upon arrival. Performed reboot to resolve problem. Calibration failed repeatedly due to PCo2 sensitivity. Sensor cassette replaced. QC B and C failed on number of oximetry values. Carried out complete oximetry test and thb calibration to resolve. | SWO-153664 | No | Replaced inlet probe. Temperature out of range upon arrival. Performed reboot to resolve problem. Calibration failed repeatedly due to PCo2 sensitivity. Sensor cassette replaced. QC B and C failed on number of oximetry values. Carried out complete oximetry test and thb calibration to resolve. | Replaced inlet probe. Temperature out of range upon arrival. Performed reboot to resolve problem. Calibration failed repeatedly due to PCo2 sensitivity. Sensor cassette replaced. QC B and C failed on number of oximetry values. Carried out complete oximetry test and thb calibration to resolve. | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX PLUS | Investigated data logs and carried out manual QCs, checked QC setup. All ok. | Investigated data logs and carried out manual QCs, checked QC setup. All ok. | SWO-154580 | Yes | Investigated data logs and carried out manual QCs, checked QC setup. All ok. | Investigated data logs and carried out manual QCs, checked QC setup. All ok. | Re classified RMED | courtesy
ABL90 FLEX PLUS | Flushed analyzer. Reinstalled consumables. | Flushed analyzer. Reinstalled consumables. | SWO-153771 | No | Flushed analyzer. Reinstalled consumables. | Flushed analyzer. Reinstalled consumables. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Limpieza del circuito con agua. Sustitucion de inlet gasket. Chequeo con programa de servicio. Cal, qc y muestras. | B | D1CAE23C683CCCD2 | No | Cleaning the circuit with water. Substitution of inlet gasket. Check with service program. Cal, qc and samples. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Comprobación del estado del analizador. Se comprueban registros, se procede a desmontar módulo de entrada y retirar nipple roto, se sustituye por uno nuevo y se procede a la realización de lavados para comprobar funcionamiento, se realiza ajuste de SL y cal de bombas. Se realiza una Cal. para comprobar funcionamiento y Qc´s. Funcionamiento correcto. | Sustitución de Nipple | D1CAC8646838A085 | No | Checking the status of the scanner. Records are checked, input module is disassembled and nipple removed broken, replaced with a new one and washing is carried out to check operation, SL and pump lime adjustment is performed. A Cal. is performed to check operation and Qc's. Proper operation. | Nipple substitution | Defect | HW | Inlet module | Nipple kit broken
ABL825 FLEX |  | B | D1CC3920DE43F0E9 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL815 FLEX | CPU Unit und Monitor ersetzt, Monitor kalibriert, akz. Kal abgewartet, Kal u. QC i.O. MAC-Adresse: 00-04-5F-90-16-39 Servicedump gezogen | CPU Unit und Monitor ersetzt, Monitor kalibriert, akz. Kal abgewartet, Kal u. QC i.O. MAC-Adresse: 00-04-5F-90-16-39 Servicedump gezogen | SWO-153682 | No | CPU unit and monitor replaced, monitor calibrated, etc. Kal waited, Kal & QC i.O. MAC address: 00-04-5F-90-16-39 Service dump pulled | CPU unit and monitor replaced, monitor calibrated, etc. Kal waited, Kal & QC i.O. MAC address: 00-04-5F-90-16-39 Service dump pulled | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Inletmodul geprüft, Einlassnadel ausgetauscht, Funktionskontrolle ok | Inletmodul geprüft, Einlassnadel ausgetauscht, Funktionskontrolle ok | SWO-153722 | No | Inlet module tested, inlet needle replaced, function control ok | Inlet module tested, inlet needle replaced, function control ok | Defect | HW | Inlet module | Probe
ABL825 FLEX | Fails OXI test in service | Replaced defective OXI module PCB. Complete OXI alignment and tHB in service on all three ABLs | D1CB10D568030293 | No | Fails OXI test in service | Replaced Defective OXYGEN module PCB. Complete OXYGEN Alignment and tHB in service on all three ABLs | Defect | HW | Oxi module | PCB, Oxi
AQT90 FLEX | Blijkt dat systeem uit gaat van een volle, nieuwe blanco cupcartridge. Echter de eerste vier cups zijn reeds verbruikt. Bij de vijfde aanvraag systeem reiniging gaat het goed. | Blijkt dat systeem uit gaat van een volle, nieuwe blanco cupcartridge. Echter de eerste vier cups zijn reeds verbruikt. Bij de vijfde aanvraag systeem reiniging gaat het goed. | SWO-153789 | No | Turns out that system goes out of a full, new blank Cupcartridge. However the first four cups are already consumed. The fifth application system cleaning goes well. | Turns out that system goes out of a full, new blank Cupcartridge. However the first four cups are already consumed. The fifth application system cleaning goes well. | User error | Consumables | Test Kit | Cartridge, blank
ABL80 CO-OX | PO2 was failing qc high on all levels. Customer has replaced both sensor cassette and solution pack with no resolve. | Failed zero check for PO2. Instrument is down. Analog board, sensor conn bd and spline board on order. | D1CFED64403240E5 | No | PO2 was failing qc high on all levels. Customer has replaced both sensor cassette and solution pack with no resolve. | Failed zero check for PO2. Instrument is down. Analog board, sensor conn bd and spline board on order.
ABL80 CO-OX | PO2 was failing qc high on all levels. Customer has replaced both sensor cassette and solution pack with no resolve. | Failed zero check for PO2. Instrument is down. Analog board, sensor conn bd and spline board on order. | D1CFF0DB401940E8 | No | PO2 was failing qc high on all levels. Customer has replaced both sensor cassette and solution pack with no resolve. | Failed zero check for PO2. Instrument is down. Analog board, sensor conn bd and spline board on order.
ABL90 FLEX | Oximetrie heb gecontroleerd, gereinigd en opnieuw heb afgesteld. | Oximetrie heb gecontroleerd, gereinigd en opnieuw heb afgesteld. | SWO-154046 | Yes | Oximetry have been checked, cleaned and re-tuned. | Oximetry have been checked, cleaned and re-tuned.
ABL90 | Customer flushed system multiple times before arriving onsite. No change in status. I tried to flush multiple times with no change. Ran oximetry complete test. Found neon intensity error and spectrometer noise error. | Replaced hemolyzer for troubleshooting purposes. Ran tHb calibration. | D1CC2B6ADE2D3365 | No | Customer flushed system multiple times before arriving onsite. No change in status. I tried to flush multiple times with no change. Ran oximetry complete test. Found neon intensity error and spectrometer noise error. | Replaced hemolyzer for troubleshooting purposes. Ran tHb calibration. | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | Replaced power supply and fitted battery and test ok. Battery P0123N072 | Replaced power supply and fitted battery and test ok. Battery P0123N072 | SWO-154117 | No | Replaced power supply and fitted battery and test ok. Battery P0123N072 | Replaced power supply and fitted battery and test ok. Battery P0123N072 | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | Replaced power supply and fitted battery test ok. Battery P0123N046 | Replaced power supply and fitted battery test ok. Battery P0123N046 | SWO-154118 | No | Replaced power supply and fitted battery test ok. Battery P0123N046 | Replaced power supply and fitted battery test ok. Battery P0123N046 | Defect | HW | Power Supply | Power Supply
ABL90 FLEX PLUS |  | B | D1CCB2F7DE3CC10F | No |  |  | Assisted Replacement | SP | SP | Leak current
ABL90 | REVISIÓN DE HISTÓRICOS. SUSTITUIR AGUJA DE MUESTRAS Y JUNTA SONDA. REALIZAR TEST TÉCNICOS DE FUNCIONAMIENTO. CALIBRACIONES Y CONTROLES | REVISIÓN DE HISTÓRICOS. SUSTITUIR AGUJA DE MUESTRAS Y JUNTA SONDA. REALIZAR TEST TÉCNICOS DE FUNCIONAMIENTO. | D1CC06D8DE4E9078 | No | HISTORICAL REVIEW. SUBSTITUTE SAMPLE NEEDLE AND SONDA JOINT. PERFORM TECHNICAL OPERATION TESTS. CALIBRATIONS AND CONTROLS | HISTORICAL REVIEW. SUBSTITUTE SAMPLE NEEDLE AND SONDA JOINT. PERFORM TECHNICAL OPERATION TESTS. | Replaced | HW | Inlet module | Probe bend
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-154174 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Lac | 378
AQT90 FLEX | 更换903-329后修复 by jesson | 更换903-329后修复 by jesson | SWO-153611 | No | Repair by jesson after replacing 903-329 | Repair by jesson after replacing 903-329 | Defect | HW | Aspiration Module | Needle Wash Unit | 1218
AQT90 FLEX | 更换光单903-364 修复 by jesson | 更换光单903-364 修复 by jesson | SWO-153327 | Yes | Replace light sheet 903-364 repair by jesson | Replace light sheet 903-364 repair by jesson | Re classified RMED | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364一个正常。by jesson | 更换903-364一个正常。by jesson | SWO-153330 | Yes | Replace 903-364 one normal. by jesson | Replace 903-364 one normal. by jesson | Re classified RMED | HW | Process Module | Optical Unit | 131
ABL800 FLEX | Replace the Crea consumable. | Replace the Crea consumable. | SWO-154438 | No | Replace the Crea consumable. | Replace the Crea consumable. | User error | Consumables | Consumables | Crea
ABL825 FLEX |  | B | D1CC2C90DE52D991 | Yes
ABL825 FLEX | Demontage und Wartung Ventile am BG Modul Meßkammer. gereinigt Elektrode membraniert ( pO2, pCO2, Lac , Glu, Ref) | Demontage und Wartung Ventile am BG Modul Meßkammer. gereinigt Elektrode membraniert ( pO2, pCO2, Lac , Glu, Ref) | SWO-153985 | No | Disassembly and maintenance valves on the BG module Measuring chamber. Cleaned Electrode membranes ( pO2, pCO2, Lac , Glu, Ref) | Disassembly and maintenance valves on the BG module Measuring chamber. Cleaned Electrode membranes ( pO2, pCO2, Lac , Glu, Ref) | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX |  | B | D1CC3B2DDE45F0EC | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL810 FLEX | Einlass und Abfall laufen aus. Flüssigkeiten werden nicht korrekt transportiert durch einen Überdruck in den Fließwegen. Abfallstutzen verstopft - Reinigung mit Hypochlorid und Alkohol lösen die Verstopfung nicht, Abflusseinheit zusätzlich in Wasser eingelegt auch hierdurch kein Erfolg. Der ganze Abfluss und die Abdeckung sind durch Blut verklebt. Lösen der Mutter des Abfallstutzen nicht möglich - Plastik bricht. Beide Ersatzteile heute nicht mehr lieferbar. | Ersatz Abdeckung inkl Schlauchstutzen zur Reparaturüberbrückung eingebaut. | SWO-154159 | No | Inlet and waste run out. Liquids are not transported correctly due to an overpressure in the flow paths. Waste nozzles clogged - cleaning with hypochloride and alcohol does not solve the blockage, drain unit additionally inserted in water also this does not succeed. The whole drain and the cover are glued by blood. Loosening of the mother of the waste nozzle not possible - plastic breaks. Both spare parts are no longer available today. | Replacement cover incl hose socket for repair bridging installed. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | FlexQ-, Inlet- und Autocheck-Module ausgebaut, BCR ersetzt. Alle Module installiert, justiert und BCR überprüft, OK. Neustart, Kalibrationen, Funktionstest. >Berechnung wegen äußerer Gewalteinwirkung< | FlexQ-, Inlet- und Autocheck-Module ausgebaut, BCR ersetzt. Alle Module installiert, justiert und BCR überprüft, OK. Neustart, Kalibrationen, Funktionstest. >Berechnung wegen äußerer Gewalteinwirkung< | SWO-154125 | No | FlexQ, Inlet and Autocheck modules removed, BCR replaced. All modules installed, adjusted and BCR checked, OK. Restart, calibrations, function test. >Calculation due to external force | FlexQ, Inlet and Autocheck modules removed, BCR replaced. All modules installed, adjusted and BCR checked, OK. Restart, calibrations, function test. >Calculation due to external force | User error | HW | Basic Unit | Barcode window broken
ABL825 FLEX | Nach Neustart wegen Ersatz des BCRs reagieren alle Tasten bzw. Bildschirmeingaben verzögert, teilweise mit einigen Sekunden Verzögerung. Das Problem wurde etwa 30 Minuten durchgängig beobachtet => Neustart, anschließend Funktionen wieder normal. | Nach Neustart wegen Ersatz des BCRs reagieren alle Tasten bzw. Bildschirmeingaben verzögert, teilweise mit einigen Sekunden Verzögerung. Das Problem wurde etwa 30 Minuten durchgängig beobachtet => Neustart, anschließend Funktionen wieder normal. | SWO-154128 | No | After rebootdue for replacement of the BcR, all buttons or screen input react slated, sometimes with a few seconds delay. The problem was observed continuously for about 30 minutes => Restart, then functions back to normal. | After rebootdue for replacement of the BcR, all buttons or screen input react slated, sometimes with a few seconds delay. The problem was observed continuously for about 30 minutes => Restart, then functions back to normal. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Tägliche Fehlermeldungen Oxi-LS Error => Lufteinlass komplett und hartnäckig verstopft => von blau nach schwarz durchgespült und durchgeblasen. Weitere Aktionen im Rahmen der HJW. | Tägliche Fehlermeldungen Oxi-LS Error => Lufteinlass komplett und hartnäckig verstopft => von blau nach schwarz durchgespült und durchgeblasen. Weitere Aktionen im Rahmen der HJW. | SWO-154076 | No | Daily error messages Oxi-LS Error => Air inlet completely and stubbornly clogged => from blue to black rinsed and blown through. Further actions within the framework of the HJW. | Daily error messages Oxi-LS Error => Air inlet completely and stubbornly clogged => from blue to black rinsed and blown through. Further actions within the framework of the HJW. | Cleaned | Fluidic pathway | Fluidic pathway | Air intake
ABL825 FLEX | Einlassnadel ohne äußere Einwirkung bananenkrumm nach oben verbogen, ersetzt und defekte Nadel zur Versendung an RDEU mitgenommen. Weitere Aktionen im Rahmen der HJW. | Einlassnadel ohne äußere Einwirkung bananenkrumm nach oben verbogen; ersetzt und defekte Nadel zur Versendung an RDEU mitgenommen. Weitere Aktionen im Rahmen der HJW. | SWO-154078 | No | Inlet needle bent upwards without external influence banana crumb; defective needle for dispatch to RDEU. Further actions within the framework of the HJW. | Inlet needle bent upwards without external influence banana crumb; defective needle for dispatch to RDEU. Further actions within the framework of the HJW. | Defect | HW | Inlet module | Probe bend
ABL815 FLEX | Lactat Elektrode geprüft und neu membraniert, Funktionskontrolle ok | Lactat Elektrode geprüft und neu membraniert, Funktionskontrolle ok | SWO-154296 | No | Lactate electrode tested and re-membraneed, function control ok | Lactate electrode tested and re-membraneed, function control ok | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL80 FLEX BASIC | 更换903-334一个正常。 by jesson | 更换903-334一个正常。 by jesson | SWO-153914 | No | Replace 903-334 one normal. by jesson | Replace 903-334 one normal. by jesson
ABL80 CO-OX | PO2 was high on all levels of qc. Zero check failed indicating the PO2 gain was out of spec and analog board needed to be replaced. Will also proactively replace spline pcb and sensor conn bd. | Replaced analog board, sens conn bd and spline pcb. Performed zero check - PASS. Verified instrument operation. | D1CFEBA3405840DC | No | PO2 was high on all levels of qc. Zero check failed indicating the PO2 gain was out of spec and analog board needed to be replaced. Will also proactively replace spline pcb and sensor conn bd. | Replaced analog board, sens conn bd and spline pcb. Performed zero check - PASS. Verified instrument operation.
ABL80 CO-OX | PO2 was high on all levels of qc. Zero check failed indicating the PO2 gain was out of spec and analog board needed to be replaced. Will also proactively replace spline pcb and sensor conn bd. | Replaced analog board, sens conn bd and spline pcb. Performed zero check - PASS. Verified instrument operation. | D1CFF15B405C40EB | No | PO2 was high on all levels of qc. Zero check failed indicating the PO2 gain was out of spec and analog board needed to be replaced. Will also proactively replace spline pcb and sensor conn bd. | Replaced analog board, sens conn bd and spline pcb. Performed zero check - PASS. Verified instrument operation.
ABL80 FLEX | Cleaned SC by cleaning solution fibrine found in SC , Placed SC For cooling for an hour and reinstalled, SC Installation Completed, Problem resolve. | Cleaned SC by cleaning solution fibrine found in SC , Placed SC For cooling for an hour and reinstalled, SC Installation Completed, Problem resolve. | SWO-154441 | No | Cleaned SC by cleaning solution fibrine found in SC , Placed SC For cooling for an hour and reinstalled, SC Installation Completed, Problem resolve. | Cleaned SC by cleaning solution fibrine found in SC , Placed SC For cooling for an hour and reinstalled, SC Installation Completed, Problem resolve.
ABL90 FLEX | Defecte batterij heb vervangen. Lot nr. Defecte accu: P0061N007 Lot nr. Nieuwe accu: P0131N012 | Defecte batterij heb vervangen. Lot nr. Defecte accu: P0061N007 Lot nr. Nieuwe accu: P0131N012 | SWO-154388 | No | Defective battery. Lot Nr. Defective Battery: P0061N007 Lot Nr. New Battery: P0131N012 | Defective battery. Lot Nr. Defective Battery: P0061N007 Lot Nr. New Battery: P0131N012 | Defect | HW | Basic Unit | Battery
ABL90 | Initially spoke to Ted Fuertes who stated shortly after (several days) installing a new solution pack, the analyzer is reporting change solution pack . Claims not to have been trained in troubleshooting this error and did not have the flushing device. I was referred to the RT on duty, Norbit. | Norbit had more complete information. He arrived for his shift with the analyzer in troubleshooting mode. I explained what actions would help, he was not able to locate the flushing device but was able to fabricate an acceptable device. I provided instructions on the process, he completed the steps, reported the actions appeared to resolve the issue. I asked he call back to report the status, he never called back. | D1CF422540021889 | No | Initially spoke to Ted Fuertes who stated shortly after (several days) installing a new solution pack, the analyzer is reporting change solution pack . Claims not to have been trained in troubleshooting this error and did not have the flushing device. I was referred to the RT on duty, Norbit. | Norbit had more complete information. He arrived for his shift with the analyzer in troubleshooting mode. I explained what actions would help, he was not able to locate the flushing device but was able to fabricate an acceptable device. I provided instructions on the process, he completed the steps, reported the actions appeared to resolve the issue. I asked he call back to report the status, he never called back.
ABL90 FLEX | Voeding en batterij unit heb vervangen. | Voeding en batterij unit heb vervangen. | SWO-154393 | No | Have replaced the power supply and battery unit. | Have replaced the power supply and battery unit. | Defect | HW | Basic Unit | Battery
ABL90 FLEX | Voeding en batterij units heb vervangen. | Voeding en batterij unit heb vervangen. | SWO-154680 | No | Power supply and battery units have been replaced. | Have replaced the power supply and battery unit. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | Datalogs & Service Dump erstellt und an TS geschickt, CCF erhoben | Datalogs & Service Dump erstellt und an TS geschickt, CCF erhoben | SWO-154287 | No | Datalogs & Service Dump created and sent to TS, CCF collected | Datalogs & Service Dump created and sent to TS, CCF collected | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS |  | B | D1CD7D70540F44FA | No |  |  | Adjustment | HW | Auto Inlet module | Inlet drive gear
AQT90 FLEX | 更换903-329备件，修复 by jesson | 更换903-329备件，修复 by jesson | SWO-153890 | No | Replace 903-329 spare parts and repair by jesson | Replace 903-329 spare parts and repair by jesson | Defect | HW | Aspiration Module | Needle Wash Unit | 1259
AQT90 FLEX | 更换580-186和903-036后修复 byjesson | 更换580-186和903-036后修复 byjesson | SWO-153593 | No | Repair byjesson after replacing 580-186 and 903-036 | Repair byjesson after replacing 580-186 and 903-036 | Defect | HW | Cover and Chassis | PCB, replenisment PCB
AQT90 FLEX | 自行更换了636-550后修复 by jesson | 自行更换了636-550后修复 by jesson | SWO-154186 | No | Repair by jesson after replacing 636-550 | Repair by jesson after replacing 636-550 | Defect | HW | Process Module | Process Unit sensor | 5784
ABL810 FLEX | Ersatz Abdeckung inkl Schlauchstutzen zur Reparaturüberbrückung eingebaut. | Ersatz Abdeckung inkl Schlauchstutzen zur Reparaturüberbrückung eingebaut. | SWO-154465 | No | Replacement cover incl hose socket for repair bridging installed. | Replacement cover incl hose socket for repair bridging installed. | Defect | HW | Basic Unit | Waste connector
ABL810 FLEX | Tausch des AutoCheck Moduls, siehe letzten Servicereport. Erneuerung der Friction Unit im Einlassmodul. LS Justage und Pumpenkal durchgeführt, mehrmals einkalibriert. Leichtes Nässen bei Einlassdichtung reduziert. Qualitätskontroll Ampullen werden richtig geöffnet und angesaugt. Qualitätskontrollen bei pCO2 sind zielwertiger, pO2 Qualitätskontroll Werte werden höher gemessen. Tausch des pH BG Moduls nicht erfolgreich, da Lösung am blauen Schlauch zwischen pH Modul und Einlassmodul entweicht und Qualitätskontrolle abgebrochen wird mit Meldung kein vorderes Luftsegment... Rückbau des pH Moduls, Einbau des eigenen Moduls. Ursache liegt im Einlassmodul. Keines auf dem Fahrzeug. Platine für Einlassmodul bestellt. Weiterhin Probleme mit Probenansaugung. Rückbau Friction Unit, Proben werden wieder richtig angesaugt, Steuerung Einlassmodul nicht in Ordnung, da Lösung bei Einlassdichtung herausfließt. Kalibration und Qualitätskontrolle in Ordnung. | Tausch des AutoCheck Moduls, siehe letzten Servicereport. Erneuerung der Friction Unit im Einlassmodul. LS Justage und Pumpenkal durchgeführt, mehrmals einkalibriert. Leichtes Nässen bei Einlassdichtung reduziert. Qualitätskontroll Ampullen werden richtig geöffnet und angesaugt. Qualitätskontrollen bei pCO2 sind zielwertiger, pO2 Qualitätskontroll Werte werden höher gemessen. Tausch des pH BG Moduls nicht erfolgreich, da Lösung am blauen Schlauch zwischen pH Modul und Einlassmodul entweicht und Qualitätskontrolle abgebrochen wird mit Meldung kein vorderes Luftsegment... Rückbau des pH Moduls, Einbau des eigenen Moduls. Ursache liegt im Einlassmodul. Keines auf dem Fahrzeug. Platine für Einlassmodul bestellt. Weiterhin Probleme mit Probenansaugung. Rückbau Friction Unit, Proben werden wieder richtig angesaugt, Steuerung Einlassmodul nicht in Ordnung, da Lösung bei Einlassdichtung herausfließt. Kalibration und Qualitätskontrolle in Ordnung. | SWO-154386 | No | Replacement of the AutoCheck module, see last service report. Renewal of the Friction Unit in the inlet module. LS adjustment and pump scale, calibrated several times. Slight wetting when inlet seal is reduced. Quality control ampoules are opened and sucked in correctly. Quality controls at pCO2 are more targeted, pO2 quality control values are measured higher. Replacement of the pH BG module not successful, because solution on the blue hose between pH module and inlet module escapes and quality control is aborted with message no front air segment... Dismantling of the pH module, installation of the own module. The cause is the inlet module. None on the vehicle. board for inlet module. Further problems with sample suction. Dismantling Friction Unit, samples are sucked in properly again, control inlet module is not in order, as solution flows out at inlet seal. Calibration and quality control in order. | Replacement of the AutoCheck module, see last service report. Renewal of the Friction Unit in the inlet module. LS adjustment and pump scale, calibrated several times. Slight wetting when inlet seal is reduced. Quality control ampoules are opened and sucked in correctly. Quality controls at pCO2 are more targeted, pO2 quality control values are measured higher. Replacement of the pH BG module not successful, because solution on the blue hose between pH module and inlet module escapes and quality control is aborted with message no front air segment... Dismantling of the pH module, installation of the own module. The cause is the inlet module. None on the vehicle. board for inlet module. Further problems with sample suction. Dismantling Friction Unit, samples are sucked in properly again, control inlet module is not in order, as solution flows out at inlet seal. Calibration and quality control in order. | Revisit required | HW | Inlet module | Inlet module | Leaking
ABL810 FLEX | Allgemeine Kontrolle ok. Drucker ausgebaut und überprüft ok. Verbindungskabel nochmals neu eingesteckt i.o. Drucker i.o. System i.o. | Allgemeine Kontrolle ok. Drucker ausgebaut und überprüft ok. Verbindungskabel nochmals neu eingesteckt i.o. Drucker i.o. System i.o. | SWO-154253 | No | General control ok. Printer removed and checked ok. Connection cable re-plugged in i.o. Printer i.o. System i.o. | General control ok. Printer removed and checked ok. Connection cable re-plugged in i.o. Printer i.o. System i.o. | Adjustment | HW | Basic Unit | Cable, Printer
ABL825 FLEX |  | B | D1CD7AE55416628A | No |  |  | Defect | HW | FlexQ Module | FlexQ Module
ABL835 FLEX | Oximetrie Modul geprüft, Lampeneinheit (902-458r2670N005) ausgetauscht und justiert, tHb Kalibration, Lautstärke am Monitor eingestellt, Funktionskontrolle ok | Oximetrie Modul geprüft, Lampeneinheit (902-458r2670N005) ausgetauscht und justiert, tHb Kalibration, Lautstärke am Monitor eingestellt, Funktionskontrolle ok | SWO-154294 | No | Oximetry module tested, lamp unit (902-458r2670N005) replaced and adjusted, tHb calibration, volume set on monitor, function control ok | Oximetry module tested, lamp unit (902-458r2670N005) replaced and adjusted, tHb calibration, volume set on monitor, function control ok | Defect | HW | Oxi module | Lamp unit
ABL835 FLEX | No halogen lamp voltage | Ordered new OXI module will replace on Monday Dec 19th | D1D01A24405C7C2C | No | No halogen lamp voltage | Ordered new OXI module will replace on Monday Dec 19th | Revisit required | HW | Oxi module | Oxi module
AQT90 FLEX | changed the slider and the belts on the motor, used from other analyser on the testing purpose. | changed the slider and the belts on the motor, used from other analyser on the testing purpose. | SWO-154421 | No | changed the slider and the belts on the motor, used from other analyser on the testing purpose. | changed the slider and the belts on the motor, used from other analyser on the testing purpose. | Defect | HW | Cover and Chassis | Inlet Door
AQT90 FLEX | Frau Frau Führbach am 15.12.2016 empfohlen das Reagenzienpack zu erneuern. Seitdem ist der Fehler nicht mehr aufgetreten. Kunden am 16.12.2016 besucht. Service Dump und Datensicherung durchgeführt. Motoren neu gefettet. Reagenzien Pack mitgenommen um es zur Zentrale zu senden. | Frau Frau Führbach am 15.12.2016 empfohlen das Reagenzienpack zu erneuern. Seitdem ist der Fehler nicht mehr aufgetreten. Kunden am 16.12.2016 besucht. Service Dump und Datensicherung durchgeführt. Motoren neu gefettet. Reagenzien Pack mitgenommen um es zur Zentrale zu senden. | SWO-154466 | Yes | Ms. Führbach on 15.12.2016 recommended to renew the reagent pack. The error has not occurred since. customers on 16.12.2016. Service dump and data backup performed. motors regreased. Reagenzien Pack to send it to the head office. | Ms. Führbach on 15.12.2016 recommended to renew the reagent pack. The error has not occurred since. customers on 16.12.2016. Service dump and data backup performed. motors regreased. Reagenzien Pack to send it to the head office. | Re classified RMED | Collecting data
ABL800 BASIC |  | B | SWO-154462 | No |  |  | Defect | HW | Oxi module | Oxi module
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-154880 | No | Replacement test card repair | Replacement test card repair
ABL90 FLEX | Service mode check - clot, solution pack => Fail Liquid flow line flush work. Service mode check => OK Calibration => Fail (Glu, Lac) Sensor cassette replace Calibration, Control => OK | Service mode check - clot, solution pack => Fail Liquid flow line flush work. Service mode check => OK Calibration => Fail (Glu, Lac) Sensor cassette replace Calibration, Control => OK | SWO-154479 | No | Service mode check - clot, solution pack => Fail Liquid flow line flush work. Service mode check => OK Calibration => Fail (Glu, Lac) Sensor cassette replace Calibration, Control => OK | Service mode check - clot, solution pack => Fail Liquid flow line flush work. Service mode check => OK Calibration => Fail (Glu, Lac) Sensor cassette replace Calibration, Control => OK | Assisted Replacement | SC | SC | Glu/Lac
ABL80 | Se desmonto el equipo y se coloco correctamente el cable del display. finalmente se comprobó correcto funcionamiento. | B | D1DB6BF6661E1F77 | No | The equipment is unmounted and the display cable is properly placed. it was finally proven to be functioning properly.
ABL80 OSM | CONFIRMED ISSUE WHEN ONSITE. PERFORMED SYSTEM CYCLE, SERVICE TEST AND CALIBRATED LIQUID SENSOR VIA DIAGNOSTICS. ALL PASSED. | 1. Unload the solution pack from the analyzer (take it out without unloading via software) 2. Use a syringe and from the solution pack rear ports, aspirate liquids (about 3-4ml) from each ports. Do not aspirate from the lowest port as this is waste port. 3. Reload the solution pack 4. Shutdown and restart the analyzer to recalibrate the liquid sensor and prime all liquids in the pack | D1D17F78B652C4F3 | No | CONFIRMED ISSUE WHEN ONSITE. PERFORMED SYSTEM CYCLE, SERVICE TEST AND CALIBRATED LIQUID SENSOR VIA DIAGNOSTICS. ALL PASSED. | 1. Unload the solution pack from the analyzer (take it out without unloading via software) 2. Use a syringe and from the solution pack rear ports, aspirate liquids (about 3-4ml) from each ports. Do not aspirate from the lowest port as this is waste port. 3. Reload the solution pack 4. Shutdown and restart the analyzer to recalibrate the liquid sensor and prime all liquids in the pack
ABL80 FLEX BASIC |  | B | SWO-154537 | No
ABL80 FLEX Basic version |  | B | D1D1540CB64E9781 | No
ABL90 FLEX | folyadék áramlási tesztek: OK, a 17-én telepített oldat chip-je nem kommunikál a készülékkel, szerviz tesztben sem. Analyzer ID PCB és a csatlakozás ellenőrzése, oldat újra telepítése után: 1405 - összeférhetetlen adatbázis hiba üzenet. Újra indítás után a hiba jelzés megszűnik, kalibrációk, QC mérés pO2 QC hiba, SC csere, kalibrációk, QC mérés ********************************************* pO2 továbbra is rossz, oldat csere szükséges!! ********************************************* | folyadék áramlási tesztek: OK, a 17-én telepített oldat chip-je nem kommunikál a készülékkel, szerviz tesztben sem. Analyzer ID PCB és a csatlakozás ellenőrzése, oldat újra telepítése után: 1405 - összeférhetetlen adatbázis hiba üzenet. Újra indítás után a hiba jelzés megszűnik, kalibrációk, QC mérés pO2 QC hiba, SC csere, kalibrációk, QC mérés ********************************************* pO2 továbbra is rossz, oldat csere szükséges!! ********************************************* | SWO-154683 | No | Liquid flow tests: OK, The chip of the solution installed on the 17 does not communicate with the device, nor in a service test. Analyzer ID PCB and connection check, solution re-installed after: 1405-incompatible database error message. After re-booting, the fault indicator disappears, calibrations, QC measurement pO2 QC Error, SC replacement, calibrations, QC measurement ********************************************* pO2 remains bad, solution replacement is required!! ********************************************* | Liquid flow tests: OK, The chip of the solution installed on the 17 does not communicate with the device, nor in a service test. Analyzer ID PCB and connection check, solution re-installed after: 1405-incompatible database error message. After re-booting, the fault indicator disappears, calibrations, QC measurement pO2 QC Error, SC replacement, calibrations, QC measurement ********************************************* pO2 remains bad, solution replacement is required!! ********************************************* | Assisted Replacement | SP | SP | FT
ABL90 | analyzer was down due to a cracked inlet gasket holder ABL 90 kept requesting a flush and to call service tech. sensor is expired will need to be replaced. prof test was ran on 11 21 2016 leading up to the test 24 hours prior all QC and cals were in. There were 2 cals out for po2 on the 19th and 2 QCs for po2 and K+ that were out. when looking at the failed samples they were run on the wrong syringe mode. | replaced cracked inlet gasket holder and replaced expired sensor. Customer is ordering a new kit and will run on the correct syringe mode patient sample mode not prof. test mode. | D1D11F65B6153D3A | No | analyzer was down due to a cracked inlet gasket holder ABL 90 kept requesting a flush and to call service tech. sensor is expired will need to be replaced. prof test was ran on 11 21 2016 leading up to the test 24 hours prior all QC and cals were in. There were 2 cals out for po2 on the 19th and 2 QCs for po2 and K+ that were out. when looking at the failed samples they were run on the wrong syringe mode. | replaced cracked inlet gasket holder and replaced expired sensor. Customer is ordering a new kit and will run on the correct syringe mode patient sample mode not prof. test mode. | Defect | HW | Inlet module | Inlet Gasket
ABL90 | unit was all green in ready mode no cal errors or qc errors present. upon further investigation of the failed samples it looks to be that they were ran in the wrong syringe mode | Customer is re ordering the test and will run it in the correct syringe mode patient sample mode not prof test mode | D1D12B69B61E3D3F | No | unit was all green in ready mode no cal errors or qc errors present. upon further investigation of the failed samples it looks to be that they were ran in the wrong syringe mode | Customer is re ordering the test and will run it in the correct syringe mode patient sample mode not prof test mode | User error | Sample Handling | Sample Handling | Sample Handling
ABL90 FLEX | New printer fitted and now the analyser is printing correctly. | New printer fitted and now the analyser is printing correctly. | SWO-154752 | No | New printer fitted and now the analyser is printing correctly. | New printer fitted and now the analyser is printing correctly. | Defect | HW | Printer module | Printer module
ABL90 FLEX PLUS | Gerät geprüft. Fliesswege mehrmals gespült , Betriebsmittel erneut installiert.- Keine Änderung Einlasshalter , Nadel und Dichtung geprüft getauscht. Diverse Diagnoseprogramme im Service aufgerufen. Analysator demontiert und neues Schlauchkit eingebaut und neu gestartet. Alles ohne Erfolg. Analysator erneut demontiert. Hemolyzer ausgebaut und mit warmen Wasser diverse Male gespült. Gerät wieder zusammengebaut,Betriebsmittel wieder installiert. Neustart und Startphase abgewartet. | Gerät geprüft. Fliesswege mehrmals gespült , Betriebsmittel erneut installiert.- Keine Änderung Einlasshalter , Nadel und Dichtung geprüft getauscht. Diverse Diagnoseprogramme im Service aufgerufen. Analysator demontiert und neues Schlauchkit eingebaut und neu gestartet. Alles ohne Erfolg. Analysator erneut demontiert. Hemolyzer ausgebaut und mit warmen Wasser diverse Male gespült. Gerät wieder zusammengebaut,Betriebsmittel wieder installiert. Neustart und Startphase abgewartet. | SWO-155130 | No | device checked. Flow paths rinsed several times, equipment reinstalled.- No change Inlet holder, needle and seal checked replaced. Various diagnostic programs are called in the service. Analyzer disassembled and new hose kit installed and restarted. All without success. Analyzer dismantled again. Hemolyzer removed and rinsed with warm water several times. Device reassembled,equipment reinstalled. restart and start-up phase awaited. | device checked. Flow paths rinsed several times, equipment reinstalled.- No change Inlet holder, needle and seal checked replaced. Various diagnostic programs are called in the service. Analyzer disassembled and new hose kit installed and restarted. All without success. Analyzer dismantled again. Hemolyzer removed and rinsed with warm water several times. Device reassembled,equipment reinstalled. restart and start-up phase awaited. | Cleaned | Fluidic pathway | Oxi module | Hemolyzer
ABL90 FLEX PLUS |  | B | D1D15975B6049935 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS |  | B | D1D1529FB6539931 | No |  |  | Assisted Replacement | SC | SC | Chip code error
AQT90 FLEX | 连续更换2个903-364定标不过。又更换一个903-364修复(by zhenjie) | 之前更换2个备用件903-364定标不过，后面又追加903-364一个(by zhenjie) | SWO-154542 | Yes | Replace 2 903-364 calibrations in a row. Replace another 903-364 repair (by-zhenjie) | Replace 2 spareparts 903-364 prescription but with 903-364 (by zhenjie) | Re classified RMED | HW | Process Module | Optical Unit | 131
ABL800 FLEX | Clean the chamber. Run protein removal multiple time & Clean inlet gasket. replace the membrane. | Clean the chamber. Run protein removal multiple time & Clean inlet gasket. replace the membrane. | SWO-154772 | No | Clean the chamber. Run protein removal multiple time & Clean inlet gasket. replace the membrane. | Clean the chamber. Run protein removal multiple time & Clean inlet gasket. replace the membrane. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX |  | B | D1D10F87B60423D9 | No |  |  | Defect | HW | Basic Unit | Barcode reader, Basic Unit
ABL810 FLEX | Barometer-Einstellungen berichtigt und korrigiert - Funktionsüberprüfung etc. ohne Beanstandung ! | Barometer-Einstellungen berichtigt und korrigiert - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-154736 | No | Barometer settings corrected and corrected - function check etc. without complaint ! | Barometer settings corrected and corrected - function check etc. without complaint ! | Setup | SW | Barometric pressure | Barometer
ABL810 FLEX | Analysator heruntergefahren - Erneuerung Inlet PCB (Inkonsistente Software) - Software Pflege - Funktionsüberprüfung etc. brachte keine Änderung der QC Ergebnisse - Rückmontage Inlet PCB - Barometerwerte Überprüfung Höhenwerte von 750 m auf 350 m berichtigt und Korrektur durchgeführt - Funktionsüberprüfung etc. ohne Beanstandung ! | Analysator heruntergefahren - Erneuerung Inlet PCB (Inkonsistente Software) - Software Pflege - Funktionsüberprüfung etc. brachte keine Änderung der QC Ergebnisse - Rückmontage Inlet PCB - Barometerwerte Überprüfung Höhenwerte von 750 m auf 350 m berichtigt und Korrektur durchgeführt - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-154737 | No | Analyzer shut down - Renewal Inlet PCB (Inconsistent Software) - Software Maintenance - Functional Review etc. brought no change in QC results - Reassembly Inlet PCB - Barometer values check altitude values from 750 m to 350 m corrected and correction performed - function check etc. without complaint ! | Analyzer shut down - Renewal Inlet PCB (Inconsistent Software) - Software Maintenance - Functional Review etc. brought no change in QC results - Reassembly Inlet PCB - Barometer values check altitude values from 750 m to 350 m corrected and correction performed - function check etc. without complaint ! | Setup | SW | Analyzer Control Module | Barometer | pO2 biased
ABL810 FLEX | Barometerwerte berichtigt und korrigiert - Funktionsüberprüfung etc. ohne Beanstandung ! Siehe SWO von Labor-Analysator | Barometerwerte berichtigt und korrigiert - Funktionsüberprüfung etc. ohne Beanstandung ! Siehe SWO von Labor-Analysator | SWO-154735 | No | Barometer values corrected and corrected - function check etc. without complaint ! See SWO by Laboratory Analyzer | Barometer values corrected and corrected - function check etc. without complaint ! See SWO by Laboratory Analyzer | Setup | SW | Barometric pressure | Barometer
ABL820 FLEX | PCO2 was failing peer group 2 on all 3 levels of QC. Customer has had to remembrane the PCO2 electrode every couple of days. Instrument was installed in May 2016, so the focus will be on the measuring chamber and electrodes. | Proactively replaced the sealing tip ring for pH and replaced the dummy electrode for Chloride. Replaced PCO2 electrode under warranty. Remembraned PCO2 electrode using membrane lot 815. Customer was using lot 788. Verified instrument operation. | D1D134F6B6451C00 | No | PCO2 was failing peer group 2 on all 3 levels of QC. Customer has had to remembrane the PCO2 electrode every couple of days. Instrument was installed in May 2016, so the focus will be on the measuring chamber and electrodes. | Proactively replaced the sealing tip ring for pH and replaced the dummy electrode for Chloride. Replaced PCO2 electrode under warranty. Remembraned PCO2 electrode using membrane lot 815. Customer was using lot 788. Verified instrument operation. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Feder neu eingebaut, Gerät kalibriett, | Feder neu eingebaut, Gerät kalibriett, | SWO-154728 | No | Spring reinstalled, device calibriett, | Spring reinstalled, device calibriett, | Defect | HW | pH/BG Module | Electrode connector
ABL800 BASIC | Replaced New PO2 Electrode | B | SWO-155927 | No | Replaced New PO2 Electrode |  | Assisted Replacement | Electrode | Electrode | pO2
ABL825 FLEX | CONFIRMED ISSUE WHEN ONSITE. PERFORMED OXI COMPLETE TEST AND FAILED LAMP MODULE DUE TO NEON INTENSITY LOW. | PROCEED TO REPLACE OXY MODULE PCB and NEON LAMP. | D1D15332B627B61B | No | CONFIRMED ISSUE WHEN ONSITE. PERFORMED OXI COMPLETE TEST AND FAILED LAMP MODULE DUE TO NEON INTENSITY LOW. | PROCEED TO REPLACE OXY MODULE PCB and NEON LAMP. | Defect | HW | Oxi module | Lamp unit
ABL825 FLEX | FOUND WASTE LINE AND PROBE BLOCKED WITH BLOOD | REPLACED WASTE LINE AND PROBE | D1D1A2C9B612479A | No | FOUND WASTE LINE AND PROBE BLOCKED WITH BLOOD | REPLACED WASTE LINE AND PROBE | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Clot
ABL835 FLEX | Fails OXI test in service | put in new OXI module getting Neon position. Replaced lamp unit, did not help. Changed spectrophometer. Now passes OXI test in service | D1D272482C06E31B | Yes | Fails OXI test in service | Put in new OXY modules getting Neon position. Replaced lamp Unit, did not help. Changed Spectrophometer. Now passes OXI test in service | Re classified RMED | OOB
ABL827 FLEX | It was aborting calibrations because the cal 1 bottle was empty and cal 2 was almost empty. When new bottles were put on it stopped aborting the calibrations and the Crea started to pass calibrations straight away. It looks like there was something wrong with the additive for the cal solutions. | It was aborting calibrations because the cal 1 bottle was empty and cal 2 was almost empty. When new bottles were put on it stopped aborting the calibrations and the Crea started to pass calibrations straight away. It looks like there was something wrong with the additive for the cal solutions. | SWO-154782 | No | It was aborting calibrations because the cal 1 bottle was empty and cal 2 was almost empty. When new bottles were put on it stopped aborting the calibrations and the Crea started to pass calibrations straight away. It looks like there was something wrong with the additive for the cal solutions. | It was aborting calibrations because the cal 1 bottle was empty and cal 2 was almost empty. When new bottles were put on it stopped aborting the calibrations and the Crea started to pass calibrations straight away. It looks like there was something wrong with the additive for the cal solutions. | Assisted Replacement | Consumables | Solution bottle | Crea
AQT90 FLEX |  | B | D1D16379B6332CED | No |  |  | Defect | SW | Analyzer Control Module | Off/On | 1999
ABL80 FLEX CO-OX | Analysator gereinigt, Manifold überprüft, Funktionskontrolle | Analysator gereinigt, Manifold überprüft, Funktionskontrolle | SWO-154826 | No | Analyzer cleaned, Manifold checked, function control | Analyzer cleaned, Manifold checked, function control
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-154865 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | LP Wechsel | LP Wechsel | SWO-154923 | No | LP Change | LP Change | User error | SP Empty | SP Empty | SP Empty
ABL90 | flushed device, there was some restriction after 4 flushes unit was flowing as it should, re ran cals and QC all passed external QC failing PO2 and Ca low 140 range 144-164 p02 ca low 1.4 range 1.78-2.18 re ran QC still failing re calibrated and re ran internal QC external still failing. did a visual fluid check with a clear tube in place of the sensor, fluid is flowing with no air bubbles - changed sensor cassette | replacement of sensor cassette | D1D2B6892C16D7D0 | No | flushed device, there was some restriction after 4 flushes unit was flowing as it should, re ran cals and QC all passed external QC failing PO2 and Ca low 140 range 144-164 p02 ca low 1.4 range 1.78-2.18 re ran QC still failing re calibrated and re ran internal QC external s till failing. did a visual fluid check with a clear tube in place of the sensor, fluid is flowing with no air bubbles - changed sensor cassette | replacement of sensor cassette | Assisted Replacement | SC | SC | FT
ABL90 | The inlet and connector gaskets, pinch valve tubing and waste tubing were replaced the other day. | Replaced the pinch valve and cleaned the pO2 sensor. | D1D286A52C62E56D | No | The inlet and connector gaskets, pinch valve tubing and waste tubing were replaced the other day. | Replaced the pinch valve and cleaned the pO2 sensor. | Defect | HW | Oxi module | Pinch valve
ABL90 | The inlet and connector gaskets, pinch valve tubing, and waste tubing have already been replaced. | Replaced the pinch valve because it was sticking intermittently and cleaned the pO2 sensor. | D1D27CFF2C56E566 | No | The inlet and connector gaskets, pinch valve tubing, and waste tubing have already been replaced. | Replaced the pinch valve because it was sticking intermittently and cleaned the pO2 sensor. | Defect | HW | Oxi module | Pinch valve
ABL90 FLEX | Gerinnsel durch Spülung der Einlassnadel und Einlassdichtung entfernt --> jedoch ohne Erfolg, Einlassnadel und Einlassdichtung erneuert, Detect remove im Service --> i.O., Systemcheck, System und Funktion i.O. | Gerinnsel durch Spülung der Einlassnadel und Einlassdichtung entfernt --> jedoch ohne Erfolg, Einlassnadel und Einlassdichtung erneuert, Detect remove im Service --> i.O., Systemcheck, System und Funktion i.O. | SWO-155090 | No | Removes clots by flushing the inlet needle and inlet seal --> but without success, inlet needle and inlet seal renewed, Detect remove in service --> i.O., system check, system and function i.O. | Removes clots by flushing the inlet needle and inlet seal --> but without success, inlet needle and inlet seal renewed, Detect remove in service --> i.O., system check, system and function i.O. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | 之前更换2个备用件903-364定标不过，后面又追加903-364一个(by zhenjie) | 之前更换2个备用件903-364定标不过，后面又追加903-364一个(by zhenjie) | SWO-154548 | Yes | Replace 2 spareparts 903-364 prescription but with 903-364 (by zhenjie) | Replace 2 spareparts 903-364 prescription but with 903-364 (by zhenjie) | Re classified RMED | HW | Process Module | Optical Unit | 131
ABL825 FLEX | Tausch Ventile BG - EL Modul Austausch der Gasventile Überprüfung Sensor Gasunit | Tausch Ventile BG - EL Modul Austausch der Gasventile Überprüfung Sensor Gasunit | SWO-154983 | No | Replacement valves BG - EL module Replacement of gas valves Verification Sensor Gasunit | Replacement valves BG - EL module Replacement of gas valves Verification Sensor Gasunit | Defect | HW | Gas module | Gas valve 1 and 2
ABL825 FLEX | Referenzelektrode neu membraniert und Messkammer gereinigt, System und Funktion i.O. | Referenzelektrode neu membraniert und Messkammer gereinigt, System und Funktion i.O. | SWO-154973 | No | Reference electrode re-membraneed and measuring chamber cleaned, system and function in the | Reference electrode re-membraneed and measuring chamber cleaned, system and function in the | Assisted Replacement | Electrode | Membrane | Reference
ABL810 FLEX | El Met Modul eingebaut und mit den entsprechenden Einstellungen hinzugefügt. Startphase, Kalibrierung und QC's durchgeführt. | B | SWO-155049 | No | El Met module built-in and added with the appropriate settings. start-up phase, calibration and QC's. |  | Assisted Replacement | Electrode | Electrode | pCO2
ABL815 FLEX | Barcode einwandfrei lesbar aber Monitor reagiert nicht, mit Verzögerung und nur sporadisch auf Berührungen, Videos der Fehlfunktion angefertigt. Testweise den - kaum verschmutzten(!) - LED Rand des Monitors mit trockenem Tuch abgewischt, Funktion nun einwandfrei! | Barcode einwandfrei lesbar aber Monitor reagiert nicht, mit Verzögerung und nur sporadisch auf Berührungen; Videos der Fehlfunktion angefertigt. Testweise den - kaum verschmutzten(!) - LED Rand des Monitors mit trockenem Tuch abgewischt, Funktion nun einwandfrei! | SWO-155022 | No | Barcode flawlessly readable but monitor does not respond, with delay and only sporadically to touches; videos of the malfunction. Test-wise the - hardly dirty(!) - LED edge of the monitor wiped with dry cloth, function now flawless! | Barcode flawlessly readable but monitor does not respond, with delay and only sporadically to touches; videos of the malfunction. Test-wise the - hardly dirty(!) - LED edge of the monitor wiped with dry cloth, function now flawless! | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Elektrode ersetzt. Weiteres Verfahren im Rahmen der heutigen HJW durchgeführt. | Elektrode ersetzt. Weiteres Verfahren im Rahmen der heutigen HJW durchgeführt. | SWO-154919 | No | electrode replaced. Further procedures carried out within the framework of today's HJW. | electrode replaced. Further procedures carried out within the framework of today's HJW. | Assisted Replacement | Electrode | Electrode | pCO2
ABL825 FLEX | Laktat Elektrode (E7077 XB04) geprüft und ausgetauscht, Funktionskontrolle ok | Laktat Elektrode (E7077 XB04) geprüft und ausgetauscht, Funktionskontrolle ok | SWO-155001 | No | Lactate electrode (E7077 XB04) tested and replaced, function control ok | Lactate electrode (E7077 XB04) tested and replaced, function control ok | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX | Had customer change the waste pump tubes and inlet gasket-No change | Replaced listed parts Cleaned waste lines and test for leaking-OK | D1D266E12C5A880B | No | Had customer change the waste pump tubes and inlet gasket-No change | Replaced listed parts Cleaned waste lines and test for leaking-OK | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | Visited site to obtain sample patient information from site, service dump from machine and test operational parts of the machine which could affect patient samples. Tech line opened and machine information sent for investigation. Machine remains operational for use. | Visited site to obtain sample patient information from site, service dump from machine and test operational parts of the machine which could affect patient samples. Tech line opened and machine information sent for investigation. Machine remains operational for use. | SWO-155091 | No | Visited site to obtain sample patient information from site, service dump from machine and test operational parts of the machine which could affect patient samples. Tech line opened and machine information sent for investigation. Machine remains operational for use. | Visited site to obtain sample patient information from site, service dump from machine and test operational parts of the machine which could affect patient samples. Tech line opened and machine information sent for investigation. Machine remains operational for use. | Unknown | Performance | Techline | Techline
AQT90 FLEX | Neustart | Neustart | SWO-155036 | No | Restart | Restart | Defect | SW | Analyzer Control Module | His/lis
AQT90 FLEX | Solution Pack geliefert | Solution Pack geliefert | SWO-155037 | Yes | Solution Pack Delivered | Solution Pack Delivered | Re classified RMED | Logistical upSWO
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-155170 | No | Repair after replacing test card | Repair after replacing test card
ABL80 FLEX BASIC |  | B | SWO-155119 | No
ABL80 FLEX BASIC | replaced the solution pack with new, now instrument working fine. | replaced the solution pack with new, now instrument working fine. | SWO-155990 | No | replaced the solution pack with new, now instrument working fine. | replaced the solution pack with new, now instrument working fine.
ABL80 CO-OX | Upon arrival, analysis was disabled due to a possible fluid blockage | Performed a flush while manually activating each valve. Removed oxi module and removed and flushed all tubings. Reassembled analyzer. Ran a few System Cycles and then all 3 levels of manual QC and all passed. | D1DE29A5522FAFB3 | No | Upon arrival, analysis was disabled due to a possible fluid blockage | Performed a flush while manually activating each valve. Removed oxi module and removed and flushed all tubings. Reassembled analyzer. Ran a few System Cycles and then all 3 levels of manual QC and all passed.
ABL90 FLEX | LP XV03 durch LP Lot XR05 ersetzt, Reklamationsausdruck erstellt und an RDEU gemailt. Funktion OK | LP XV03 durch LP Lot XR05 ersetzt, Reklamationsausdruck erstellt und an RDEU gemailt. Funktion OK | SWO-155440 | No | LP XV03 replaced by LP Lot XR05, complaint expression created and emailed to RDEU. Function OK | LP XV03 replaced by LP Lot XR05, complaint expression created and emailed to RDEU. Function OK | Assisted Replacement | SP | SP | FT
ABL90 |  | B | D1D3D04FA25A4E19 | No |  |  | Defect | HW | PM Kit | PM Kit
ABL90 | Replaced the inlet gasket, connector and pinch valve tubings and waste tubing. Analyzer worked for five days then the errors returned. Replaced the pinch valve and the analyzer worked for 1 day. Cleaned the pO2 detector and the analyzer worked for a day before the errors returned. User also ran all three levels of QC (in the QC mode) and the Hb results were completly missing on all three levels and no errors were printed or displayed. Ran the external QC again and the analyzer worked as it should. | I have now replaced the entire wet section from a demo analyzer because the 903-183 Sensor Interface Board is not available in stock and there are none in the field. The analyzer is working at this time, but the analyzer will usually work for one to five days after performing service before the errors return. User has stated that if the errors return again they want the analyzer out of the hospital, either replaced or refunded and they will get an analyzer from another company. | D1D52FC0185490A1 | No | Replaced the inlet gasket, connector and pinch valve tubings and waste tubing. Analyzer worked for five days then the errors returned. Replaced the pinch valve and the analyzer worked for 1 day. Cleaned the pO2 detector and the analyzer worked for a day before the errors returned. User also ran all three levels of QC (in the QC mode) and the Hb results were completly missing on all three levels and no errors were printed or displayed. Ran the external QC again and the analyzer worked as it should. | I have now replaced the entire wet section from a demo analyzer because the 903-183 Sensor Interface Board is not available in stock and there are none in the field. The analyzer is working at this time, but the analyzer will usually work for one to five days after performing service before the errors return. User has stated that if the errors return again they want the analyzer out of the hospital, either replaced or refunded and they will get an analyzer from another company. | Revisit required | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX PLUS |  | B | D1D3B585A24496BF | Yes
ABL90 FLEX PLUS |  | B | D1D3A84CA20251AC | Yes |  |  | Re classified RMED | Logistal PER
ABL90 FLEX PLUS | Reinstalled analyser software, analyser froze twice during reinstall, asked customer to monitor. | analyser had to be shutdown with pc Button, no sol pack changes in the log, no user actions, screen saver screen frozen in the middle, users cleaned screen but couldn't connect via aqure monitor status shown as disconnected. inlet was closed. escalated to JD he will visit and reinstall XPE and AU SW | SWO-155376 | No | Reinstalled analyser software, analyser froze twice during reinstall, asked customer to monitor. | analyser had to be shutdown with pc Button, no sol pack changes in the log, no user actions, screen saver screen frozen in the middle, users cleaned screen but couldn't connect via aqure monitor status shown as disconnected. inlet was closed. escalated to JD he will visit and reinstall XPE and AU SW | Revisit required | SW | SW | SW
ABL90 | Se les indica que cambien gasket SP y. cal oK pero valores pO2 en QC fuera de rango. Tras cambiar otro SP al principio pO2 sigue igual, luego se arregla y da QC B y C fuera de rango en Ca. Comprobación y limpieza equipo, cal y Qc ahora bien. Formación personal | Limpieza. Parece que podia haber algun resto | D1D3AB34A217E59F | No | They are instructed to change SP and lime oK gasket but pO2 values in out-of-range QC. After switching another SP to the beginning pO2 remains the same, then it is fixed and gives QC B and C out of range in Ca. Checking and cleaning equipment, lime and Qc now well. Personal training | Cleaning. Looks like there could be some remnants. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX |  | B | SWO-155132 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
ABL835 FLEX | Se sustituye láser de autocheck. Se comprueba que lee correctamente todos los códigos del carrusel. | B | D1DB6F3F66451F7F | No | Autocheck laser is replaced. It is verified that it reads all the codes on the carousel correctly. |  | Defect | HW | Auto Check Module | Barcode reader, AC
ABL825 FLEX | Analyzer is only being used to test the interface. Found rinse solution not being pumped to the pH BG module rinse valve. | Cleaned the pH BG module rinse and gas valves. Remembraned reference and Ca electrodes because they were overdue. Performed a tHb calibration because it was overdue. | D1D3B7E1A230EA05 | No | Analyzer is only being used to test the interface. Found rinse solution not being pumped to the pH BG module rinse valve. | Cleaned the pH BG module rinse and gas valves. Remembraned reference and Ca electrodes because they were overdue. Performed a tHb calibration because it was overdue. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL810 FLEX | Einspielen der entsprechenden Astm.ini nicht erfolgreich. Analysator bleibt im Laden der Applikations Software stehen. Neueinstellungen der Applikations Software auch nicht geholfen. Erst einspielen einer Astm.ini von einem Analysator im Haus der angebunden ist brachte den Erfolg. Mit Herrn Handl von Synlab gesprochen. Daten gesendet, Daten kommen an. Funktion wieder in Ordnung. | Einspielen der entsprechenden Astm.ini nicht erfolgreich. Analysator bleibt im Laden der Applikations Software stehen. Neueinstellungen der Applikations Software auch nicht geholfen. Erst einspielen einer Astm.ini von einem Analysator im Haus der angebunden ist brachte den Erfolg. Mit Herrn Handl von Synlab gesprochen. Daten gesendet, Daten kommen an. Funktion wieder in Ordnung. | SWO-155358 | No | Playing the corresponding Astm.ini not successful. Analyzer stops in the loading of the application software. New settings of the application software also did not help. Only playing an Astm.ini by an analyzer in the house that is connected brought the success. Spoke to Mr. Handl of Synlab. Data sent, data arrives. Function back in order. | Playing the corresponding Astm.ini not successful. Analyzer stops in the loading of the application software. New settings of the application software also did not help. Only playing an Astm.ini by an analyzer in the house that is connected brought the success. Spoke to Mr. Handl of Synlab. Data sent, data arrives. Function back in order. | Unknown | His/lis | His/lis | His/lis
ABL810 FLEX | Reinigung der Ranbereiches nicht erfolgreich. Polieren des Randes mit Polierpaste. Rand wieder sauber. Reaktion besser, aber nicht gut genug. Neukalibration des Displays im Service. Funktion wieder normal, Anwender zufrieden. Kalibration und Qualitätskontrolle in Ordnung. | Reinigung der Ranbereiches nicht erfolgreich. Polieren des Randes mit Polierpaste. Rand wieder sauber. Reaktion besser, aber nicht gut genug. Neukalibration des Displays im Service. Funktion wieder normal, Anwender zufrieden. Kalibration und Qualitätskontrolle in Ordnung. | SWO-155361 | No | Cleaning of the ran area not successful. Polishing the edge with polishing paste. Edge clean again. Reaction better, but not good enough. Recalibration of the display in service. Function back to normal, users satisfied. Calibration and quality control in order. | Cleaning of the ran area not successful. Polishing the edge with polishing paste. Edge clean again. Reaction better, but not good enough. Recalibration of the display in service. Function back to normal, users satisfied. Calibration and quality control in order. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | 更换910-275修复(by zhenjie) | 更换910-275修复(by zhenjie) | SWO-154798 | No | Replacement 910-275 repair (by zhenjie) | Replacement 910-275 repair (by zhenjie) | Defect | HW | Printer module | Printer module
AQT90 FLEX | COMPROBACIÓN DE APERTURA Y CIERRE DE LA PUERTA DE ENTRADA DE MUESTRAS, AJUSTE DE GUIAS Y DE SUS POLEAS, ENGRASE DE LAS GUIAS, COMPROBACIÓN DEL CORRECTO FUNCIONAMIENTO, COMPROBACIÓN DE REGISTROS DE DATOS DEL EQUIPO. | B | D1D3B461A2434F39 | Yes | OPENING AND CLOSING CHECK OF THE SAMPLE INPUT DOOR, ADJUSTMENT OF GUIDES AND ITS POLEAs, GRASE THE GUIDES, CHECK THE CORRECT OPERATION, CHECK OF DATA RECORDS OF THE EQUIPMENT.
AQT90 FLEX |  | B | D1D3F658A24C694A | No |  |  | Defect | Consumables | LQC Kit | LQC Kit | Myo
TCM/Tosca | Replaced the TOSCA module, reseated the CF card. | unit | D1D40586A24BAF7F | No | Replaced the TOSCA module, reseated the CF card. | unit | User error | HW | Analyzer dropped
TCM/Tosca | I was able to reproduce error. | replaced CPU and error was fixed | D1D3E00FA23B3C91 | No | I was able to reproduce error. | replaced CPU and error was fixed | Defect | HW | PCB, Main CPU | PCB, Main CPU
ABL80 FLEX BASIC |  | B | SWO-155500 | No
ABL80 FLEX CO-OX | Gerät gereinigt und LP-Anschlüsse mit einer Luer-Spritze geweitet geöffnet, anschließend Gerät in Betrieb genommen | Gerät gereinigt und LP-Anschlüsse mit einer Luer-Spritze geweitet geöffnet, anschließend Gerät in Betrieb genommen | SWO-155785 | No | Device cleaned and LP connections widened with a Luer syringe, then put the device into operation | Device cleaned and LP connections widened with a Luer syringe, then put the device into operation
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-155455 | No | Replace the test card | Replace the test card
ABL90 | All tubings and the pinch valve have been replaced, the pO2 detector has been replaced, and the same error intermittently occurs. | Temporarily exchanged the entire sensor section to test the analyzer since no sensor interface PCBs were available. | D1D96CFBF03820A2 | No | All tubings and the pinch valve have been replaced, the pO2 detector has been replaced, and the same error intermittently occurs. | Temporarily exchanged the entire sensor section to test the analyzer since no sensor interface PCBs were available. | Revisit required | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX PLUS |  | B | D1D4E9F3182A5AA6 | Yes |  |  | Re classified RMED | Logistal PER
ABL90 FLEX PLUS | Startphase QC und Kalibration. | Startphase QC und Kalibration. | SWO-155580 | No | Start phase QC and calibration. | Start phase QC and calibration. | Defect | HW | Auto Inlet module | Inlet Gasket | 1342
ABL90 FLEX PLUS |  | B | D1D4E872185A5AA3 | Yes |  |  | Re classified RMED | Logistal PER
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-155812 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | SC | 1343
AQT90 FLEX | 更换903-364后修复(by zhenjie) | 更换903-364后修复(by zhenjie) | SWO-155428 | No | Repair after replacement 903-364 (by zhenjie) | Repair after replacement 903-364 (by zhenjie) | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | 更换备用件903-210修复(by zhenjie) | 更换备用件903-210修复(by zhenjie) | SWO-155746 | No | Replace spare parts 903-210 repair (by zhenjie) | Replace spare parts 903-210 repair (by zhenjie) | Defect | HW | Process Module | PCB, Cup dryer | 17
AQT90 FLEX | Verificação dos registos de dados. Teste de verificação de unidade optica OK. Realização de Kit anual e substituição de roda processadora e unidade de lavagem. Inicialização. Calibração de lote de Myo. Verificação com QC. O erro mantem-se, Myo abaixo do intervalo. Substituição da unidade optica. Calibração de Myo, Tnl e CKMB. Verificação com QC. O erro mantem-se. | B | D1D4D266181486D0 | No | Verification of data records. Optical drive Verification Test OK. Realization of annual Kit and replacement of processing wheel and washing unit. Startup. Myo Batch calibration. Check with QC. The error keeps, Myo below the interval. Replacement of the Optica unit. Calibration of Myo, TNL and CKMB. Check with QC. The error keeps. |  | Defect | HW | Process Module | Optical Unit | Myo - LQC
ABL825 FLEX | Systeem doorgespoeld demiwater. Na twee calibraties controles gemeten, OK. | Systeem doorgespoeld demiwater. Na twee calibraties controles gemeten, OK. | SWO-155770 | No | System washed down demi water. After two calibrations controls measured, OK. | System washed down demi water. After two calibrations controls measured, OK. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL835 FLEX | DMS起動不良。CPU交換（MB662⇒DMS910）、HDD交換、ｸﾘｰﾝｲﾝｽﾄｰﾙ実施、LVDS PCB交換、CDDtoUSB交換、ﾌﾟﾘﾝﾀｰﾍｯﾄﾞ交換。 電解質QC外れの為、Ref電極交換。 | DMS起動不良。CPU交換（MB662⇒DMS910）、HDD交換、ｸﾘｰﾝｲﾝｽﾄｰﾙ実施、LVDS PCB交換、CDDtoUSB交換、ﾌﾟﾘﾝﾀｰﾍｯﾄﾞ交換。 電解質QC外れの為、Ref電極交換。 | SWO-156031 | No | DMS startup failure. CPU replacement (MB662⇒DMS910), HDD replacement, clean installation, LVDS PCB replacement, CDDtoUSB replacement, printer head replacement. Ref electrode replacement for electrolyte QC out. | DMS startup failure. CPU replacement (MB662⇒DMS910), HDD replacement, clean installation, LVDS PCB replacement, CDDtoUSB replacement, printer head replacement. Ref electrode replacement for electrolyte QC out. | Defect | HW | DMS_PC | CPU refurb
ABL815 FLEX | Leih OXI Modul eingesetzt, Kunden OXI Modul zur Reparatur abgezogen. Lösungen laufen aus dem Spritzen Einlass, Einlassdichtung erneuert, Luftloch gespült. | Lampeneinheit nachgestellt, keinen Erfolg. Neon Position schwankt stark, scheinbar durch Erschütterung. Spectrometer erneuert. Neon Position stabil bei ~ 0,400. | SWO-155716 | No | Loan OXI module used, customer OXI module deducted for repair. Solutions run out of the syringe inlet, inlet seal renewed, air hole flushed. | lamp unit, no success. Neon position fluctuates strongly, apparently due to shock. Spectrometer. Neon position stable at .400. | Defect | HW | Oxi module | Oxi module
ABL825 FLEX | Erneut Ausreißer nach unten bei QC L3 Messung, mit Frau Dietzel vereinbart, dass das Gerät weiter beobachtet wird und ggf. kommende Woche ein Datenextrakt erfolgt, der bei RDEU ausgewertet werden soll zwecks Abstimmung der weiteren Vorgehensweise. | Erneut Ausreißer nach unten bei QC L3 Messung; mit Frau Dietzel vereinbart, dass das Gerät weiter beobachtet wird und ggf. kommende Woche ein Datenextrakt erfolgt, der bei RDEU ausgewertet werden soll zwecks Abstimmung der weiteren Vorgehensweise. | SWO-155692 | No | Again outliers down at QC L3 measurement; With Ms. Dietzel, she agreed that the device will continue to be monitored and, if necessary, a data extract will be carried out next week, which will be evaluated at RDEU in order to coordinate the further procedure. | Again outliers down at QC L3 measurement; With Ms. Dietzel, she agreed that the device will continue to be monitored and, if necessary, a data extract will be carried out next week, which will be evaluated at RDEU in order to coordinate the further procedure. | Unknown | Performance | QC | pCO2
ABL815 FLEX | Erneut Ca Driften sowie abnormale Elektroden-Stati pH und E´lyte sowie Grenzverletzungen bei Ca nach unten bei L1 und L3 => Referenzelektrode sowie Kontaktkappe ersetzt. Neustart, Initialisierung Kalibrationen, Stati nun satt im Bereich. Funktionstest und Endkontrolle L1 und L3 Messung => Ca Werte nun genau mittig. | Erneut Ca Driften sowie abnormale Elektroden-Stati pH und E´lyte sowie Grenzverletzungen bei Ca nach unten bei L1 und L3 => Referenzelektrode sowie Kontaktkappe ersetzt. Neustart, Initialisierung Kalibrationen, Stati nun satt im Bereich. Funktionstest und Endkontrolle L1 und L3 Messung => Ca Werte nun genau mittig. | SWO-155696 | No | Again Ca drifting as well as abnormal electrode stati pH and E'lyte as well as boundary injuries at Ca replaced downwards at L1 and L3 => reference electrode as well as contact cap. Restart, initialization calibrations, stati now full in the range. Functional test and final control L1 and L3 measurement => Ca values now exactly in the middle. | Again Ca drifting as well as abnormal electrode stati pH and E'lyte as well as boundary injuries at Ca replaced downwards at L1 and L3 => reference electrode as well as contact cap. Restart, initialization calibrations, stati now full in the range. Functional test and final control L1 and L3 measurement => Ca values now exactly in the middle. | Assisted Replacement | Electrode | Electrode | Reference
ABL825 FLEX | Oxi Modul erneuert, Gerät kalibriert. | Oxi Modul erneuert, Gerät kalibriert. | SWO-155731 | No | Oxi module renewed, device calibrated. | Oxi module renewed, device calibrated. | Defect | HW | Oxi module | Oxi module
ABL825 FLEX | ABL rebootet, danach Scannerfunktion OK. | ABL rebootet, danach Scannerfunktion OK. | SWO-155760 | No | ABL reboots, then scanner function OK. | ABL reboots, then scanner function OK. | Defect | SW | DMS_PC | Off/On
AQT90 FLEX |  | B | D1D544B118571C74 | Yes
ABL80 FLEX CO-OX | Fehlersuche, Austausch Lösungspack, Funktionskontrolle , neuer Lösungspack undicht ( Lot: 229743), nach Reinigung und erneuter Installierung ok Eine spätere Blutmessung ergab einen pO2 Wert von 21,2 kPa, QC und Kalibrierungen iO. , neue Sensorkassette installiert | Sr. Katrin bittet um einen erneute Überprüfung des ABL | SWO-155776 | No | Troubleshooting, replacement solution pack, function control, new solution pack leaking ( lot: 229743), after cleaning and reinstallation ok A later blood measurement showed a pO2 value of 21.2 kPa, QC and calibrations iO. , new sensor cartridge installed | Sr. Katrin asks for a re-examination of the ABL
ABL90 FLEX | Touch Screen aus UKE alt Gerät erneuert Kalibration und QC durchgeführt | Touch Screen aus UKE alt Gerät erneuert Kalibration und QC durchgeführt | SWO-155814 | No | Touch screen from UKE old device renewed calibration and QC performed | Touch screen from UKE old device renewed calibration and QC performed | Defect | HW | DMS_PC | Touch screen
ABL90 |  | B | D1D640668E46B7C2 | No |  |  | Defect | HW | Sample mixer | Sample mixer
ABL90 FLEX | Drucker-Einheit (P0154N009) eingebaut und (P0176N012) ausgebaut, mehrere Ausdrucke durchgeführt --> i.O., System und Funktion i.O. | Drucker-Einheit (P0154N009) eingebaut und (P0176N012) ausgebaut, mehrere Ausdrucke durchgeführt --> i.O., System und Funktion i.O. | SWO-155822 | No | Printer unit (P0154N009) built-in and (P0176N012) removed, multiple printouts performed --> i.O., system and function i.O. | Printer unit (P0154N009) built-in and (P0176N012) removed, multiple printouts performed --> i.O., system and function i.O. | Defect | HW | Printer module | Printer module
ABL90 FLEX PLUS | Replaced bent probe. Done complete service test all ok after redoing clot detect clot remove. | Replaced bent probe. Done complete service test all ok after redoing clot detect clot remove. | SWO-156495 | No | Replaced bent probe. Done complete service test all ok after redoing clot detect clot remove. | Replaced bent probe. Done complete service test all ok after redoing clot detect clot remove. | Replaced | HW | Auto Inlet module | Probe bend
ABL815 FLEX | Lampeneinheit nachgestellt, keinen Erfolg. Neon Position schwankt stark, scheinbar durch Erschütterung. Spectrometer erneuert. Neon Position stabil bei ~ 0,400. | Lampeneinheit nachgestellt, keinen Erfolg. Neon Position schwankt stark, scheinbar durch Erschütterung. Spectrometer erneuert. Neon Position stabil bei ~ 0,400. | SWO-155894 | No | lamp unit, no success. Neon position fluctuates strongly, apparently due to shock. Spectrometer. Neon position stable at .400. | lamp unit, no success. Neon position fluctuates strongly, apparently due to shock. Spectrometer. Neon position stable at .400. | Defect | HW | Oxi module | Spectrometer
ABL800 BASIC | Provided the document regarding HCO3- calculation from reference manual and added HCO3-(P) in the patient report layout . | Provided the document regarding HCO3- calculation from reference manual and added HCO3-(P) in the patient report layout . | SWO-155971 | Yes | Provided the document regarding HCO3- calculation from reference manual and added HCO3-(P) in the patient report layout . | Provided the document regarding HCO3- calculation from reference manual and added HCO3-(P) in the patient report layout . | Re classified RMED | courtesy
AQT90 FLEX |  | B | D1D678DF8E532135 | No |  |  | Defect | SW | Analyzer Control Module | Off/On | 167
AQT90 FLEX | NEEDLE WASH VE FİLİTRENİN TIKALI OLDUĞU TESPİT EDİLDİ.TIKANIKLIK GİDERİLEREK SORUN ÇÖZÜLDÜ. | NEEDLE WASH VE FİLİTRENİN TIKALI OLDUĞU TESPİT EDİLDİ.TIKANIKLIK GİDERİLEREK SORUN ÇÖZÜLDÜ. | SWO-156698 | Yes | NEEDLE WASH AND THE FILM WAS CONFIRMED TO BE TIKALI.THE PROBLEM WAS SOLVED BY SENDING TIKANIKLIK. | NEEDLE WASH AND THE FILM WAS CONFIRMED TO BE TIKALI.THE PROBLEM WAS SOLVED BY SENDING TIKANIKLIK.
AQT90 FLEX | Parameter hinzugefügt und Testmessungen laufen gelassen. | Parameter hinzugefügt und Testmessungen laufen gelassen. | SWO-155833 | No | parameters added and test measurements run. | parameters added and test measurements run. | Setup | SW | Patient Layout | Patient Layout
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-156044 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX | 更换903-158后修复(by zhenjie) | 更换903-158后修复(by zhenjie) | SWO-155754 | No | Repair after replacement 903-158 (by zhenjie) | Repair after replacement 903-158 (by zhenjie) | Defect | HW | Basic Unit | Solution Pack ID PCB
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-155983 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | Leakage current
AQT90 FLEX | 更换903-364后修复(by zhenjie) | 更换903-364后修复(by zhenjie) | SWO-155786 | Yes | Repair after replacement 903-364 (by zhenjie) | Repair after replacement 903-364 (by zhenjie) | Re classified RMED | HW | Process Module | Optical Unit | 131
ABL827 FLEX | ｲﾝﾚｯﾄﾌﾟﾛｰﾌﾞ交換。ｲﾝﾚｯﾄﾗﾊﾞｰﾁｭｰﾌﾞ交換。廃液ﾎﾟﾝﾌﾟﾁｭｰﾌﾞ交換。 廃液ライン清掃。 ｲﾝﾚｯﾄから水滴が噴き出す現象と推測される。 | B | SWO-157123 | No | Inlet probe replacement. Inlet rubber tube replacement. Waste pump tube replacement. Waste line cleaning. It is presumed that the phenomenon that water droplets spout from the inlet. |  | Defect | HW | Inlet module | Probe
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-156050 | No | Repair after replacing test card | Repair after replacing test card
ABL80 FLEX BASIC |  | B | SWO-156083 | No
ABL90 |  | B | D1DA0933F024B45E | Yes |  |  | Re classified RMED | Phone support
ABL90 FLEX | IP Adressen überprüft - keine IP Adressen mehr eingetragen, Pat. Berichte seit 24.12. 07:59 nicht mehr übertragen, 24.12. 07:59 Power on Neustart des ABL durchgeführt, IP Adressen eingetragen, Speicherbefehl ausgeführt, Neustart des ABL, LP90 wurde kurz vorher eingesetzt und machte Probleme bei Init., LP90 war nicht richtig aktiviert, LP90 komplett aktiviert, LP90 eingesetzt, Startphase von 15 Minuten ist gestartet, nach Startphase Glu weiterhin in Level C fehlerhaft, Tausch der SC90 und Initialisierung, System- und Funktionskontrolle in Ordnung | IP Adressen überprüft - keine IP Adressen mehr eingetragen, Pat. Berichte seit 24.12. 07:59 nicht mehr übertragen, 24.12. 07:59 Power on Neustart des ABL durchgeführt, IP Adressen eingetragen, Speicherbefehl ausgeführt, Neustart des ABL, LP90 wurde kurz vorher eingesetzt und machte Probleme bei Init., LP90 war nicht richtig aktiviert, LP90 komplett aktiviert, LP90 eingesetzt, Startphase von 15 Minuten ist gestartet, nach Startphase Glu weiterhin in Level C fehlerhaft, Tausch der SC90 und Initialisierung, System- und Funktionskontrolle in Ordnung | SWO-156125 | No | IP addresses checked - no more IP addresses entered, Pat. Reports since 24.12. 07:59 no longer transmitted, 24.12. 07:59 Power on restart of the ABL performed, IP addresses entered, memory command executed, restart of the ABL, LP90 was used shortly before and caused problems with Init., LP90 was not enabled correctly, LP90 fully activated, LP90 inserted, start phase of 15 minutes is started, after start phase Glu continues to be faulty in level C, swapping sc90 and initialization, system and function control all right | IP addresses checked - no more IP addresses entered, Pat. Reports since 24.12. 07:59 no longer transmitted, 24.12. 07:59 Power on restart of the ABL performed, IP addresses entered, memory command executed, restart of the ABL, LP90 was used shortly before and caused problems with Init., LP90 was not enabled correctly, LP90 fully activated, LP90 inserted, start phase of 15 minutes is started, after start phase Glu continues to be faulty in level C, swapping sc90 and initialization, system and function control all right | Assisted Replacement | SC | SC | Glu
ABL90 | The solution pack has continuous Replace Solution Pack error messages, and the sensor cassette has intermittent Leakage Detected errors. | Replace the solution pack and sensor cassette. | D1D988A0F04920AB | No | The solution pack has continuous Replace Solution Pack error messages, and the sensor cassette has intermittent Leakage Detected errors. | Replace the solution pack and sensor cassette. | Assisted Replacement | SC | SC | FT
ABL90 FLEX PLUS | Sensorkassette wurde nach telefonischer Empfehlung vorab getauscht (Reklamation der SC durch MarWa). Bei Einsatz vor Ort wurde der Schlauch des Quetschventils erneuert. Quetschventil ohne Beanstandung. Analysator manuell gespült. Nadel überprüft. Servicetests und pO2 completeTest i.O. Kalibration der noch instabilen Parameter und wiederholen der QC. Glu und Lac noch knapp über Grenzwert, Kalibration i.O. Testmessung durch Personal durchgeführt, i.O. Hinweis auf die Präanalytik und Messmethodik | Sensorkassette wurde nach telefonischer Empfehlung vorab getauscht (Reklamation der SC durch MarWa). Bei Einsatz vor Ort wurde der Schlauch des Quetschventils erneuert. Quetschventil ohne Beanstandung. Analysator manuell gespült. Nadel überprüft. Servicetests und pO2 completeTest i.O. Kalibration der noch instabilen Parameter und wiederholen der QC. Glu und Lac noch knapp über Grenzwert, Kalibration i.O. Testmessung durch Personal durchgeführt, i.O. Hinweis auf die Präanalytik und Messmethodik | SWO-156123 | No | Sensor cassette was exchanged in advance according to telephone recommendation (complaint of the SC by MarWa). When used on site, the hose of the crushing valve was renewed. Squeezing valve without complaint. Analyzer manually rinsed. Needle checked. Service tests and pO2 completeTest in the o. Calibration of the still unstable parameters and repeat the QC. Glu and Lac still just above limit, calibration i.O. Test measurement carried out by personnel, i.e. Reference to preanalysis and measurement methodology | Sensor cassette was exchanged in advance according to telephone recommendation (complaint of the SC by MarWa). When used on site, the hose of the crushing valve was renewed. Squeezing valve without complaint. Analyzer manually rinsed. Needle checked. Service tests and pO2 completeTest in the o. Calibration of the still unstable parameters and repeat the QC. Glu and Lac still just above limit, calibration i.O. Test measurement carried out by personnel, i.e. Reference to preanalysis and measurement methodology | User error | Measurement | Measurement | Pre analytics
ABL90 FLEX |  | B | SWO-156087 | No |  |  | Defect | HW | Main PCB | Main PCB
AQT90 FLEX | 更换903-364后修复 | B | SWO-156072 | Yes | Repair after replacement 903-364 |  | Re classified RMED | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364修复(by andrew) | 更换903-364修复(by andrew) | SWO-156049 | Yes | Replacement 903-364 repair (by andrew) | Replacement 903-364 repair (by andrew) | Re classified RMED | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-156088 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL800 BASIC |  | B | SWO-156089 | No |  |  | Assisted Replacement | Electrode | Electrode | Na+
ABL800 BASIC | Clean the reference electrode chamber thoroughly ,clean the gasket,run 2 point he calibration,after successful completion run 2-3 samples ,got the desired results. | Clean the reference electrode chamber thoroughly ,clean the gasket,run 2 point he calibration,after successful completion run 2-3 samples ,got the desired results. | SWO-157034 | No | Clean the reference electrode chamber thoroughly ,clean the gasket,run 2 point he calibration,after successful completion run 2-3 samples ,got the desired results. | Clean the reference electrode chamber thoroughly ,clean the gasket,run 2 point he calibration,after successful completion run 2-3 samples ,got the desired results. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-156218 | No | Replacement test card repair | Replacement test card repair
ABL80 FLEX Basic version |  | B | D1DB45E96633C608 | No
ABL9 | Gerät überprüft, Sensorkassette gespült, Kontakte kontrolliert, SC reinstalliert, Servicedaten gesichert | Gerät überprüft, Sensorkassette gespült, Kontakte kontrolliert, SC reinstalliert, Servicedaten gesichert | SWO-156464 | No | Device checked, sensor cartridge flushed, contacts checked, SC reinstalled, service data backed up | Device checked, sensor cartridge flushed, contacts checked, SC reinstalled, service data backed up
ABL90 FLEX | LP erneuert | LP erneuert | SWO-156180 | No | LP renewed | LP renewed | Assisted Replacement | SP | SP Expired | SP Expired
ABL90 FLEX | SC erneuert | SC erneuert | SWO-156850 | No | SC renewed | SC renewed | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | Gerät gespült. Alle Wege frei.Inlet Gasket und Nadel ok. SP eingebaut. Gerät ok. | Gerät gespült. Alle Wege frei.Inlet Gasket und Nadel ok. SP eingebaut. Gerät ok. | SWO-156342 | No | device rinsed. All paths clear. Inlet Gasket and needle ok. SP built-in. Device ok. | device rinsed. All paths clear. Inlet Gasket and needle ok. SP built-in. Device ok. | Assisted Replacement | SP | SP | FT
ABL90 | Defective battery took out the power supply | Repalced SP and Battery | D1DB85056606F16C | No | Defective battery took out the power supply | Repalced SP and Battery | Defect | HW | Power Supply | Power Supply
ABL90 FLEX PLUS | replaced missing inlet gasket and straightened bent probe. Done back flush. All ok. | replaced missing inlet gasket and straightened bent probe. Done back flush. All ok. | SWO-156502 | No | replaced missing inlet gasket and straightened bent probe. Done back flush. All ok. | replaced missing inlet gasket and straightened bent probe. Done back flush. All ok. | User error | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Stripped and found metal latch attached to upper sensor bay door is bent downwards. Bent slightly upwards as to align photo sensor pick up with slit in latch. Door tested ok. | Stripped and found metal latch attached to upper sensor bay door is bent downwards. Bent slightly upwards as to align photo sensor pick up with slit in latch. Door tested ok. | SWO-156507 | No | Stripped and found metal latch attached to upper sensor bay door is bent downwards. Bent slightly upwards as to align photo sensor pick up with slit in latch. Door tested ok. | Stripped and found metal latch attached to upper sensor bay door is bent downwards. Bent slightly upwards as to align photo sensor pick up with slit in latch. Door tested ok. | Adjustment | HW | Basic Unit | Sensor cassette Door
ABL90 FLEX PLUS | Gerät geprüft. Touchscreen aus GS geholt und ausgetauscht. Startphase 50Min. abgewartet. | Gerät geprüft. Touchscreen aus GS geholt und ausgetauscht. Startphase 50Min. abgewartet. | SWO-156648 | No | device checked. Touchscreen is taken from GS and replaced. start-up phase 50min. | device checked. Touchscreen is taken from GS and replaced. start-up phase 50min. | Defect | HW | DMS_PC | Touch screen
ABL827 FLEX | Arrived to do software upgrade, but encountered 0931 errors after upgrade | Upgraded software to ver 6.16. Replaced membrane valves. Replaced one dummy electrode in Met II chamber. replaced measuring chamber and liquid sensor. | D1DB98226650F06C | Yes | Arrived to do software upgrade, but encountered 0931 errors after upgrade | Upgraded software to ver 6.16. Replaced membrane valves. Replaced one dummy electrode in Met II chamber. replaced measuring chamber and liquid sensor. | Re classified RMED | SWU
ABL800 BASIC | OXi Controller PCB交換（903-654 R0002 N004） | B | SWO-156385 | No | OXi Controller PCB交換（903-654 R0002 N004） |  | Defect | HW | Oxi module | PCB, Oxi module
ABL815 FLEX | Inletmodul demontiert, geprüft und gereinigt, Funktionskontrolle ok | Inletmodul demontiert, geprüft und gereinigt, Funktionskontrolle ok | SWO-156308 | No | Inlet module dismantled, tested and cleaned, function control ok | Inlet module dismantled, tested and cleaned, function control ok | Cleaned | Fluidic pathway | Inlet module | Fluidic pathway
ABL80 FLEX BASIC | 更换902-972后修复 by andrew | 更换902-972后修复 by andrew | SWO-156211 | No | Repair by andrew after replacement 902-972 | Repair by andrew after replacement 902-972
ABL80 CO-OX | Customer tried another SP it did not help Fedex parts failed to be delivered to the appropiate address. | I found the problem to be the latch(Customer Service sent parts to the wrong place) Pushed SP to at least get it to work until Jose gets here tomorrow. Calibrated unit a couple of times, but SP would fail to work Performed a Zero cal Reinstalled the SP I checked zero point and it has shifted again Jose will replaced parts tomorrow. | D1DC856DDC3B24A0 | No | Customer tried another SP it did not help Fedex parts failed to be delivered to the appropiate address. | I found the problem to be the latch(Customer Service sent parts to the wrong place) Pushed SP to at least get it to work until Jose gets here tomorrow. Calibrated unit a couple of times, but SP would fail to work Performed a Zero cal Reinstalled the SP I checked zero point and it has shifted again Jose will replaced parts tomorrow.
ABL90 FLEX | Sensorkassette leer | Sensorkassette leer | SWO-156563 | No | Sensorkassette read | Sensorkassette read | User error | Sample expired | Sample expired | Sample expired
ABL90 | The sensor section of the analyzer was temporarily exchanged for testing purposes. The analyzer has worked well since that exchange. Now the sensor interface PCB needs to be installed on the original sensor section and then reinstalled into the analyzer. | Replaced the sensor interface board and reinstalled the sensor section into the analyzer. External QC not available for testing. | D1DCC0F7DC33D576 | No | The sensor section of the analyzer was temporarily exchanged for testing purposes. The analyzer has worked well since that exchange. Now the sensor interface PCB needs to be installed on the original sensor section and then reinstalled into the analyzer. | Replaced the sensor interface board and reinstalled the sensor section into the analyzer. External QC not available for testing. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 |  | B | D1DD9C2F5236B39A | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Flushed inlet LS and flow path with protein remover and then flushed with water. Run LS adjust program and this is passing read 183 ok. Replaced pinch valve tube as worn. | Customer has tried to fix nothing worked. FSE required . | SWO-156609 | No | Flushed inlet LS and flow path with protein remover and then flushed with water. Run LS adjust program and this is passing read 183 ok. Replaced pinch valve tube as worn. | Customer has tried to fix nothing worked. FSE required . | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX PLUS | Allgemeine Kontrolle i.o Gerät gespült i.o Einlassdichtung und Nadel gereinigt i.o Zwischen Lösungspack und Einlassdichtung Verschmutzung gereinigt i.o Kal. und QC i.o | Allgemeine Kontrolle i.o Gerät gespült i.o Einlassdichtung und Nadel gereinigt i.o Zwischen Lösungspack und Einlassdichtung Verschmutzung gereinigt i.o Kal. und QC i.o | SWO-156534 | No | General control i.o Device rinsed i.o Inlet seal and needle cleaned i.o Between solution pack and inlet seal Pollution cleaned i.o Kal. and QC i.o | General control i.o Device rinsed i.o Inlet seal and needle cleaned i.o Between solution pack and inlet seal Pollution cleaned i.o Kal. and QC i.o | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-156672 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Reference
AQT90 FLEX | 更换903-364后修复by andrew | 更换903-364后修复by andrew | SWO-156092 | Yes | Repair by Andrew after replacement 903-364 | Repair by Andrew after replacement 903-364 | Re classified RMED | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-329和903-364后修复 by andrew | 更换903-329和903-364后修复 by andrew | SWO-156075 | Yes | Repair by Andrew after replacing 903-329 and 903-364 | Repair by Andrew after replacing 903-329 and 903-364 | Re classified RMED | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | 之前借的都有问题，现在重新借886-639，886-638 | SWO-156420 | No |  | There have been problems with previous loans, and now borrow 886-639, 886-638 | Defect | HW | Cover and Chassis | Inlet Door
ABL825 FLEX | Konfiguration geprüft und geändert, Funktionskontrolle ok | Konfiguration geprüft und geändert, Funktionskontrolle ok | SWO-156519 | No | Configuration checked and changed, function control ok | Configuration checked and changed, function control ok | Setup | SW | His/lis | His/lis
ABL820 FLEX | CHECKOUT OF SYSTEM COULD NOT IDENTIFY ISSUE. | REPLACED AUTOCHECK RIBBON CABLE AND UPGRADED SOFTWARE TO 6.16 | D1DC7C68DC201001 | No | CHECKOUT OF SYSTEM COULD NOT IDENTIFY ISSUE. | REPLACED AUTOCHECK RIBBON CABLE AND UPGRADED SOFTWARE TO 6.16 | Defect | HW | Auto Check Module | FPC Cable
ABL80 CO-OX | The analyzer had a leakage which also caused message1117 Solution pack verification not performed error. | Found the SP door hinge defective, replaced hinge. Replaced the Analog boarD, the sensor connector board and, the spine board. Replaced a new SC and, SP. Also, oximetry is failing showing error 1218 air contamination in oxi module, replaced the hemolyzer unit. Performed the ctHb calibration, passing with ranges. | D1DDF050521B8EE6 | No | The analyzer had a leakage which also caused message1117 Solution pack verification not performed error. | Found the SP door hinge defective, replaced hinge. Replaced the Analog boarD, the sensor connector board and, the spine board. Replaced a new SC and, SP. Also, oximetry is failing showing error 1218 air contamination in oxi module, replaced the hemolyzer unit. Performed the ctHb calibration, passing with ranges.
ABL90 FLEX | 4 SC und 2 LPs ausgeliefert. | 4 SC und 2 LPs ausgeliefert. | SWO-156815 | Yes | 4 SC and 2 LPs. | 4 SC and 2 LPs. | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS |  | B | D1DDA968525E51D8 | Yes
ABL90 FLEX PLUS |  | B | D1DDAB03520E51DD | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Printer unit vervangen | B | SWO-156783 | No | Replace Printer Unit |  | Defect | HW | Printer module | Printer module
ABL90 FLEX | 更换903-158修复by andrew | 更换903-158修复by andrew | SWO-156466 | No | Replacement 903-158 Repair by Andrew | Replacement 903-158 Repair by Andrew | Defect | HW | Basic Unit | Solution Pack PCB Id
ABL90 FLEX | 使用备件903-135修复 by andrew | 使用备件903-135修复 by andrew | SWO-156425 | No | Repair by andrew with spare parts 903-135 | Repair by andrew with spare parts 903-135 | Unknown | Unknown | Unknown | SC Simulator
AQT90 FLEX | 更换903-364后修复 by andrew | 更换903-364后修复 by andrew | SWO-156635 | Yes | Repair by andrew after replacing 903-364 | Repair by andrew after replacing 903-364 | Re classified RMED | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换备用光单修复by andrew | 更换备用光单修复by andrew | SWO-156705 | Yes | Replace the spare light sheet repair by andrew | Replace the spare light sheet repair by andrew | Re classified RMED | HW | Process Module | Optical Unit | 131
ABL825 FLEX |  | B | D1DD910D523DBF44 | No |  |  | User error | Consumables | Consumables | Waste bottle
ABL825 FLEX | 重启机器修复仪器 by andrew | 重启机器修复仪器 by andrew | SWO-156413 | No | Restart the machine repair instrument by andrew | Restart the machine repair instrument by andrew | Defect | HW | Power supply | Power supply
ABL835 FLEX | Skruer strammet på hemolyzeren, service og alm. tHb kalibreret, slutafprøvet. | Skruer strammet på hemolyzeren, service og alm. tHb kalibreret, slutafprøvet. | SWO-156790 | No | Screws tightened on the Hemolyzeren, service and Alm. tHb calibrated, final tested. | Screws tightened on the Hemolyzeren, service and Alm. tHb calibrated, final tested. | Defect | HW | Oxi module | Hemolyzer
ABL800 FLEX | Try to Reinstall Software. After reinstalling again after half an hour start hanging again. Need to wait till upgraded driver and software will come. External Keyboard and mouse attached for temporary use. | Try to Reinstall Software. After reinstalling again after half an hour start hanging again. Need to wait till upgraded driver and software will come. External Keyboard and mouse attached for temporary use. | SWO-156997 | No | Try to Reinstall Software. After reinstalling again after half an hour start hanging again. Need to wait till upgraded driver and software will come. External Keyboard and mouse attached for temporary use. | Try to Reinstall Software. After reinstalling again after half an hour start hanging again. Need to wait till upgraded driver and software will come. External Keyboard and mouse attached for temporary use. | Revisit required | HW | DMS_PC | Touch screen
ABL815 FLEX | Einlassdichtung erneuert, Gerät kalibriert, Gerät im OP aufgestellt. Gerät konfiguriert. Alt- Gerät in der Medizintechnik abgegeben. | B | SWO-156816 | No | Inlet seal renewed, device calibrated, device installed in the operating room. device configured. old device in medical technology. |  | Defect | HW | Inlet module | Inlet Gasket
AQT90 FLEX | Reiniging diverse sensoren systeem | Reiniging diverse sensoren systeem | SWO-156777 | No | Cleaning various sensors system | Cleaning various sensors system | Defect | SW | Analyzer Control Module | Off/On | 1999
AQT90 FLEX | Gerät überprüft, Infrarotsensor gesäubert, Daten gelöscht, Konfiguration für Archivierung geändert, Messungen eingegeben | Gerät überprüft, Infrarotsensor gesäubert, Daten gelöscht, Konfiguration für Archivierung geändert, Messungen eingegeben | SWO-156929 | No | Device checked, infrared sensor cleaned, data deleted, configuration changed for archiving, measurements entered | Device checked, infrared sensor cleaned, data deleted, configuration changed for archiving, measurements entered | Cleaned | HW | Aspiration Module | Barcode reader, Inlet
AQT90 FLEX |  | B | D1DDC81B524AE281 | No |  |  | Defect | SW | SW | Slow response
AQT90 FLEX |  | B | D1DDEBBA522B3BE6 | No |  |  | Adjustment | HW | Process Module | PCB, Cup dryer
ABL80 FLEX BASIC | 更换备件902-972修复 by andrew | 更换备件902-972修复 by andrew | SWO-156409 | No | Replacement spare parts 902-972 repair by andrew | Replacement spare parts 902-972 repair by andrew
ABL80 FLEX BASIC |  | 开错了 | SWO-156900 | No |  | It's wrong.
ABL90 FLEX | Gerät durchgespült, Startphase abgewartet -> Fehler bleibt, im Serviceprog. Fluidtransport alle Programme durchgeführt-> div. Prog. mit Fehler, an Pumpenschlauch des LP rote Verfärbung, Pumpenschlauch perforiert->es ist nur noch ein LP im gesamten Haus vorrätig-> Kunde ersetzt diesen am Montag, Gerät ausgeschaltet | Gerät durchgespült, Startphase abgewartet -> Fehler bleibt, im Serviceprog. Fluidtransport alle Programme durchgeführt-> div. Prog. mit Fehler, an Pumpenschlauch des LP rote Verfärbung, Pumpenschlauch perforiert->es ist nur noch ein LP im gesamten Haus vorrätig-> Kunde ersetzt diesen am Montag, Gerät ausgeschaltet | SWO-156996 | No | Device flushed through, start phase awaited -> Error remains, in the service program. Fluid transport all programs performed-> div. Prog. with errors, on pump hose of the LP red discoloration, pump hose perforated->it is only one LP in stock in the entire house-> Customer replaces this on Monday, device switched off | Device flushed through, start phase awaited -> Error remains, in the service program. Fluid transport all programs performed-> div. Prog. with errors, on pump hose of the LP red discoloration, pump hose perforated->it is only one LP in stock in the entire house-> Customer replaces this on Monday, device switched off | Assisted Replacement | SP | SP | FT
AQT90 FLEX | 更换842-344修复 by andrew | 更换842-344修复 by andrew | SWO-156667 | No | Replacement 842-344 Repair by Andrew | Replacement 842-344 Repair by Andrew | Defect | HW | Aspiration Module | Nipple broken
ABL80 CO-OX | SE CAMBIA SENSOR CONNECTOR BOARD, SE COMPRUEBA FUNCIONAMIENTO CON CAL Y QC | SE CAMBIA SENSOR CONNECTOR BOARD, SE COMPRUEBA FUNCIONAMIENTO CON CAL Y QC | D1DF00FCC80B4DED | Yes | SENSOR CONNECTOR BOARD IS CHANGED, CAL AND QC OPERATION CHECKS | SENSOR CONNECTOR BOARD IS CHANGED, CAL AND QC OPERATION CHECKS
ABL825 FLEX | Alle Monitor- Eingabefunktionen ok, keine Verzögerung und keinerlei Aussetzer. LED-Ränder alle sauber. Monitor-Sektor der Start Taste in anderen Menüs überprüft - Funktion einwandfrei. Laut Auftraggeber (Medizintechnik) wurde das Gerät bereits durch Anwender sowie durch Techniker neu gestartet aber keine dementsprechende Einträge in der Datei Aktivitäten gefunden (ABL wurde NICHT neu gestartet). Gerät aus- und eingeschaltet, Starttaste funktioniert wieder. | Alle Monitor- Eingabefunktionen ok, keine Verzögerung und keinerlei Aussetzer. LED-Ränder alle sauber. Monitor-Sektor der Start Taste in anderen Menüs überprüft - Funktion einwandfrei. Laut Auftraggeber (Medizintechnik) wurde das Gerät bereits durch Anwender sowie durch Techniker neu gestartet aber keine dementsprechende Einträge in der Datei Aktivitäten gefunden (ABL wurde NICHT neu gestartet). Gerät aus- und eingeschaltet, Starttaste funktioniert wieder. | SWO-156982 | No | All monitor input functions ok, no delay and no dropouts. LED edges all clean. Monitor sector of start button checked in other menus - function flawless. According to the client (medical technology), the device has already been restarted by users as well as by technicians, but no corresponding entries in the Activities file have been found (ABL has NOT been restarted). Device off and on, start button works again. | All monitor input functions ok, no delay and no dropouts. LED edges all clean. Monitor sector of start button checked in other menus - function flawless. According to the client (medical technology), the device has already been restarted by users as well as by technicians, but no corresponding entries in the Activities file have been found (ABL has NOT been restarted). Device off and on, start button works again. | Defect | SW | DMS_PC | Off/On
ABL825 FLEX | Error in intermittent but came back again Saw small air bubble under pH electode | Replaced inlet probe, inlet tube and pH BG lower liquid sensor again Replaced inlet nib, associated liquid sensor and pH electrode-new Replaced inlet nib and liquid sensor | D1DF2369C83F56D0 | No | Error in intermittent but came back again Saw small air bubble under pH electode | Replaced inlet probe, inlet tube and pH BG lower liquid sensor again Replaced inlet nib, associated liquid sensor and pH electrode-new Replaced inlet nib and liquid sensor | Assisted Replacement | Electrode | Electrode | pH
ABL90 FLEX | LP Wechsel | LP Wechsel | SWO-157093 | No | LP Change | LP Change | User error | Sample expired | Sample expired | Sample expired
ABL9 | Gerät überprüft, auf der Windowsebene Uhrzeit und Datum korrigiert, Reinstallation nicht möglich, da Verbrauchsmaterial nicht mehr akzeptiert wird, Datum und Uhrzeit zurückgestellt, QC Bereiche eingegeben | Gerät überprüft, auf der Windowsebene Uhrzeit und Datum korrigiert, Reinstallation nicht möglich, da Verbrauchsmaterial nicht mehr akzeptiert wird, Datum und Uhrzeit zurückgestellt, QC Bereiche eingegeben | SWO-157230 | No | Device checked, time and date corrected at Windows level, reinstallation not possible because consumables are no longer accepted, date and time deferred, QC ranges entered | Device checked, time and date corrected at Windows level, reinstallation not possible because consumables are no longer accepted, date and time deferred, QC ranges entered
ABL90 FLEX | アンプルダイレクトでの吸引で問題なく実施 | アンプルダイレクトでの吸引で問題なく実施 | SWO-158049 | No | Suction with ampoule direct without any problems | Suction with ampoule direct without any problems | User error | Wrong QC Adapter | Wrong QC Adapter | Wrong QC Adapter
ABL90 |  | B | D1E2D74B2A077F48 | No |  |  | Defect | HW | Inlet module | Inlet Gasket
ABL90 |  | B | D1E2A4432A267409 | Yes |  |  | Re classified RMED | Double entry
ABL90 FLEX | Replaced split Oxi pinch valve tube. Replaced faulty hemolyser. | Replaced split Oxi pinch valve tube. Replaced faulty hemolyser. | SWO-157241 | No | Replaced split Oxi pinch valve tube. Replaced faulty hemolyser. | Replaced split Oxi pinch valve tube. Replaced faulty hemolyser. | Defect | HW | Oxi module | Hemolyzer
ABL90 | Se acude, con ayuda de PBM, se chequean todas las conexiones y servicios de Aqure. Limpieza de colas de salida. Equipos empiezan a transmitir OK | se chequean todas las conexiones y servicios de Aqure. Limpieza de colas de salida. | D1E302062A16A2E2 | No | All Aqure connections and services are checked with the help of PBM. Cleaning of outbound queues. Teams start broadcasting OK | all Aqure connections and services are checked. Cleaning of outbound queues. | Setup | SW | His/lis | His/lis
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-160765 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | Leaking
AQT90 FLEX | 更换处理包，测试卡，清理内部灰尘后修复 by jesson | 更换处理包，测试卡，清理内部灰尘后修复 by jesson | SWO-157396 | No | Replace the processing pack, test the card, clean the internal dust and repair by jesson | Replace the processing pack, test the card, clean the internal dust and repair by jesson | Cleaned | HW | Storage Module | Sensor, storage wheel | 5689
ABL825 FLEX | kampós szelep cseréje, teljes OXI teszt, esedékes membrán cserék és tHb kalibráció elvégzése, QC mérés | B | SWO-157155 | No | Toggle valve replacement, full OXY test, due membrane substitutions and tHb calibration, QC measurement |  | Defect | HW | Oxi module | Pinch valve
ABL815 FLEX | OXI Modul 902-833R1966N010 installiert, Leihmodul abgezogen. Lac. QC Fehler beseitigt, fehlenden Dichtring ersetzt, Membran erneuert. OXI Complete Test und Service tHb Kal. i.O. | OXI Modul 902-833R1966N010 installiert, Leihmodul abgezogen. Lac. QC Fehler beseitigt, fehlenden Dichtring ersetzt, Membran erneuert. OXI Complete Test und Service tHb Kal. i.O. | SWO-157237 | No | OXI module 902-833R1966N010 installed, loan module deducted. Lac. QC error eliminated, missing sealing ring replaced, membrane renewed. OXI Complete Test and Service tHb Kal. Ok. | OXI module 902-833R1966N010 installed, loan module deducted. Lac. QC error eliminated, missing sealing ring replaced, membrane renewed. OXI Complete Test and Service tHb Kal. Ok. | Revisit required | HW | Oxi module | Oxi module
AQT90 FLEX | entfernt und ersetzt, Gerät geöffnet und abgebrochene Teile und Proberöhrchen entfernt, Patmessungen durchgeführt, i.O. | entfernt und ersetzt, Gerät geöffnet und abgebrochene Teile und Proberöhrchen entfernt, Patmessungen durchgeführt, i.O. | SWO-157355 | No | removed and replaced, device opened and broken parts and sample tubes removed, pat measurements performed, i.O. | removed and replaced, device opened and broken parts and sample tubes removed, pat measurements performed, i.O. | Defect | HW | Aspiration Module | Tube holder
AQT90 FLEX | Limpieza del la unidad de Hct con hipoclorito segun tip. Comprobacion con qc y varias muestras. | B | D1E314862A038D61 | No | Cleaning the Hct unit with hypochlorite according to tip. Check with qc and multiple samples. |  | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1258
ABL80 FLEX BASIC |  | B | SWO-157393 | No
ABL90 FLEX | Thb kal durchgeführt | Thb kal durchgeführt | SWO-157385 | Yes | Thb kal performed | Thb kal performed | Re classified RMED | ECKHI Email 24.jan-2117 | Oxi module | tHb Cal
ABL90 FLEX | Replaced black pinch valve tubing with new | Replaced black pinch valve tubing with new | SWO-157489 | No | Replaced black pinch valve tubing with new | Replaced black pinch valve tubing with new | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX | Jeg har skiftet inletpakningsholder. Der er skiftet væskepakke. Alle test under Fluid transport i service er kørt. Der er kørt kal. og QCer. Der er kørt tHb kal. Alt er OK igen. | Jeg har skiftet inletpakningsholder. Der er skiftet væskepakke. Alle test under Fluid transport i service er kørt. Der er kørt kal. og QCer. Der er kørt tHb kal. Alt er OK igen. | SWO-157515 | No | I've changed the inlet pack holder. A liquid package has been switched. All tests during FLUID transport in service are run. There's a run of the Kal. and QCer. A tHb Kal has been run. Everything is OK again. | I've changed the inlet pack holder. A liquid package has been switched. All tests during FLUID transport in service are run. There's a run of the Kal. and QCer. A tHb Kal has been run. Everything is OK again. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Fliesswege gespült ,Dichtung überprüft und erneut eingesetzt. SP erneut eingesetzt Test wiederholt, Aktivitäten kontrolliert. Kalibrierungen und QC durchgeführt. | Fliesswege gespült ,Dichtung überprüft und erneut eingesetzt. SP erneut eingesetzt Test wiederholt, Aktivitäten kontrolliert. Kalibrierungen und QC durchgeführt. | SWO-157494 | No | Flow paths rinsed, seal checked and re-inserted. SP re-used test repeatedly, controlling activities. calibrations and QC. | Flow paths rinsed, seal checked and re-inserted. SP re-used test repeatedly, controlling activities. calibrations and QC. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | Set up as per request, adult and baby report layouts available. | B | SWO-157800 | No | Set up as per request, adult and baby report layouts available. |  | Setup | SW | Patient Layout | Patient Layout
ABL90 FLEX PLUS | Installed new auto inlet motor assembly as part of tech line case to see if this resolves the issue. Customer will monitor the unit. Part 903-344 not on CRM. | Installed new auto inlet motor assembly as part of tech line case to see if this resolves the issue. Customer will monitor the unit. Part 903-344 not on CRM. | SWO-157452 | No | Installed new auto inlet motor assembly as part of tech line case to see if this resolves the issue. Customer will monitor the unit. Part 903-344 not on CRM. | Installed new auto inlet motor assembly as part of tech line case to see if this resolves the issue. Customer will monitor the unit. Part 903-344 not on CRM. | Defect | HW | Auto Inlet module | Motor assembly | Grub
AQT90 FLEX | 5784故障修复后，机器开机后出现错误故障1020，检查机器状态，机器的孵育盘和环境温度没有显示，更换样机上的636-652后修复 by jesson | 5784故障修复后，机器开机后出现错误故障1020，检查机器状态，机器的孵育盘和环境温度没有显示，更换样机上的636-652后修复 by jesson | SWO-157344 | No | 5784 fault repair, the machine after the start-up error failure 1020, check the status of the machine, the machine incubation disk and ambient temperature is not shown, replace the prototype 636-652 after repair by jesson | 5784 fault repair, the machine after the start-up error failure 1020, check the status of the machine, the machine incubation disk and ambient temperature is not shown, replace the prototype 636-652 after repair by jesson | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken)
AQT90 FLEX | 自行更换备用光单（R0358N008）修复 by andrew | 自行更换备用光单（R0358N008）修复 by andrew | SWO-157336 | No | Replace your own backup light sheet (R0358N008) to repair by Andrew | Replace your own backup light sheet (R0358N008) to repair by Andrew | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 自行更换备用光单修复 | 自行更换备用光单（R0358N008）修复 by andrew | SWO-157345 | Yes | Replace your own backup light sheet repair | Replace your own backup light sheet (R0358N008) to repair by Andrew | Re classified RMED | Double entry | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364后修复by jesson | 更换903-364后修复by jesson | SWO-157314 | No | Repair by jesson after replacement 903-364 | Repair by jesson after replacement 903-364 | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 使用备用机备件903-364修复 by jesson | 使用备用机备件903-364修复 by jesson | SWO-157365 | No | Repair by jesson with spare parts 903-364 | Repair by jesson with spare parts 903-364 | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | Inspected flex Q module and found that it was slightly out of alignment possibly causing the error. | Performed the alignment procedure in the service program. | D1E3ED6BA0293B35 | Yes | Inspected flex Q module and found that it was slightly out of alignment possibly causing the error. | Performed the alignment procedure in the service program.
ABL837 FLEX |  | B | D1E4343DA0343A65 | No |  |  | Defect | HW | Inlet module | Probe bend
ABL800 BASIC | Replaced the oxi-module sent by Prakatheesh Sir, restarted the machine done the Complete Oximetry Test, after a successful 2 point calibration run 3 samples, results are OK, machine is working fine. Kept the machine under observation for the next 2 days. | Replaced the oxi-module sent by Prakatheesh Sir, restarted the machine done the Complete Oximetry Test, after a successful 2 point calibration run 3 samples, results are OK, machine is working fine. Kept the machine under observation for the next 2 days. | SWO-158256 | No | Replaced the oxi-module sent by Prakatheesh Sir, restarted the machine done the Complete Oximetry Test, after a successful 2 point calibration run 3 samples, results are OK, machine is working fine. Kept the machine under observation for the next 2 days. | Replaced the oxi-module sent by Prakatheesh Sir, restarted the machine done the Complete Oximetry Test, after a successful 2 point calibration run 3 samples, results are OK, machine is working fine. Kept the machine under observation for the next 2 days. | Revisit required | HW | Oxi module | Oxi module
ABL815 FLEX | Proteinentfernungsprogramm durchgeführt, LS-Justage, fehlende untere Kappe der PO2-Elektrode ersetzt, Messkammer gereinigt, Kal, QC | Ventile von den Messmodulen kontrolliert -> beide Air Ventile hingen unten fest und Membran vom pH BG Modul durchstoßen beide Ventile ausgetauscht LS Justage + Pumpenkal | SWO-157670 | No | Protein removal program performed, LS adjustment, missing lower cap of PO2 electrode replaced, measuring chamber cleaned, Kal, QC | Valves controlled by the measuring modules -> both air valves hung tightly at the bottom and penetrate membrane from the pH BG module both valves replaced LS Adjustment + Pump Scale | Defect | HW | pH/BG Module | Valve, air | Membrane
ABL825 FLEX | Standzeit der Lac Membranen aktuell etwa 3 bis 7 Tage. Elektrode sowie Kontaktkappe auf Korrosion untersucht und Reinigungsverfahren mit Hypochlorid und Aqua Dest durchgeführt, beide Dichtringe ersetzt, membraniert, initialisiert, Kalibrationen (Empf. ca. 500pA mM) und QC Messung, OK | Nachtrag SWO Dichtungen | SWO-157530 | No | Life of the lac membranes currently about 3 to 7 days. Electrode and contact cap tested for corrosion and cleaning procedures performed with hypochloride and aqua dest, both sealing rings replaced, membraneized, initialized, calibrations (damp approx. 500pA mM) and QC measurement, OK | Supplement SWO Seals | Cleaned | Electrode | Electrode | Lac | Cleaned e-tip (378)
ABL825 FLEX | Nachtrag SWO Dichtungen | Nachtrag SWO Dichtungen | SWO-157548 | Yes | Supplement SWO Seals | Supplement SWO Seals | Re classified RMED | Double entry
ABL825 FLEX | Upon arrival, found unit setup using Dynamic Parameter feature which is not needed. | Setup all patine parameter profiles using all parameters wanted in each profile to eliminate need to choose parameters. | D1E41109A03738C3 | No | Upon arrival, found unit setup using Dynamic Parameter feature which is not needed. | Setup all patine parameter profiles using all parameters wanted in each profile to eliminate need to choose parameters. | Setup | Patient Layout | Patient Layout | Patient Layout
ABL825 FLEX | checked data logs for errors, checked flow of solutiuons, checked pumps and tubing, waste tubing disconnected from connector | removed and replaced waste connector tubing to waste bottle, | D1E48093A0218119 | No | checked data logs for errors, checked flow of solutiuons, checked pumps and tubing, waste tubing disconnected from connector | removed and replaced waste connector tubing to waste bottle, | Defect | HW | Basic Unit | Waste tubing
ABL80 FLEX BASIC | Solution Pack Not Inserted properly. Waist Leak Also Found. There is a manufacturing fault in Solution Pack. Door closed properly. and Puted a paper piece to retail. | Solution Pack Not Inserted properly. Waist Leak Also Found. There is a manufacturing fault in Solution Pack. Door closed properly. and Puted a paper piece to retail. | SWO-158409 | No | Solution Pack Not Inserted properly. Waist Leak Also Found. There is a manufacturing fault in Solution Pack. Door closed properly. and Puted a paper piece to retail. | Solution Pack Not Inserted properly. Waist Leak Also Found. There is a manufacturing fault in Solution Pack. Door closed properly. and Puted a paper piece to retail.
ABL90 | Troubleshot and found that after the customer had powered off the analyzer, the static IP settings and defaulted back to DHCP. | Corrected the settings and performed the EWFMGR Commit. Then I powered off the analyzer and powered it back on to make sure the settings remained. | D1E5CD44166271EF | No | Troubleshot and found that after the customer had powered off the analyzer, the static IP settings and defaulted back to DHCP. | Corrected the settings and performed the EWFMGR Commit. Then I powered off the analyzer and powered it back on to make sure the settings remained. | Setup | SW | His/lis | His/lis
ABL90 | When analyzer is unplugged from mains, it switches to battery but when plugged back to mains, the power supply will not switch back to mains. | Replaced power supply. | D1E528C8163E113A | No | When analyzer is unplugged from mains, it switches to battery but when plugged back to mains, the power supply will not switch back to mains. | Replaced power supply. | Defect | HW | Power Supply | Power Supply
ABL90 | When analyzer is unplugged from mains, it switches to battery but when plugged back to mains, the power supply will not switch back to mains. | Replaced power supply. | D1E52A691657113E | No | When analyzer is unplugged from mains, it switches to battery but when plugged back to mains, the power supply will not switch back to mains. | Replaced power supply. | Defect | HW | Power Supply | Power Supply
ABL90 | Checked battery, will not function. | Replaced battery. | D1E52B1016281141 | No | Checked battery, will not function. | Replaced battery. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | Kein Fehler feststellbar, kurzeinweisung, Änderungswünsche Konfiguration aufgenommen und mit Labor besprochen, Pico50 Probennehmer könnten das Problem beheben. | Kein Fehler feststellbar, kurzeinweisung, Änderungswünsche Konfiguration aufgenommen und mit Labor besprochen, Pico50 Probennehmer könnten das Problem beheben. | SWO-157881 | No | No error detectable, short instruction, change requests configuration included and discussed with laboratory, Pico50 samplers could fix the problem. | No error detectable, short instruction, change requests configuration included and discussed with laboratory, Pico50 samplers could fix the problem. | No fault found | No fault found | No fault found | Wetability
ABL90 FLEX | Inlet Probe ersetzt da verbogen, Inlet Flow getestet. Kein weiteres, technisches Problem erkennbar. Funktionskontrolle, Probenmessungen. | Inlet Probe ersetzt da verbogen, Inlet Flow getestet. Kein weiteres, technisches Problem erkennbar. Funktionskontrolle, Probenmessungen. | SWO-157880 | No | Inlet sample replaces there bent, Inlet Flow tested. No further technical problem recognizable. Functional control, sample measurements. | Inlet sample replaces there bent, Inlet Flow tested. No further technical problem recognizable. Functional control, sample measurements. | Replaced | HW | Inlet module | Probe bend
ABL90 FLEX PLUS |  | B | D1E53520161111FC | No |  |  | Adjustment | HW | Auto Inlet module | Inlet drive gear | Grub
ABL90 FLEX PLUS | Einlass lässt sich elektrisch nicht verstellen. Gerät demontiert und Einlass mechanisch in die richtige Position gebracht. Neustart durchgeführt und Einlass im Service kalibriert. Wegen unterbrochener Spülung, Sensorcassette erneuert. Trotz diverser Kal und QC ist Glu und Lac außer Bereich. Anschließend musste der Lösungspack auch erneuert werden. Diverse Kal und QC erneut durchgeführt. Uhrzeit Gestellt , Date Time Einstellungen vorgenommen und abgespeichert. | Einlass lässt sich elektrisch nicht verstellen. Gerät demontiert und Einlass mechanisch in die richtige Position gebracht. Neustart durchgeführt und Einlass im Service kalibriert. Wegen unterbrochener Spülung, Sensorcassette erneuert. Trotz diverser Kal und QC ist Glu und Lac außer Bereich. Anschließend musste der Lösungspack auch erneuert werden. Diverse Kal und QC erneut durchgeführt. Uhrzeit Gestellt , Date Time Einstellungen vorgenommen und abgespeichert. | SWO-158337 | No | Inlet cannot be electrically adjusted. device and mechanically placed the inlet in the correct position. restart and inlet calibrated in the service. Due to interrupted flushing, sensor cassette renewed. Despite various Kal and QC, Glu and Lac is out of range. Subsequently, the solution pack had to be renewed. Several Kal and QC performed again. Time Set, Date Time settings made and saved. | Inlet cannot be electrically adjusted. device and mechanically placed the inlet in the correct position. restart and inlet calibrated in the service. Due to interrupted flushing, sensor cassette renewed. Despite various Kal and QC, Glu and Lac is out of range. Subsequently, the solution pack had to be renewed. Several Kal and QC performed again. Time Set, Date Time settings made and saved. | Adjustment | HW | Auto Inlet module | Inlet drive gear
ABL90 FLEX | 更换测试卡修复 | 更换测试卡修复 | SWO-157661 | No | Replacement test card repair | Replacement test card repair | Assisted Replacement | SC | SC | Reference
ABL90 FLEX | 更换测试卡修复 | 更换测试卡修复 | SWO-157663 | No | Replacement test card repair | Replacement test card repair | Adjustment | SC | SC | Cal
AQT90 FLEX | 更换636-550修复 by jesson | 更换636-550修复 by jesson | SWO-157359 | No | Replace 636-550 repair by jesson | Replace 636-550 repair by jesson | Defect | HW | Process Module | Process Unit sensor | 5257
AQT90 FLEX | 更换配件886-625修复 by jesson | 更换配件886-625修复 by jesson | SWO-157644 | No | Replacement accessory 886-625 repair by jesson | Replacement accessory 886-625 repair by jesson | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken)
AQT90 FLEX | 清洁处理盘灰尘无效，更换903-364后修复 by jesson | 清洁处理盘灰尘无效，更换903-364后修复 by jesson | SWO-157353 | No | Cleaning the dust of the treatment disc is invalid, repair by jesson after replacing 903-364 | Cleaning the dust of the treatment disc is invalid, repair by jesson after replacing 903-364 | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | 更换903-224修复 by jesson | 更换903-224修复 by jesson | SWO-157303 | Yes | Replace 903-224 Fix by Jesson | Replace 903-224 Fix by Jesson | Re classified RMED | OOB
ABL800 FLEX | Replace the Crea consumable & recalibrate system. | Replace the Crea consumable & recalibrate system. | SWO-158476 | No | Replace the Crea consumable & recalibrate system. | Replace the Crea consumable & recalibrate system. | Assisted Replacement | Consumables | Solution bottle | Crea
ABL835 FLEX |  | B | D1E781B60261A2B5 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | Gateway Adresse korrigiert. Übertragung hergestellt, Testmessung mit Datensendung iO. | Gateway Adresse korrigiert. Übertragung hergestellt, Testmessung mit Datensendung iO. | SWO-157777 | No | Gateway address corrected. Transmission, test measurement with data transmission iO. | Gateway address corrected. Transmission, test measurement with data transmission iO. | Setup | SW | His/lis | His/lis
AQT90 FLEX | cleaned process wheel. cleaned optical unit. measured darkness value. run a test analyzer working fine. | cleaned process wheel. cleaned optical unit. measured darkness value. run a test analyzer working fine. | SWO-158465 | No | cleaned process wheel. cleaned optical unit. measured darkness value. run a test analyzer working fine. | cleaned process wheel. cleaned optical unit. measured darkness value. run a test analyzer working fine. | Cleaned | HW | Process Module | Process wheel
ABL80 FLEX | Done Hydration Procedure, Placed SC for Cooling in Fridge for 6 Hours, Try to Reinstalled SC But problem Remains, Need to replace SC with New One, PER Sent for Replacement. | Done Hydration Procedure, Placed SC for Cooling in Fridge for 6 Hours, Try to Reinstalled SC But problem Remains, Need to replace SC with New One, PER Sent for Replacement. | SWO-158488 | No | Done Hydration Procedure, Placed SC for Cooling in Fridge for 6 Hours, Try to Reinstalled SC But problem Remains, Need to replace SC with New One, PER Sent for Replacement. | Done Hydration Procedure, Placed SC for Cooling in Fridge for 6 Hours, Try to Reinstalled SC But problem Remains, Need to replace SC with New One, PER Sent for Replacement.
ABL80 FLEX BASIC | 更换902-972后修复 by jesson | 更换902-972后修复 by jesson | SWO-157720 | No | Repair by jesson after replacement 902-972 | Repair by jesson after replacement 902-972
ABL80 FLEX BASIC | 现场更换试剂包修复 | 现场更换试剂包修复 | SWO-158272 | No | On-site replacement reagent kit repair | On-site replacement reagent kit repair
ABL90 FLEX | Printer Unit csere. Kalibrálás, QC mérés. | B | SWO-158076 | No | Printer Unit replacement. Calibration, QC measurement. |  | Defect | HW | Printer module | Printer module
ABL90 FLEX | Einlass überarbeitet. Dichtung durch Kd. ersetzt. Einlassnadel deformiert. Nadel wieder gerichtet. Einlass getestet ok. Nach einigen Reinstallationen SP endlich ok. Kal 40 Min abgewartet. | ABL ohne LP neu gestartet - dabei fällt Sr. Heidi auf, das die Probennadel verbogen ist. Da sie keine mehr vor Ort hat, hat sie die Nadel gerichtet. Das Startverfahren wird wieder abgebrochen. Sr. Heidi schaut auf den anderen Stationen, ob sie dort eine neue Probennadel bekommt und tauscht die dann aus. Bei weiteren Fragen wird sie sich wieder melden. | SWO-158185 | No | Entrance revised. Seal replaced by Kd. inlet needle deformed. Needle again directed. Inlet tested ok. After some reinstallations SP finally ok. Kal 40 min waited. | ABL without LP restarted - Sr. Heidi notices that the sample needle is bent. Since she no longer has one on site, she has the needle. The startup procedure is aborted again. Sr. Heidi looks at the other stations to see if she gets a new rehearsal needle there and then exchanges them. She will contact me again if she has any further questions. | Replaced | HW | Inlet module | Probe bend
ABL90 |  | B | D1E6BF698C50F79F | Yes |  |  | Re classified RMED | Remote Support
ABL90 FLEX | Konfiguration gesichert, CF-Karte erneuert, Windows-Software und Analysatorsoftware V3.1MR8 installiert, LP und SC wieder installiert, gesicherte Konfiguration wieder installiert, | Konfiguration gesichert, CF-Karte erneuert, Windows-Software und Analysatorsoftware V3.1MR8 installiert, LP und SC wieder installiert, gesicherte Konfiguration wieder installiert, | SWO-158104 | No | Configuration backed up, CF card renewed, Windows software and analyzer software Installed V3.1MR8, LP and SC reinstalled, Backed up configuration reinstalled, | Configuration backed up, CF card renewed, Windows software and analyzer software Installed V3.1MR8, LP and SC reinstalled, Backed up configuration reinstalled, | Defect | HW | DMS_PC | CF Card
ABL90 FLEX | Fehlersuche, Fließwege gespült, div. Programme im Service durchgef., LP neu initialisiert, Neustart, Startphase abgewartet, Kal. + QC i.O. | Fehlersuche; Fließwege gespült; div. Programme im Service durchgef.; LP neu initialisiert; Neustart; Startphase abgewartet; Kal. + QC i.O. | SWO-158087 | No | Troubleshooting; Flow paths rinsed; various programs in the service; LP reinitialized; Restart; start-up phase awaited; Kal. + QC in the first place | Troubleshooting; Flow paths rinsed; various programs in the service; LP reinitialized; Restart; start-up phase awaited; Kal. + QC in the first place | Assisted Replacement | SP | SP | FT
ABL90 FLEX PLUS |  | B | D1E7A90202291069 | No |  |  | No fault found | SC | SC | pCO2
ABL90 FLEX PLUS | Changed interval date to 2 months, then back to 3, which fixed issue. Reported back to Nick so that he can do the same to the other affected analysers. | Changed interval date to 2 months, then back to 3, which fixed issue. Reported back to Nick so that he can do the same to the other affected analysers. | SWO-158544 | No | Changed interval date to 2 months, then back to 3, which fixed issue. Reported back to Nick so that he can do the same to the other affected analysers. | Changed interval date to 2 months, then back to 3, which fixed issue. Reported back to Nick so that he can do the same to the other affected analysers. | No fault found | SW | Oxi module | tHb Cal
ABL90 FLEX PLUS |  | B | D1E6B4318C41C7A1 | No |  |  | Defect | HW | Oxi module | Hemolyzer
AQT90 FLEX | 定标卡批号发货错误，已补发，故此单关单处理。 | 定标卡批号发货错误，已补发，故此单关单处理。 | SWO-158019 | No | The bid card lot number shipping error, has been reissued, so this single clearance processing. | The bid card lot number shipping error, has been reissued, so this single clearance processing. | Assisted Replacement | Consumables | Test Kit | Cartridge | Myo
AQT90 FLEX |  | 定标卡批号发货错误，已补发，故此单关单处理。 | SWO-158047 | Yes |  | The bid card lot number shipping error, has been reissued, so this single clearance processing. | Re classified RMED | Double entry
ABL835 FLEX | Comprobar estado de los tubos, bombas y válvulas, cambiar tubos de bombas, calibrar y probar con controles. | Cambiar los tubos de las bombas. | D1E6BCF18C05CA94 | No | Check the status of tubes, pumps and valves, change pump tubes, calibrate and test with controls. | Change the pump tubes. | Defect | HW | pH/BG Module | Pump tubes
ABL837 FLEX | Elektrodeconnector udskiftet, slutafprøvet. SEJ sender ny referenceelektrode. | Elektrodeconnector udskiftet, slutafprøvet. SEJ sender ny referenceelektrode. | SWO-158140 | No | Elektrodeconnector replaced, final tested. COOL sends new reference electrode. | Elektrodeconnector replaced, final tested. COOL sends new reference electrode. | Assisted Replacement | Electrode | Electrode | Reference
ABL815 FLEX | Ventile von den Messmodulen kontrolliert -> beide Air Ventile hingen unten fest und Membran vom pH BG Modul durchstoßen beide Ventile ausgetauscht LS Justage + Pumpenkal | Ventile von den Messmodulen kontrolliert -> beide Air Ventile hingen unten fest und Membran vom pH BG Modul durchstoßen beide Ventile ausgetauscht LS Justage + Pumpenkal | SWO-158071 | No | Valves controlled by the measuring modules -> both air valves hung tightly at the bottom and penetrate membrane from the pH BG module both valves replaced LS Adjustment + Pump Scale | Valves controlled by the measuring modules -> both air valves hung tightly at the bottom and penetrate membrane from the pH BG module both valves replaced LS Adjustment + Pump Scale | Defect | HW | pH/BG Module | Valve, air
ABL837 FLEX | Elektrodeconnector udskiftet, slutafprøvet. SEJ sender ny referenceelektrode. | Elektrodeconnector udskiftet, slutafprøvet. SEJ sender ny referenceelektrode. | SWO-158139 | Yes | Elektrodeconnector replaced, final tested. COOL sends new reference electrode. | Elektrodeconnector replaced, final tested. COOL sends new reference electrode. | Re classified RMED | Double entry
ABL805 FLEX | Regelmäßig QC Ausreißer nach unten bei pCO2 in Level drei. Erneute Anfrage an RDEU gestartet. Ggf. erneuter Datenextrakt und erneuter Tausch der Elektrode (zuletzt ersetzt am 26.08.2016). | Regelmäßig QC Ausreißer nach unten bei pCO2 in Level drei. Erneute Anfrage an RDEU gestartet. Ggf. erneuter Datenextrakt und erneuter Tausch der Elektrode (zuletzt ersetzt am 26.08.2016). | SWO-158149 | Yes | Regularly QC outliers down at pCO2 in level three. Re-enquiry to RDEU started. If necessary, re-data extract and re-exchange of the electrode (last replaced on 26.08.2016). | Regularly QC outliers down at pCO2 in level three. Re-enquiry to RDEU started. If necessary, re-data extract and re-exchange of the electrode (last replaced on 26.08.2016). | Re classified RMED | Courtesy / Follow Up
ABL825 FLEX | Looked in datalogs and found that status is reading 6.672, Sensitivity is reading 98.3 and Status 0 is reading 7.175 | Cleaned measuring chamber, looked at reference and PH electrodes. Reference electrode was lot 31-05, it is over 2 years old. Customer had new one. Replaced reference electrode and new reference membrane. | D1E68E118C37A1AA | No | Looked in datalogs and found that status is reading 6.672, Sensitivity is reading 98.3 and Status 0 is reading 7.175 | Cleaned measuring chamber, looked at reference and PH electrodes. Reference electrode was lot 31-05, it is over 2 years old. Customer had new one. Replaced reference electrode and new reference membrane. | User error | Electrode | Electrode | Reference | Expired
ABL800 BASIC | cleaned chamber. run protien removal. shaked co2 electrode. analyzer working fine. | cleaned chamber. run protien removal. shaked co2 electrode. analyzer working fine. | SWO-158463 | No | cleaned chamber. run protien removal. shaked co2 electrode. analyzer working fine. | cleaned chamber. run protien removal. shaked co2 electrode. analyzer working fine. | Cleaned | Fluidic pathway | Fluidic pathway | pCO2
ABL837 FLEX | Ny gas 1+2 monteret, slutafprøvet. | Ny gas 1+2 monteret, slutafprøvet. | SWO-158133 | No | New gas 1 + 2 mounted, final tested. | New gas 1 + 2 mounted, final tested. | User error | Consumables | Gas empty | Gas empty
ABL837 FLEX | Service tHb og alm. tHb kalibreret, slutafprøvet. | Service tHb og alm. tHb kalibreret, slutafprøvet. | SWO-158136 | No | Service tHb og alm. tHb kalibreret, slutafprøvet. | Service tHb og alm. tHb kalibreret, slutafprøvet. | Adjustment | HW | Oxi module | tHb cal
ABL837 FLEX | Service tHb og alm. tHb kalibreret, slutafprøvet. | Service tHb og alm. tHb kalibreret, slutafprøvet. | SWO-158135 | No | Service tHb og alm. tHb kalibreret, slutafprøvet. | Service tHb og alm. tHb kalibreret, slutafprøvet. | Adjustment | HW | Oxi module | tHb cal
ABL825 FLEX | Messkammer und Einlassdichtung gereinigt, Verbogene Einlasssonde gerichtet. Gerät mehrfach kalibriert. | Messkammer und Einlassdichtung gereinigt, Verbogene Einlasssonde gerichtet. Gerät mehrfach kalibriert. | SWO-158111 | No | Measuring chamber and inlet seal cleaned, bent inlet probe directed. calibrated several times. | Measuring chamber and inlet seal cleaned, bent inlet probe directed. calibrated several times. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Found inlet motor and friction detector not assembled correctly, loose | Tight and made sure in place properly Test in service inlet program--OK Problem returned on 1-9-2017 Replaced friction detector | D1E66B5D8C3DA3ED | No | Found inlet motor and friction detector not assembled correctly, loose | Tight and made sure in place properly Test in service inlet program--OK Problem returned on 1-9-2017 Replaced friction detector | Defect | HW | Inlet module | Friction detector unit | Screws not tightened
AQT90 FLEX |  | B | D1E6E9988C17B64B | No |  |  | Defect | HW | Process Module | PCB, Cup dryer
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-159794 | No | Repair after replacing test card | Repair after replacing test card
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-158325 | No | Replacement test card repair | Replacement test card repair
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-158315 | No | Replacement test card repair | Replacement test card repair
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-158284 | No | Replacement test card repair | Replacement test card repair
ABL90 | Confirmed reported issue by phone with Dr. Kim | Ordered replacement probe and will schedule on site visit 1 9 2017 | D1EB93A2644C04BE | Yes | Confirmed reported issue by phone with Dr. Kim | Ordered replacement probe and will schedule on site visit 1 9 2017 | Re classified RMED
ABL90 | Confirmed problem. Preesure test errors. | Replaced the inlet items. | D1E827DB0219B59A | No | Confirmed problem. Preesure test errors. | Replaced the inlet items. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Bluteingabedichtung wurde nicht korrekt angebracht. Dichtung ersetzt, alle test im Service Menü erfolgreich Neustart, Kalibration und QC-Kontrolle nicht erfogreich da Sensor länger nicht konditioniert wurde. Daher Sensor ersetzt. | Bluteingabedichtung wurde nicht korrekt angebracht. Dichtung ersetzt, alle test im Service Menü erfolgreich Neustart, Kalibration und QC-Kontrolle nicht erfogreich da Sensor länger nicht konditioniert wurde. Daher Sensor ersetzt. | SWO-158521 | No | Blood entry seal was not applied correctly. Seal replaced, all test in the service menu successfully restart, calibration and QC control not achieved because sensor has not been conditioned for longer. Therefore, sensor replaced. | Blood entry seal was not applied correctly. Seal replaced, all test in the service menu successfully restart, calibration and QC control not achieved because sensor has not been conditioned for longer. Therefore, sensor replaced. | User error | HW | Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D1E7EDD90222AB0B | Yes |  |  | Re classified RMED | Remote Support
ABL90 FLEX | 更换测试卡及试剂包后修复 | 更换测试卡及试剂包后修复 | SWO-160179 | No | Repair after replacing test cards and reagent kits | Repair after replacing test cards and reagent kits | Assisted Replacement | SC | SC | pCO2 | Sensitivity
ABL90 FLEX | 更换试剂包修复 | 更换试剂包修复 | SWO-158349 | No | Replacement reagent kit repair | Replacement reagent kit repair | Assisted Replacement | SP | SP | FT
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-158222 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | Chip
AQT90 FLEX | 更换备件636-550修复 | 更换备件636-550修复 | SWO-158218 | No | Replacement spare parts 636-550 repair | Replacement spare parts 636-550 repair | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken)
ABL800 BASIC | Check PO2 electrode and clean with brush and clean all the chamber,clean inlet gasket and waste line and run Protein removal and 2-point calibration which is Ok.Run QC which is OK and PO2 status is also OK and problem has been resolved. | Check PO2 electrode and clean with brush and clean all the chamber,clean inlet gasket and waste line and run Protein removal and 2-point calibration which is Ok.Run QC which is OK and PO2 status is also OK and problem has been resolved. | SWO-158470 | No | Check PO2 electrode and clean with brush and clean all the chamber,clean inlet gasket and waste line and run Protein removal and 2-point calibration which is Ok.Run QC which is OK and PO2 status is also OK and problem has been resolved. | Check PO2 electrode and clean with brush and clean all the chamber,clean inlet gasket and waste line and run Protein removal and 2-point calibration which is Ok.Run QC which is OK and PO2 status is also OK and problem has been resolved. | Cleaned | Fluidic pathway | Fluidic pathway | pO2
ABL800 BASIC | 更换910-275修复 by jesson | 更换910-275修复 by jesson | SWO-158278 | No | Replace 910-275 Repair by jesson | Replace 910-275 Repair by jesson | Defect | HW | Printer module | Printer module
ABL815 FLEX | Verstopften Belüftungsstutzen gespült Ref.- und Ca- Leihelektroden eingebaut Dichtring der pH-Elektrode ersetzt Probelauf | Verstopften Belüftungsstutzen gespült Ref.- und Ca- Leihelektroden eingebaut Dichtring der pH-Elektrode ersetzt Probelauf | SWO-158450 | No | Clogged ventilation nozzle flushed Ref.- and Ca- rental electrodes installed Sealing ring of the pH electrode replaced Rehearsal | Clogged ventilation nozzle flushed Ref.- and Ca- rental electrodes installed Sealing ring of the pH electrode replaced Rehearsal | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX | 0676 inlet sensitivity out of range errors in activity log. | Replaced friction detector. Made corrections to set up to default to arterial sample type. Remembraned lactate electrode. | D1E7BE47022BE682 | No | 0676 inlet sensitivity out of range errors in activity log. | Replaced friction detector. Made corrections to set up to default to arterial sample type. Remembraned lactate electrode. | Defect | HW | Inlet module | Friction detector unit
ABL815 FLEX | Fehlersuche: Keine Gerinnsel o.Ä. feststellbar, pCO2 Elektrode mit Cleaner gereinigt und vorerst mit der selben Hülse membraniert. Mehrere Kal, Empf. abfallend von 92% Richtung Untergrenze => Elektrode ersetzt, Initialisierung, Kal und QC Messung, OK | Fehlersuche: Keine Gerinnsel o.Ä. feststellbar; pCO2 Elektrode mit Cleaner gereinigt und vorerst mit der selben Hülse membraniert. Mehrere Kal, Empf. abfallend von 92% Richtung Untergrenze => Elektrode ersetzt, Initialisierung, Kal und QC Messung, OK | SWO-158383 | No | Troubleshooting: No clots or similar can be detected; pCO2 electrode cleaned with cleaner and membraned with the same sleeve for the time being. Multiple Kal, Empf. sloping from 92% towards the lower limit => electrode replaced, initialization, cal and QC measurement, OK | Troubleshooting: No clots or similar can be detected; pCO2 electrode cleaned with cleaner and membraned with the same sleeve for the time being. Multiple Kal, Empf. sloping from 92% towards the lower limit => electrode replaced, initialization, cal and QC measurement, OK | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | ヘモライザーユニット交換 OXIコンプリートテスト確認 コンタミ除去・エアライン・測定チャンバー清掃 Glu・Lac・Refメンブラン交換 CAL・tHbCAL・StHbCAL・QC測定データ確認 | ヘモライザーユニット交換 OXIコンプリートテスト確認 コンタミ除去・エアライン・測定チャンバー清掃 Glu・Lac・Refメンブラン交換 CAL・tHbCAL・StHbCAL・QC測定データ確認 | SWO-158453 | No | Hemolyzer unit replacement OXI Complete Test Confirmation Contamination removal, airline measurement chamber cleaning Glu, Lac, Ref, Membrane Replacement Cal, tHbCAL, StHbCAL, QC measurement data confirmation | Hemolyzer unit replacement OXI Complete Test Confirmation Contamination removal, airline measurement chamber cleaning Glu, Lac, Ref, Membrane Replacement Cal, tHbCAL, StHbCAL, QC measurement data confirmation | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX |  | B | D1E7BCA6024271C9 | Yes |  |  | Re classified RMED | Delivery
ABL815 FLEX | Inletmodul demontiert und geprüft, Abfallweg geprüft, Ersatzmodul installiert, Nullpunkt Kalibration durchgeführt, Funktionskontrolle ok | Inletmodul demontiert, geprüft und gereinigt, Funktionskontrolle ok | SWO-158392 | No | Inlet module dismantled and tested, waste path tested, replacement module installed, zero calibration performed, function control ok | Inlet module dismantled, tested and cleaned, function control ok | Defect | HW | Inlet module | Inlet module
AQURE | Skiftet V12 | Rettet ABL navn i filter fra HJ-KBA-FRH til HJ-FRH-KBA. | SWO-158296 | Yes | Changed V12 | Corrected ABL name in filter from HJ-KBA-FRH to HJ-FRH-KBA. | Re classified RMED | AQURE
ABL80 FLEX BASIC | removed casssette. reinstallled cassette and solution pack. cassette has installed successfully. | removed casssette. reinstallled cassette and solution pack. cassette has installed successfully. | SWO-158460 | No | removed casssette. reinstallled cassette and solution pack. cassette has installed successfully. | removed casssette. reinstallled cassette and solution pack. cassette has installed successfully.
ABL80 FLEX BASIC | remove the touch panel and cleaned it properly,then upgraded software but still the touch screen was not working. | replace the touch panel with the new one and it was working properly | SWO-160079 | No | remove the touch panel and cleaned it properly,then upgraded software but still the touch screen was not working. | replace the touch panel with the new one and it was working properly
ABL80 FLEX BASIC | replace the touch panel with the new one and it was working properly | replace the touch panel with the new one and it was working properly | SWO-160080 | Yes | replace the touch panel with the new one and it was working properly | replace the touch panel with the new one and it was working properly | Re classified RMED | Double entry | same as day before
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-159173 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-158702 | No | Replacement test card repair | Replacement test card repair
ABL90 FLEX | Visula Flow Check rendben, Tűtartó kimosása, Felszívótű ellenőrzése, Oldatpakk újrainstallálása, Szervíztesztek, Indítás, Kalibrációk, QC mérés | B | SWO-158694 | No | The Visula Flow Check is OK, you can wash the needle holder, check the suction needle, reinstall Solution Pack, service tests, start, calibration, QC measurement |  | Assisted Replacement | SP | SP | FT
ABL90 | Confirmed reported issue by visual inspection. | Worked with Dr. Kim, had him replace probe so he can train other staff members. | D1EB9528644F04C0 | No | Confirmed reported issue by visual inspection. | Worked with Dr. Kim, had him replace probe so he can train other staff members. | Replaced | HW | Inlet module | Probe bend
ABL90 FLEX | Analysator im Service überprüft (Adjust Manifold, Pressure Test, Clot detect remove, Pumpenkal., SP Check, LS Justage), Startphase, System und Funktion i.O. | Analysator im Service überprüft (Adjust Manifold, Pressure Test, Clot detect remove, Pumpenkal., SP Check, LS Justage), Startphase, System und Funktion i.O. | SWO-158678 | No | Analyzer checked in service (Adjust Manifold, Pressure Test, Clot detect remove, Pump Calca, SP Check, LS Adjustment), start phase, system and function i.O. | Analyzer checked in service (Adjust Manifold, Pressure Test, Clot detect remove, Pump Calca, SP Check, LS Adjustment), start phase, system and function i.O. | Adjustment | SW | Service pgm | Manifold
ABL90 FLEX PLUS | Sensorkassette und Lösungspack erneuert. Defekte SC und LP wird reklamiert. | Sensorkassette und Lösungspack erneuert. Defekte SC und LP wird reklamiert. | SWO-159045 | No | Sensor cassette and solution pack. Defective SC and LP are complained about. | Sensor cassette and solution pack. Defective SC and LP are complained about. | Assisted Replacement | SC | SC | Glu/Lac
ABL90 FLEX | 更换测试卡修复 | 更换测试卡修复 | SWO-159801 | No | Replacement test card repair | Replacement test card repair | Assisted Replacement | SC | SC | Reference
ABL90 FLEX |  | B | SWO-158594 | No |  |  | Defect | HW | Inlet module | Inlet module
ABL90 FLEX | 更换测试卡修复 | 更换测试卡修复 | SWO-158561 | No | Replacement test card repair | Replacement test card repair | Assisted Replacement | SC | SC | pCO2 | Calibration
AQT90 FLEX |  | B | SWO-158602 | No |  |  | Defect | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX | 更换备件944-394修复(by tim) | 更换备件944-394修复(by tim) | SWO-158549 | No | Replacement spare parts 944-394 repair (by tim) | Replacement spare parts 944-394 repair (by tim) | Assisted Replacement | Consumables | Test Kit | Cartridge | PCT CAL
AQT90 FLEX | 更换备件944-394修复 | 更换备件944-394修复(by tim) | SWO-158550 | Yes | Replacement parts 944-394 repair | Replacement spare parts 944-394 repair (by tim) | Re classified RMED | Double entry
AQT90 FLEX |  | B | SWO-158626 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit
ABL825 FLEX | Fehlersuche Wartung Ventilblock BG und EL div. Pumpenschläuche erneuert, Rollenkränze überprüft, LS Justage, Pumpenkal ,Proteinentfernung tHb kal , Lac membraniert | Fehlersuche Wartung Ventilblock BG und EL div. Pumpenschläuche erneuert, Rollenkränze überprüft, LS Justage, Pumpenkal ,Proteinentfernung tHb kal , Lac membraniert | SWO-158777 | No | Troubleshooting Maintenance valve block BG and EL div. Pump hoses renewed, roller wreaths checked, LS Adjustment, Pump Scale ,Protein Removal tHb kal , Lac membraneed | Troubleshooting Maintenance valve block BG and EL div. Pump hoses renewed, roller wreaths checked, LS Adjustment, Pump Scale ,Protein Removal tHb kal , Lac membraneed | Defect | HW | pH/BG Module | Pump tubes
ABL825 FLEX |  | B | D1EB78D764581EC8 | Yes
ABL835 FLEX | Inletmodul geprüft, Probe Holder ausgetauscht, Funktionskontrolle ok | B | SWO-159054 | No | Inlet module tested, sample holder replaced, function control ok |  | Defect | HW | Inlet module | Probe arm
ABL825 FLEX | 更换O2电极修复 | 更换O2电极修复 | SWO-160260 | No | Replacement of O2 electrode repair | Replacement of O2 electrode repair | Assisted Replacement | Electrode | Electrode | pO2
ABL815 FLEX | Datensicherung, SW erneuert (XPE, Applikation), Rücksicherung Daten, Kal u. QC i.O. | B | SWO-158731 | No | Data backup, SW renewed (XPE, application), backup data, Kal & QC i.O. |  | Defect | HW | DMS_PC | Touch screen
ABL815 FLEX | Datensicherung, SW erneuert (XPE, Applikation), Rücksicherung Daten, Kal u. QC i.O. | CPU Unit und Monitor ersetzt, Monitor kalibriert, akz. Kal abgewartet, Kal u. QC i.O. MAC-Adresse: 00-04-5F-90-16-39 Servicedump gezogen | SWO-158730 | No | Data backup, SW renewed (XPE, application), backup data, Kal & QC i.O. | CPU unit and monitor replaced, monitor calibrated, etc. Kal waited, Kal & QC i.O. MAC address: 00-04-5F-90-16-39 Service dump pulled | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Reinigung der Messkammern und der Glasmembranfläche, System und Funktion i.O. | Reinigung der Messkammern und der Glasmembranfläche, System und Funktion i.O. | SWO-158806 | No | Cleaning of the measuring chambers and the glass membrane surface, system and function in the | Cleaning of the measuring chambers and the glass membrane surface, system and function in the | Assisted Replacement | Electrode | Electrode | Lac
ABL825 FLEX |  | B | SWO-158630 | Yes |  |  | Re classified RMED | OOB
ABL827 FLEX | Investigated to find 743 - Autochack Breaker error, 247 - Spectrometer A D converter failed ,269 - Hemolyzer phase regulation failed. | Replaced the breaker motor and hemolyzer. Performed complete OXI test, adjusted the lamp (3280), service THB calibration and calibrations and QC to test the autocheck function. Placed the analyzer in standby mode - Not live yet | D1EBBFAA643100B0 | No | Investigated to find 743 - Autochack Breaker error, 247 - Spectrometer A D converter failed ,269 - Hemolyzer phase regulation failed. | Replaced the breaker motor and hemolyzer. Performed complete OXI test, adjusted the lamp (3280), service THB calibration and calibrations and QC to test the autocheck function. Placed the analyzer in standby mode - Not live yet | Defect | HW | Auto Check Module | Breaker Motor
ABL80 CO-OX | Found that the sensor cassette probe was clotted and potentially the pumps. | Replaced the sensor cassette, pump heads and tubing for the analyzer. Advised Bobbie to use clot catchers. | D1ECEC4EDA2E92C0 | No | Found that the sensor cassette probe was clotted and potentially the pumps. | Replaced the sensor cassette, pump heads and tubing for the analyzer. Advised Bobbie to use clot catchers.
ABL90 FLEX | SC erneuert | SC erneuert | SWO-159017 | Yes | SC renewed | SC renewed | Re classified RMED | ECKHI Email 24.jan-2117 | SC | Reference
ABL90 FLEX | Analysator überprüft (kein Fehler feststellbar) - Erneuerung Lösungspack notwendig - Funktionsüberprüfung etc. ohne Beanstandung ! | Analysator überprüft (kein Fehler feststellbar) - Erneuerung Lösungspack notwendig - Funktionsüberprüfung etc. ohne Beanstandung ! | SWO-159050 | No | Analyzer checked (no error detectable) - renewal solution pack necessary - function check etc. without complaint ! | Analyzer checked (no error detectable) - renewal solution pack necessary - function check etc. without complaint ! | Assisted Replacement | SP | SP | pO2
ABL90 |  | B | D1ECDC6CDA19B9AA | No |  |  | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | LP erneuert - pO2 Fehler nicht mehr vorhanden. STK durchgeführt - ok. Nach Hochfahren und Kal Kalibrierfehler Glu vorhanden. SC Tausch und Reinitialisierung brachte keine Besserung. Glu nach Rücksprache mit Kunden deaktiviert. Kunde beobachtet, ob sich Glu über Nacht einläuft. Herr Dr. Herden gibt Mittwoch Rückmeldung an Techniker. Kalibration und QC's vorgenommen(alle Parameter außer Glu in Ordnung) Kunden SC und LP für Lager überlassen. Sollte Fehler noch vorhanden sein wird ein Reparaturersatzgerät geliefert, da E-Teil nicht im Bestand. | LP erneuert - pO2 Fehler nicht mehr vorhanden. STK durchgeführt - ok. Nach Hochfahren und Kal Kalibrierfehler Glu vorhanden. SC Tausch und Reinitialisierung brachte keine Besserung. Glu nach Rücksprache mit Kunden deaktiviert. Kunde beobachtet, ob sich Glu über Nacht einläuft. Herr Dr. Herden gibt Mittwoch Rückmeldung an Techniker. Kalibration und QC's vorgenommen(alle Parameter außer Glu in Ordnung) Kunden SC und LP für Lager überlassen. Sollte Fehler noch vorhanden sein wird ein Reparaturersatzgerät geliefert, da E-Teil nicht im Bestand. | SWO-159195 | No | LP renewed - pO2 error no longer exists. STK performed - ok. After ramp-up and kal calibration error Glu present. SC exchange and reinitialization brought no improvement. Glu deactivated after consultation with customers. customer observes if Glu enters overnight. Dr. Herden gives feedback to technicians on Wednesday. Calibration and QC's (all parameters except Glu in order) customers SC and LP for warehouses. If there are still errors, a repair replacement device is delivered, as E-part is not in stock. | LP renewed - pO2 error no longer exists. STK performed - ok. After ramp-up and kal calibration error Glu present. SC exchange and reinitialization brought no improvement. Glu deactivated after consultation with customers. customer observes if Glu enters overnight. Dr. Herden gives feedback to technicians on Wednesday. Calibration and QC's (all parameters except Glu in order) customers SC and LP for warehouses. If there are still errors, a repair replacement device is delivered, as E-part is not in stock. | Assisted Replacement | SP | SP | pO2
ABL90 FLEX | div. Test`s im Service durchgef., SC + LP gew., Startphase abgewartet, Kal. i.O., QC-Level A + C außer Bereich -> wegen zu kalter Lösung im LP (war im Kühlschrank), weitere Kal., QC i.O. | Fehlersuche; Fließwege gespült; div. Programme im Service durchgef.; LP neu initialisiert; Neustart; Startphase abgewartet; Kal. + QC i.O. | SWO-159044 | No | div. Test it in the service. SC + LP. start-up phase awaited; Kal. i.O.; QC-Level A + C out of range -> because of too cold solution in the LP (was in the refrigerator); more Kal.; QC in the first | Troubleshooting; Flow paths rinsed; various programs in the service; LP reinitialized; Restart; start-up phase awaited; Kal. + QC in the first place | Assisted Replacement | SP | SP | FT
ABL90 FLEX PLUS | Replaced inlet gasket and inlet probe. | Replaced inlet gasket and inlet probe. | SWO-159248 | No | Replaced inlet gasket and inlet probe. | Replaced inlet gasket and inlet probe. | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX PLUS |  | B | D1ECEA8CDA58DE3B | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX |  | B | SWO-158897 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX |  | B | SWO-158904 | No |  |  | Defect | HW | Printer module | Printer module
ABL825 FLEX | Beim Eintreffen nach Pumpenkal o.g. Fehler. ElMet Modul ausgebaut und beide Ventile sowie Pumpe überprüft, keine sichtbaren Fehler feststellbar. Beide Ventile ersetzt, weitere Arbeiten im Rahmen de HJW (s. anderer SWO). | Beim Eintreffen nach Pumpenkal o.g. Fehler. ElMet Modul ausgebaut und beide Ventile sowie Pumpe überprüft, keine sichtbaren Fehler feststellbar. Beide Ventile ersetzt, weitere Arbeiten im Rahmen de HJW (s. anderer SWO). | SWO-159078 | No | On arrival after pump scale above o.g. error. ElMet module removed and both valves and pump checked, no visible errors detectable. Both valves replaced, further work within the framework of HJW (see other SWO). | On arrival after pump scale above o.g. error. ElMet module removed and both valves and pump checked, no visible errors detectable. Both valves replaced, further work within the framework of HJW (see other SWO). | Defect | HW | El/Met Module | Pump tubes
ABL825 FLEX | Gerät demontiert und gereinigt, Elektroden geprüft, Funktionskontrolle ok | Frau D. lässt noch die Einlassdichtung ersetzen. Bleibt der Fehler bestehen wird sie sich noch einmal melden. | SWO-159107 | No | Device dismantled and cleaned, electrodes tested, function control ok | Ms. D. has the inlet seal replaced. If the error persists, she will report again. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Unplug and connect to another power outlet and still same issue. Primary fuses in power supply unit keeps blowing. | Replaced Power supply unit. Applied Windows7 software license label on analyzer. All tests are successful. | D1ECE863DA04FFD3 | No | Unplug and connect to another power outlet and still same issue. Primary fuses in power supply unit keeps blowing. | Replaced Power supply unit. Applied Windows7 software license label on analyzer. All tests are successful. | Defect | HW | Power supply | Power supply
AQT90 FLEX | Opened analyser and removed a large amount of dust from around the barcode scanner, removed tape from under solution pack and removed and refitted, no issues found. Cleaned around solution pack enclosure. Removed cartridges and tested scanner, all seems ok. Error rotate storage wheel error on data logs, operated all functions in service, no issues found. Customer ran sample with no issues. | Opened analyser and removed a large amount of dust from around the barcode scanner, removed tape from under solution pack and removed and refitted, no issues found. Cleaned around solution pack enclosure. Removed cartridges and tested scanner, all seems ok. Error rotate storage wheel error on data logs, operated all functions in service, no issues found. Customer ran sample with no issues. | SWO-159156 | No | Opened analyser and removed a large amount of dust from around the barcode scanner, removed tape from under solution pack and removed and refitted, no issues found. Cleaned around solution pack enclosure. Removed cartridges and tested scanner, all seems ok. Error rotate storage wheel error on data logs, operated all functions in service, no issues found. Customer ran sample with no issues. | Opened analyser and removed a large amount of dust from around the barcode scanner, removed tape from under solution pack and removed and refitted, no issues found. Cleaned around solution pack enclosure. Removed cartridges and tested scanner, all seems ok. Error rotate storage wheel error on data logs, operated all functions in service, no issues found. Customer ran sample with no issues. | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX BASIC |  | B | SWO-159225 | No
ABL80 FLEX BASIC |  | B | SWO-159186 | No
ABL80 FLEX BASIC | All data Base download and deleted. Analyzer Serviced and calibrated properly. Handed over in a fine working condition. | All data Base download and deleted. Analyzer Serviced and calibrated properly. Handed over in a fine working condition. | SWO-159548 | No | All data Base download and deleted. Analyzer Serviced and calibrated properly. Handed over in a fine working condition. | All data Base download and deleted. Analyzer Serviced and calibrated properly. Handed over in a fine working condition.
ABL90 FLEX | Erneuerung der Sonde, Einlassdichtungseinheit, Solution pack | B | SWO-159465 | No | Renewal of the probe, inlet seal unit, solution pack |  | Assisted Replacement | SP | SP | FT | 1176
ABL90 FLEX | SC wurde am 04.01.17 06:49Uhr ersetzt, Lac QC Level A fehlerfrei 04.01.17 23:04Uhr, Datalogs & SD erstellt und an TS geschickt | B | SWO-159383 | No | SC was replaced on 04.01.17 06:49, Lac QC Level A error-free 04.01.17 23:04, Datalogs & SD created and sent to TS |  | Assisted Replacement | SC | SC | Lac
ABL90 FLEX | Kunde wird die SC am 14.01.17 (fällig) ersetzten und sich ggf. melden wenn der Fehler weiter besteht, | B | SWO-159380 | No | Customer will replace the SC on 14.01.17 (due) and, if necessary, report it if the error persists, |  | Assisted Replacement | SC | Sc | Ca++ | LEH mail
ABL90 FLEX | SC wurde am 27.12.16 14:08 Uhr ersetzt, Glu QC Level A fehlerfrei am 27.12.16 20:51 Uhr, datalogs & SD erstellt und an TS geschickt | B | SWO-159442 | No | SC was replaced on 27.12.16 14:08, Glu QC Level A error-free on 27.12.16 20:51, datalogs & SD created and sent to TS |  | Assisted Replacement | SC | SC | Glu
ABL90 FLEX | SC wurde am 03.01.17 7:24Uhr ersetzt, Lac Level A fehlerfrei am 04.01.17 07:15Uhr, Datalogs und SD erstellt und an TS geschickt | B | SWO-159430 | No | SC was replaced on 03.01.17 7:24o, Lac Level A error-free on 04.01.17 07:15, Datalogs and SD created and sent to TS |  | Assisted Replacement | SC | SC | Lac
ABL90 FLEX PLUS |  | B | D1EE242D5045F2DA | No |  |  | Assisted Replacement | SC | SC | Reference
ABL90 FLEX PLUS | SC wurde am 11.1.17 07.30Uhr ersetzt, bei Anwesenheit Lac QC Werte in Level A & C fehlerhaft, Gerät mehrfach kalibriert, QC wiederholt, Lac QC Wert in Level A dann i.O., in Level C wert auf oberer Grenze ->info an Med-Technik, Datalogs & SD erstellt und an TS geschickt | B | SWO-159438 | No | SC was replaced on 11.1.17 07.30o, in presence Lac QC values in level A & C faulty, device repeatedly calibrated, QC repeated, Lac QC value in level A then i.O., in level C value on upper limit ->info created to Med-Technique, Datalogs & SD and sent to TS |  | Assisted Replacement | SC | SC | Lac
ABL90 FLEX PLUS | SC wurde am 10.01.17 13.20 ersetzt, Lac QC Level A fehlerfrei 10.01.17 19:39 Uhr, datalogs & SD erstellt und an TS geschickt | B | SWO-159435 | No | SC was replaced on 10.01.17 13.20, Lac QC Level A error-free 10.01.17 19:39, datalogs & SD created and sent to TS |  | Assisted Replacement | SC | SC | Lac
AQT90 FLEX | 更换备用机上的580-421后修复 | B | SWO-160155 | No | Repair after replacing 580-421 on the standby |  | Defect | HW | Storage Module | Stepper Motor, punch out
ABL835 FLEX |  | B | D1F9581E763D7799 | No |  |  | Assisted Replacement | Electrode | Electrode | pO2
ABL805 FLEX | Erneut einen Datenextrakt an RDEU Service gesendet und die Elektrode ersetzt. Gerät komplett überprüft, sichtkontrolliert und Fließverhalten getestet, keine Unregelmäßigkeiten feststellbar. Kalibrationen, Funktionstest, per Remote eine L3 Messung ausgelöst, OK | Erneut einen Datenextrakt an RDEU Service gesendet und die Elektrode ersetzt. Gerät komplett überprüft, sichtkontrolliert und Fließverhalten getestet, keine Unregelmäßigkeiten feststellbar. Kalibrationen, Funktionstest, per Remote eine L3 Messung ausgelöst, OK | SWO-159318 | No | Once again, a data extract is sent to RDEU Service and the electrode is replaced. Device completely checked, visually controlled and flow behavior tested, no irregularities detectable. Calibrations, function test, remotely triggered an L3 measurement, OK | Once again, a data extract is sent to RDEU Service and the electrode is replaced. Device completely checked, visually controlled and flow behavior tested, no irregularities detectable. Calibrations, function test, remotely triggered an L3 measurement, OK | Assisted Replacement | Electrode | Electrode | pCO2 | Grey zone
ABL815 FLEX | Neustarttaste nicht anwählbar beim Eintreffen, obwohl beide Klappen richtig sitzen. Klappen bewegt, ohne Änderung der Situation. Klappenposition im Servicemenü überprüft, alle Varianten werden korrekt angezeigt. Neustart nun möglich aber mit abnormalem Startbild => Kaltstart ausgelöst, Startverfahren und Initialisierung. Bitte morgen erneut beobachten! | Neustarttaste nicht anwählbar beim Eintreffen, obwohl beide Klappen richtig sitzen. Klappen bewegt, ohne Änderung der Situation. Klappenposition im Servicemenü überprüft, alle Varianten werden korrekt angezeigt. Neustart nun möglich aber mit abnormalem Startbild => Kaltstart ausgelöst, Startverfahren und Initialisierung. Bitte morgen erneut beobachten! | SWO-159426 | No | Restart button cannot be selected on arrival, even though both flaps are seated correctly. flaps moved without changing the situation. Flap position checked in the service menu, all variants are displayed correctly. Restart now possible but with abnormal start screen => Cold start triggered, start procedure and initialization. Please watch again tomorrow! | Restart button cannot be selected on arrival, even though both flaps are seated correctly. flaps moved without changing the situation. Flap position checked in the service menu, all variants are displayed correctly. Restart now possible but with abnormal start screen => Cold start triggered, start procedure and initialization. Please watch again tomorrow! | Unknown | HW | Inlet module | PCB, FLAP | Revisit required??
ABL825 FLEX |  | B | D1EE09A6504F3B4F | Yes |  |  | Re classified RMED | Installation
ABL825 FLEX | checked data logs for errors, checked flow of waste solutions, checked for clotts, | removed waste tubing cover and cleaned pooling blood caused by clotin waste connector, replaced waste tubing, performed tHb cal | D1EE9D295061FCA9 | No | checked data logs for errors, checked flow of waste solutions, checked for clotts, | removed waste tubing cover and cleaned pooling blood caused by clotin waste connector, replaced waste tubing, performed tHb cal | Defect | Hw | Basic Unit | Waste tubing
ABL825 FLEX | Nullpunktkalibrierung durchgeführt. Lac Elektrode neu membraniert. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL wieder Messbereit. | Nullpunktkalibrierung durchgeführt. Lac Elektrode neu membraniert. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. ABL wieder Messbereit. | SWO-159339 | No | zero-point calibration. Lac electrode re-membraneed. Multiple calibrations performed. QCs for all levels. ABL ready to measure again. | zero-point calibration. Lac electrode re-membraneed. Multiple calibrations performed. QCs for all levels. ABL ready to measure again. | Adjustment | HW | El/Met Module | Lac | Zero cal/adj
ABL825 FLEX |  | 借的备件910-275有质量问题，更换后仍出现卡纸，故重新借910-275一个，此单先关闭。(by TIM) | SWO-159204 | No |  | Borrowed spare parts 910-275 have quality problems, after replacement still appear paper jam, so re-borrow 910-275 one, this single closed first. (by TIM) | Defect | HW | Printer Module | Printer Module
AQT90 FLEX | reset the USB Cable, now working fine | reset the USB Cable, now working fine | SWO-159862 | No | reset the USB Cable, now working fine | reset the USB Cable, now working fine | Defect | SW | Analyzer Control Module | Off/On | Printer Module
ABL80 CO-OX | The analyzer works better than originally, but the pO2 still measures low on Level 1 external QC the first time it is measured. | Replaced the analog PC-board and the snesor connector board. The errors still occurred. Replaced the Spine Board and performed the Thourough Cleaning Program. | D1EF5916C6236C18 | No | The analyzer works better than originally, but the pO2 still measures low on Level 1 external QC the first time it is measured. | Replaced the analog PC-board and the snesor connector board. The errors still occurred. Replaced the Spine Board and performed the Thourough Cleaning Program.
ABL90 FLEX | Erneuerung des Solutions packs gegen konditionierte Ware lot:.XV-01 | Erneuerung des Solutions packs gegen konditionierte Ware lot:.XV-01 | SWO-159766 | Yes | Renewal of the Solution pack against conditioned goods lot:. XV-01 | Renewal of the Solution pack against conditioned goods lot:. XV-01 | Re classified RMED | ECKHI Email 24.jan-2117 | SP | Chip
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-159521 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Lac
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-161720 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Lac
ABL825 FLEX |  | B | D1EF50E6C60FF059 | No |  |  | Defect | HW | Basic Unit | Waste detector
ABL810 FLEX | E-Lyte Modul eingesetzt und Gerät entsprechend konfiguriert, Kal und QC durchgeführt | E-Lyte Modul eingesetzt und Gerät entsprechend konfiguriert, Kal und QC durchgeführt | SWO-159753 | No | E-Lyte module inserted and device configured accordingly, Kal and QC performed | E-Lyte module inserted and device configured accordingly, Kal and QC performed | Assisted Replacement | Electrode | Electrode | pCO2 | 8X / El/Met installed
AQT90 FLEX |  | B | D1EFA6D8C605CA40 | Yes |  |  | Re classified RMED | Collect data
ABL80 FLEX BASIC |  | B | SWO-159843 | No
ABL80 FLEX BASIC |  | B | SWO-159842 | No
ABL80 FLEX BASIC | Found SP not setting properly need to push in by external efforts then only smart chip connecting then again open. Solution pack issue as waist also coming out. SP not setting properly and not initialysing when try to reset and install. Need to replace SP. SP replaced and calibrated. analyser working fine. | Found SP not setting properly need to push in by external efforts then only smart chip connecting then again open. Solution pack issue as waist also coming out. SP not setting properly and not initialysing when try to reset and install. Need to replace SP. SP replaced and calibrated. analyser working fine. | SWO-160121 | No | Found SP not setting properly need to push in by external efforts then only smart chip connecting then again open. Solution pack issue as waist also coming out. SP not setting properly and not initialysing when try to reset and install. Need to replace SP. SP replaced and calibrated. analyser working fine. | Found SP not setting properly need to push in by external efforts then only smart chip connecting then again open. Solution pack issue as waist also coming out. SP not setting properly and not initialysing when try to reset and install. Need to replace SP. SP replaced and calibrated. analyser working fine.
ABL80 | Checked the data logs and verified the problem. Ordered and replaced the signal data board and sensor connector board. This did not solve the problem. | Replaced the signal data board and sensor connector board. Ordered and replaced the main distribution pcb and soution pack. | D1F0B2C83C0C5AB4 | No | Checked the data logs and verified the problem. Ordered and replaced the signal data board and sensor connector board. This did not solve the problem. | Replaced the signal data board and sensor connector board. Ordered and replaced the main distribution pcb and soution pack.
ABL90 FLEX | Nach Erneuerung des LP (lot YB-50) besteht ein Problem in der QC Lvl 1 und Lvl 3 mit pH und pCO2. LP (lot YB-50) ausgetauscht - genau das gleiche Fehlerbild. LP mit einer anderen lot Nummer aus dem Kreißsaal organisiert - Fehler behoben. Der Kunde hat eine Lieferung mit 4 LP bekommen alle mit der gleichen Lot. Die verbleibenden 2 LP nehme ich mit da diese vermutlich das gleiche Bild zeigen werden. Der Kunde bestellt neue LP. Kalibrierung und QC ok - Ampel steht auf grün. Datensicherung durchgeführt zur Auswertung der QC Probleme Anfang Januar (Ableitstrom erkannt 964 - Kunde tauschte das LP zur Problembehebung. | FSE-Einsatz gewünscht ( Anwender wird LP nochmals gegen ein neues tauschen) | SWO-160067 | No | After renewal of the LP (lot YB-50) there is a problem in the QC Lvl 1 and Lvl 3 with pH and pCO2. LP (lot YB-50) - exactly the same error image. LP organized with another lot number from the Kreißsaal - bug fixed. The customer has received a delivery with 4 LP all with the same lot. I take the remaining 2 Lp with me as they will probably show the same picture. The customer orders new LP. Calibration and QC ok - traffic light is on green. Data backup performed to evaluate QC problems at the beginning of January (Leakestrom detected 964 - customer exchanged the LP for troubleshooting. | FSE use desired (user will swap LP again for a new one) | Assisted Replacement | SP | SP | pCO2
ABL90 | Perform inspection and performance check. | Perform performance check. Run all servicve tests in service mode. | D1F0AE9B3C463A45 | Yes | Perform inspection and performance check. | Perform performance check. Run all servicve tests in service mode. | Re classified RMED | Courtesy
ABL90 FLEX | LP (XT06) eingebaut und (YB05) ausgebaut, System und Funktion i.O. | LP (XT06) eingebaut und (YB05) ausgebaut, System und Funktion i.O. | SWO-160053 | No | LP (XT06) built-in and (YB05) removed, system and function in the | LP (XT06) built-in and (YB05) removed, system and function in the | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX | LP (XT06) eingebaut und (YB05) ausgebaut, Clot detect remove, Spülung des Analysators, System und Funktion i.O. | LP (XT06) eingebaut und (YB05) ausgebaut, Clot detect remove, Spülung des Analysators, System und Funktion i.O. | SWO-160014 | No | LP (XT06) installed and (YB05) removed, Clot detect remove, flushing of the analyzer, system and function i.O. | LP (XT06) installed and (YB05) removed, Clot detect remove, flushing of the analyzer, system and function i.O. | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX | Analysator gespült und im Service überprüft (Pressure test, Clot detect remove), Startphase, System und Funktion i.O. | Analysator gespült und im Service überprüft (Pressure test, Clot detect remove), Startphase, System und Funktion i.O. | SWO-160052 | No | Analyzer rinsed and checked in the service (Pressure test, Clot detect remove), start phase, system and function i.O. | Analyzer rinsed and checked in the service (Pressure test, Clot detect remove), start phase, system and function i.O. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS |  | Er wird den ABL mal neu starten. Er möchte einen Service Einsatz und einen Wartungstermin. | SWO-160064 | Yes |  | He will restart the ABL. He wants a service assignment and a maintenance appointment. | Re classified RMED
ABL90 FLEX PLUS | Replaced PSU. Tested and checked. All working ok | Replaced PSU. Tested and checked. All working ok | SWO-159936 | No | Replaced PSU. Tested and checked. All working ok | Replaced PSU. Tested and checked. All working ok | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-161365 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Reference
ABL835 FLEX | Comprobar bombas y válvulas, cambiarlos los tubos de las bombas, calibrar y probar. | Cambiar los tubos de las bombas. | D1F0C2863C13CAC9 | No | Check pumps and valves, replace pump tubes, calibrate and test. | Change the pump tubes. | Defect | HW | pH/BG Module | Pump tubes
ABL825 FLEX | Customer rebooted analyzer several times but did not help. Checked power supply. | Replaced the CPU unit. | D1F088943C054D8E | No | Customer rebooted analyzer several times but did not help. Checked power supply. | Replaced the CPU unit. | Defect | HW | DMS_PC | CPU
ABL815 FLEX | Tausch der Pumpenschläuche. Durchführung einer LS Justage und einer Pumpenkalibration. Ein Lösungspumpenschlauch Funktion defekt. Tausch des Flachbandkabels im Autocheck Modul. Funktion nach Inbetriebnahme in Ordnung, Fehlermeldung erscheint nicht mehr unter Aktivitäten. | Tausch der Pumpenschläuche. Durchführung einer LS Justage und einer Pumpenkalibration. Ein Lösungspumpenschlauch Funktion defekt. Tausch des Flachbandkabels im Autocheck Modul. Funktion nach Inbetriebnahme in Ordnung, Fehlermeldung erscheint nicht mehr unter Aktivitäten. | SWO-160075 | No | Replacement of the pump hoses. Performing an LS adjustment and pump calibration. A solution pump hose function defective. Replacement of the ribbon cable in the Autocheck module. Function in order after commissioning, error message no longer appears under Activities. | Replacement of the pump hoses. Performing an LS adjustment and pump calibration. A solution pump hose function defective. Replacement of the ribbon cable in the Autocheck module. Function in order after commissioning, error message no longer appears under Activities. | Defect | HW | pH/BG Module | Pump tubes
ABL837 FLEX | Checked data logs. Tried to reset zero calibration. Reading no signal from electrodes. Tried electrode replacement for fault isolation. Order MetII module for replacement. | B | D1F06D363C51B1D2 | No | Checked data logs. Tried to reset zero calibration. Reading no signal from electrodes. Tried electrode replacement for fault isolation. Order MetII module for replacement. |  | Defect | HW | El/Met Module | Crea
ABL90 FLEX | oldat csere, készülék újra indítása, csere készülék beüzemelése, kalibrációk, QC mérés | B | SWO-160119 | No | replacement, device re-starting, Installation of the replacement apparatus, calibrations, QC measurement |  | User error | SP | SP Activation | SP Activation
ABL90 FLEX | 连续更换一张测试卡和一个试剂包 | 连续更换一张测试卡和一个试剂包 | SWO-160156 | No | Replace one test card and one reagent kit in a row | Replace one test card and one reagent kit in a row
AQT90 FLEX | removed the cup from incubator wheel and cleaned the wash block and cleaned all tubing, now working fine. | removed the cup from incubator wheel and cleaned the wash block and cleaned all tubing, now working fine. | SWO-160111 | No | removed the cup from incubator wheel and cleaned the wash block and cleaned all tubing, now working fine. | removed the cup from incubator wheel and cleaned the wash block and cleaned all tubing, now working fine. | Defect | Consumables | SP | SP | 1999
ABL80 FLEX BASIC | 更换测试卡修复 | 更换测试卡修复 | SWO-160219 | No | Replacement test card repair | Replacement test card repair
ABL90 FLEX | Fehlersuchverfahren in Serviceprogramm. Oxi-Schlauchventil ausgetauscht. Pumpenkal und LS- Justage. Neutart. | Fehlersuchverfahren in Serviceprogramm. Oxi-Schlauchventil ausgetauscht. Pumpenkal und LS- Justage. Neutart. | SWO-160210 | No | Error search procedure in service program. Oxi hose valve. Pump scale and LS adjustment. Neutart. | Error search procedure in service program. Oxi hose valve. Pump scale and LS adjustment. Neutart. | Defect | HW | Oxi module | Pinch valve
ABL80 FLEX BASIC |  | B | SWO-160150 | No
ABL80 FLEX BASIC | ゼロカセット、目視等で異常確認 | B | SWO-161747 | No | Check abnormalities by zero cassette, visual, etc.
ABL9 | Getät überprüft, Reklamationsdaten ausgedruckt, Datensicherung und QC durchgeführt | Getät überprüft, Reklamationsdaten ausgedruckt, Datensicherung und QC durchgeführt | SWO-160479 | No | Done checked, complaint data printed out, data backup and QC performed | Done checked, complaint data printed out, data backup and QC performed
ABL90 | Checked printer - no problem found. | Checked all tjhe analyzers and everything is working. | D1F55D4914121CC8 | No | Checked printer - no problem found. | Checked all tjhe analyzers and everything is working. | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | Analysator überprüft (kein Fehler feststellbar) - evt. Rücksprache bei neuen Erkenntnissen | Analysator überprüft (kein Fehler feststellbar) - evt. Rücksprache bei neuen Erkenntnissen | SWO-160388 | No | Analyzer checked (no error detectable) - evt. Consultation with new findings | Analyzer checked (no error detectable) - evt. Consultation with new findings | No fault found | No fault found | No fault found | pCO2
ABL90 FLEX | Lösungspack (XT06) eingebaut und (YB50) ausgebaut, System und Funktion i.O. --> Datalogs und SD gezogen | Lösungspack (XT06) eingebaut und (YB50) ausgebaut, System und Funktion i.O. --> Datalogs und SD gezogen | SWO-160386 | No | Solution pack (XT06) built-in and (YB50) removed, system and function in the --> Datalogs and SD pulled | Solution pack (XT06) built-in and (YB50) removed, system and function in the --> Datalogs and SD pulled | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX PLUS | Took UPS unit into workshop for further investigations. | B | SWO-160308 | Yes | Took UPS unit into workshop for further investigations. |  | Re classified RMED | UPS fault
ABL90 FLEX PLUS | Anwender hat bereits zweimal SP erneuert (LOT YB-50) - Analysator gespült und überprüft (kein Fehler feststellbar) - Erneuerung Anschlussdichtung - nochmalige Erneuerung SP (LOT XV05) - Analysator zeigt Kalibrierfehler pCO2 (Kalibrierdrift 2 außer Toleranz) - mehrfache Kalibration durchgeführt und QC gestartet - Analysator ohne Beanstandung ! | Anwender hat bereits zweimal SP erneuert (LOT YB-50) - Analysator gespült und überprüft (kein Fehler feststellbar) - Erneuerung Anschlussdichtung - nochmalige Erneuerung SP (LOT XV05) - Analysator zeigt Kalibrierfehler pCO2 (Kalibrierdrift 2 außer Toleranz) - mehrfache Kalibration durchgeführt und QC gestartet - Analysator ohne Beanstandung ! | SWO-160349 | No | User has already renewed SP twice (LOT YB-50) - Analyzer flushed and checked (no error detectable) - Renewal connection seal - Renewed renewal SP (LOT XV05) - Analyzer shows calibration error pCO2 (calibration drift 2 except tolerance) - multiple Calibration performed and QC started - analyzer without complaint ! | User has already renewed SP twice (LOT YB-50) - Analyzer flushed and checked (no error detectable) - Renewal connection seal - Renewed renewal SP (LOT XV05) - Analyzer shows calibration error pCO2 (calibration drift 2 except tolerance) - multiple Calibration performed and QC started - analyzer without complaint ! | Assisted Replacement | SP | SP | pCO2
AQT90 FLEX |  | B | SWO-160181 | No |  |  | Defect | HW | Process Module | Optical Unit | 135
ABL825 FLEX |  | B | D1F460BE9E2CCEC4 | No |  |  | Revisit required | Electrode | Electrode | pCO2
ABL815 FLEX | Ca- und Ref.- Elektode ersetzt, Probelauf, Kal. Und QC ok | Ca- und Ref.- Elektode ersetzt, Probelauf, Kal. Und QC ok | SWO-160423 | No | Ca- and Ref.- Elektode replaced, test run, Cal. And QC ok | Ca- and Ref.- Elektode replaced, test run, Cal. And QC ok | Assisted Replacement | Electrode | Electrode | Reference
ABL820 FLEX | Replaced the PC02 membrane and the dirift was 18 mm low on cal 2 and 10 mm low on cal 1. after 2 point passed and drift was corrected the QC came right up - before membrane values are at L1=64, L2=38.9, L3=20.9 After replacement values are at L1=66.8, L2=40.5, L3=22.9 this puts level 2 and 3 at a higher value than instrument #1 and less than 1 mm difference between both units on level 1. | Replaced Pco2 and REF membranes and cleaned dummy lower sealing ring. | D1F47ADF9E09E867 | No | Replaced the PC02 membrane and the dirift was 18 mm low on cal 2 and 10 mm low on cal 1. after 2 point passed and drift was corrected the QC came right up - before membrane values are at L1=64, L2=38.9, L3=20.9 After replacement values are at L1=66.8, L2=40.5, L3=22.9 this puts level 2 and 3 at a higher value than instrument #1 and less than 1 mm difference b etween both units on level 1. | Replaced Pco2 and REF membranes and cleaned dummy lower sealing ring. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX |  | B | D1F450C79E561547 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL825 FLEX | Membrane D788 Lot. R0759 undicht, Flüssigkeit tritt an Membrane aus. Elektrode erneut membraniert , weiterhin Empf. Fehler 50%. Mehrere Kal. durchgeführt . pCO2 Elektrode erneuert und mit gleicher Membran eingesetzt. Alle Werte o.k. | Membrane D788 Lot. R0759 undicht, Flüssigkeit tritt an Membrane aus. Elektrode erneut membraniert , weiterhin Empf. Fehler 50%. Mehrere Kal. durchgeführt . pCO2 Elektrode erneuert und mit gleicher Membran eingesetzt. Alle Werte o.k. | SWO-160482 | No | Membrane D788 Lot. R0759 leaks, liquid leaks to membrane. Electrode re-membraneed, continues empf. Error 50%. Several Kal. Performed. pCO2 electrode renewed and used with the same membrane. All values ok. | Membrane D788 Lot. R0759 leaks, liquid leaks to membrane. Electrode re-membraneed, continues empf. Error 50%. Several Kal. Performed. pCO2 electrode renewed and used with the same membrane. All values ok. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX |  | B | D1F48F4B9E067876 | No |  |  | Defect | HW | Oxi module | PCB, Oxi
ABL80 FLEX BASIC | cleaned the sensor board and signal data board, did zeroing for sensor connector board, did calibration and QC all values found with in the range instrument kept on observation for one day | compared the patient results and QC results with other instrument found all values with in the range. | SWO-162925 | No | cleaned the sensor board and signal data board, did zeroing for sensor connector board, did calibration and QC all values found with in the range instrument kept on observation for one day | compared the patient results and QC results with other instrument found all values with in the range.
ABL80 FLEX CO-OX | Analysator beim Eintreffen dicht, Gerät gereinigt, Funktionskontrolle undichter Lösungspack: Lot 229816 | Anwender hat SC gewechselt, Analysator macht keine Aktionen und zeigt Fehler Meldung 11111. Nach Service Manuel ist die Lösung 1 zu gering. Anwender tauscht nun auch den LP. Falls weiterhin Fehler, meldet sich der Anwender wieder. Analysator muss heute unbedingt laufen. | SWO-160578 | No | Analyzer on arrival dense, device cleaned, functional control of leaky solution pack: Lot 229816 | User has changed SC, analyzer does not take any actions and displays error message 11111. According to Service Manuel, solution 1 is too small. Users now also swap the LP. If errors persist, the user logs back in. Analyzer must run today.
ABL80 FLEX BASIC | 连续更换两张测试卡修复 | 连续更换两张测试卡修复 | SWO-161724 | No | Replace two test cards in a row for repair | Replace two test cards in a row for repair
ABL90 FLEX | QC durchgeführt | QC durchgeführt | SWO-160582 | Yes | QC performed | QC performed | Re classified RMED | ECKHI Email 24.jan-2117 | Unknown | pCO2 | QC Level C
ABL90 FLEX | Jahreswartung lt. Vertrag und RDEU Vorgaben durchgeführt, thb Kal durchgeführt , Kunde verwendet Kabe BK140 Kapillären zur Messung -> Kunden auf Einhaltung des Opl Vertages hingewiesen | B | SWO-160812 | No | Annual maintenance carried out according to contract and RDEU specifications, thb Kal carried out, customer uses Kabe BK140 capillary for measurement -> Customers advised to comply with the Opl Contract |  | User error | User error | User error | Capillary tubes | Kabe BK14
ABL90 FLEX PLUS |  | B | D1F6EAF38A068709 | No |  |  | Revisit required | HW | Oxi module | Oxi module | Data investigation
AQT90 FLEX |  | B | SWO-160463 | No |  |  | Cleaned | HW | Aspiration Module | HCT Sensor | 1267
ABL825 FLEX | ﾌﾘｸｼｮﾝﾃﾞｨﾃｸﾀｰﾕﾆｯﾄ交換。 | ﾌﾘｸｼｮﾝﾃﾞｨﾃｸﾀｰﾕﾆｯﾄ交換。 | SWO-160540 | No | Friction detector unit replacement. | Friction detector unit replacement. | Defect | HW | Inlet module | Friction detector unit
ABL825 FLEX | Netzwerk Verbindung geprüft, Hub neu gestartet, IT Hotline informiert | Netzwerk Verbindung geprüft, Hub neu gestartet, IT Hotline informiert | SWO-161450 | No | Network connection checked, hub restarted, IT hotline informed | Network connection checked, hub restarted, IT hotline informed | Setup | SW | His/lis | His/lis
ABL800 BASIC | Cleaned the touch screen properly,reconnect the touch screen connector and then switch on the instrument and touch screen was working properly. | Cleaned the touch screen properly,reconnect the touch screen connector and then switch on the instrument and touch screen was working properly. | SWO-161586 | No | Cleaned the touch screen properly,reconnect the touch screen connector and then switch on the instrument and touch screen was working properly. | Cleaned the touch screen properly,reconnect the touch screen connector and then switch on the instrument and touch screen was working properly. | Defect | HW | DMS_PC | Touch screen
ABL800 FLEX | QC setup done. QC ampul installed. As Lactate not calibrating ask for new lactate mambrane and QC sheduled for next visit. | Analyser Calibrate properly. QC run and found all values well with in range. Inform to managment and Dr. Indrajeet for same. Dr convensed and confirm to do all test on analyser. Analyser Handed over in a fine working condition. | SWO-163119 | No | QC setup done. QC ampul installed. As Lactate not calibrating ask for new lactate mambrane and QC sheduled for next visit. | Analyser Calibrate properly. QC run and found all values well with in range. Inform to managment and Dr. Indrajeet for same. Dr convensed and confirm to do all test on analyser. Analyser Handed over in a fine working condition. | Revisit required | HW | Auto Check Module | QC Ampul | Installed
ABL835 FLEX | pCO2 Elektrode (E788 XP-03) geprüft und ausgetauscht, Funktionskontrolle ok | pCO2 Elektrode (E788 XP-03) geprüft und ausgetauscht, Funktionskontrolle ok | SWO-161448 | No | pCO2 electrode (E788 XP-03) tested and replaced, function control ok | pCO2 electrode (E788 XP-03) tested and replaced, function control ok | Assisted Replacement | Electrode | Electrode | pCO2 | Glass breakage
ABL825 FLEX | Perform scan in service not reading all the ampoules | Replaced ribbon cable, now passes scan in service | D1F5D20E14012E12 | No | Perform scan in service not reading all the ampoules | Replaced ribbon cable, now passes scan in service | Defect | HW | Auto Check Module | FPC Cable
ABL825 FLEX | CONFIRMED ISSUE WHEN ONSITE. | INSPECT AND VERIFIED LAC SETTINGS. LAC CAL WAS PASSING WHEN ONSITE. ADVISED CUSTOMER TO CONTINUE MONITOR STATUS AND REPLACE LAC ELECTRODE IF NEEDED. | D1F5A0D3140480C0 | No | CONFIRMED ISSUE WHEN ONSITE. | INSPECT AND VERIFIED LAC SETTINGS. LAC CAL WAS PASSING WHEN ONSITE. ADVISED CUSTOMER TO CONTINUE MONITOR STATUS AND REPLACE LAC ELECTRODE IF NEEDED. | No fault found | Electrode | Membrane | Lac
AQT90 FLEX | gevşemiş olan dişli vidası sıkılarak sorun giderildi. | gevşemiş olan dişli vidası sıkılarak sorun giderildi. | SWO-160747 | Yes | the problem was fixed by tightening the threaded screw that had loosened. | the problem was fixed by tightening the threaded screw that had loosened.
AQT90 FLEX | Der Barcode wurde bereits am Gerät erkannt. Sehr wahrscheinlich wurde die Meldung übersehen, dass zu wenig Kapazitäten in der Pufferkassette waren, um eine Kalibration zu starten. Im Servicemenu zeigte das Gerät keine Fehler mehr. Pufferkassette erneuert, danach Kalibration durchgeführt. Anwender startet danach eine Qualitätskontrolle. Gerät überprüft. | Kal Kassette herausgenommen und neu eingelegt, auch jetzt wird der Barcode nicht abgefragt und die Kal nicht gestartet. Kundin startet den AQT neu. 2. Anruf: der AQT geht beim hochfahren in Fehler und fordert zum Neustart auf. Auch eine Wiederholung bringt das gleiche Ergebnis. Kundin bittet um einen Service. | SWO-160794 | No | The barcode has already been detected on the device. Most likely, the message that there was too little capacity in the buffer cartridge to start a calibration was overlooked. In the service menu, the device no longer showed any errors. Buffer cassette renewed, then calibrated. Users then start quality control. device checked. | Kal cassette removed and reinserted, even now the barcode is not queried and the cal is not started. Customer restarts the AQT. 2. Call: the AQT fails when booting up and prompts for a restart. A repetition also produces the same result. Customer asks for a service. | Unknown | Not Cat. | Not Cat. | Not Cat.
ABL80 FLEX BASIC |  | B | SWO-160817 | No
ABL80 FLEX BASIC | compared the patient results and QC results with other instrument found all values with in the range. | compared the patient results and QC results with other instrument found all values with in the range. | SWO-162927 | No | compared the patient results and QC results with other instrument found all values with in the range. | compared the patient results and QC results with other instrument found all values with in the range.
ABL90 FLEX |  | B | SWO-160997 | No |  |  | Assisted Replacement | SC | SC | Glu/Lac
ABL90 FLEX | manuell gespült, Qwetschschlauch getauscht, SC getauscht, Kaltstartphase, manuelle Kalibrierungen durchgeführt, QC,s bei PH und pCO2 nicht bestanden, LP getauscht, Kal und QC IO | manuell gespült, Qwetschschlauch getauscht, SC getauscht, Kaltstartphase, manuelle Kalibrierungen durchgeführt, QC,s bei PH und pCO2 nicht bestanden, LP getauscht, Kal und QC IO | SWO-161063 | No | manually rinsed, qwetsch hose exchanged, SC exchanged, cold start phase, manual calibrations performed, QC,s failed at PH and pCO2, LP exchanged, Kal and QC IO | manually rinsed, qwetsch hose exchanged, SC exchanged, cold start phase, manual calibrations performed, QC,s failed at PH and pCO2, LP exchanged, Kal and QC IO | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Hämolysator (R0698N009) und Schlauch am Quetschventil erneuert, System und Funktion i.O. --> SD gezogen | Hämolysator (R0698N009) und Schlauch am Quetschventil erneuert, System und Funktion i.O. --> SD gezogen | SWO-160978 | No | Hemolyser (R0698N009) and hose on the crushing valve renewed, system and function in the order --> SD pulled | Hemolyser (R0698N009) and hose on the crushing valve renewed, system and function in the order --> SD pulled | Defect | HW | Oxi module | Hemolyzer
TCM5 BASIC | Unable to explain cause of error, monitor removed for investigation, (Sensor & Mains lead left with customer). | Unable to explain cause of error, monitor removed for investigation, (Sensor & Mains lead left with customer). | SWO-161014 | No | Unable to explain cause of error, monitor removed for investigation, (Sensor & Mains lead left with customer). | Unable to explain cause of error, monitor removed for investigation, (Sensor & Mains lead left with customer).
TCM5 BASIC | System log & service dump to be evaluated. | System log & service dump to be evaluated. | SWO-161157 | No | System log & service dump to be evaluated. | System log & service dump to be evaluated.
ABL815 FLEX |  | B | D1F6F9D18A19F131 | No |  |  | Cleaned | HW | Oxi module | Hemolyzer
ABL825 FLEX | Fehlersuche Wartung Fließwege, Einlaß, Luftansaugstutzen Rollenkranz vonLösungspumpe gewartet Ventilwartung Oximodul ( scheint etwas zu haken) LS justage , Pumpenkalibrierung Austausch div. Schläuche ABL meldet nach Neustart bei Testmessung OXI Hardwarefehler, aber keine Meldung unter Systemmeldung mehrfach Neustart durchgeführt Testmessungen | Kunde bittet um einen erneuten Service. | SWO-160872 | No | Troubleshooting Maintenance flow paths, inlet, air intake nozzle Roller wreath serviced by solution pump Valve maintenance Oximodul (seems to hook a bit) LS adjustment , pump calibration Replacement of div. hoses ABL reports OXI hardware failure after restart on test measurement, but no message under system message Restarted multiple times Test measurements | Customer asks for a new service. | Defect | HW | Fluidic manifold | Pump tubes | Probably revisit
ABL825 FLEX | Customer stating that PH electrode and measuring chamber would fill up with blood numerous times. | Replaced measuring chamber, replaced electrode jacket for PH electrode and also replaced ph sealing ring and replace upper sealing gasket. | D1F6FBC58A620C76 | No | Customer stating that PH electrode and measuring chamber would fill up with blood numerous times. | Replaced measuring chamber, replaced electrode jacket for PH electrode and also replaced ph sealing ring and replace upper sealing gasket. | Defect | Electrode | Electrode | pH | Sealing ring
ABL825 FLEX | Customer stating that PH electrode and measuring chamber would fill up with blood numerous times. | Replaced measuring chamber, replaced electrode jacket for PH electrode and also replaced ph sealing ring and replace upper sealing gasket. | D1F6D9E98A570C6B | No | Customer stating that PH electrode and measuring chamber would fill up with blood numerous times. | Replaced measuring chamber, replaced electrode jacket for PH electrode and also replaced ph sealing ring and replace upper sealing gasket. | Defect | Electrode | Electrode | pH | Sealing ring
ABL825 FLEX |  | B | D1F6ECEE8A4C870B | No |  |  | Adjustment | HW | Oxi module | Lamp unit | Halogen
ABL800 BASIC | 装置引上げの為、代替機設置・動作確認 | 装置引上げの為、代替機設置・動作確認 | SWO-161040 | No | Installation and operation confirmation of alternative machines for equipment pulling up | Installation and operation confirmation of alternative machines for equipment pulling up | Unknown | Electrode | Electrode | pCO2
ABL800 BASIC | 修理完了装置入替 CAL・QC測定データ確認 | 修理完了装置入替 CAL・QC測定データ確認 | SWO-161030 | No | Replacement of repair completion equipment Cal/QC measurement data confirmation | Replacement of repair completion equipment Cal/QC measurement data confirmation | Defect | HW | Inlet module | Friction detector unit | pCO2 and CL issues as well
ABL825 FLEX | Fluid leaking from inlet. Suspect is worn pump tubing or clogged inlet gasket. | Replaced waste pump tubing. Verified instrument operation. | D1FC47FB625007AE | No | Fluid leaking from inlet. Suspect is worn pump tubing or clogged inlet gasket. | Replaced waste pump tubing. Verified instrument operation. | Defect | HW | Fluidic manifold | Pump tubes
ABL825 FLEX |  | B | D1F7ED230016CED7 | Yes |  |  | Re classified RMED | Planned re install module
ABL825 FLEX |  | Für die Optimierung der Meßergebnisse sollte eine Service-THb-Kal gestartet werden. Der Kunde bittet um Rückruf, um abzuklären, ob die Wartung vom Februar noch vor dem Ringversuch durchgeführt werden kann. | SWO-160965 | Yes |  | A service THb-Kal should be started to optimize the measurement results. The customer asks for a callback to determine whether the Maintenance from February can be carried out before the ring test. | Re classified RMED
ABL800 FLEX | Analyser Calibrate properly. QC run and found all values well with in range. Inform to managment and Dr. Indrajeet for same. Dr convensed and confirm to do all test on analyser. Analyser Handed over in a fine working condition. | Analyser Calibrate properly. QC run and found all values well with in range. Inform to managment and Dr. Indrajeet for same. Dr convensed and confirm to do all test on analyser. Analyser Handed over in a fine working condition. | SWO-163120 | No | Analyser Calibrate properly. QC run and found all values well with in range. Inform to managment and Dr. Indrajeet for same. Dr convensed and confirm to do all test on analyser. Analyser Handed over in a fine working condition. | Analyser Calibrate properly. QC run and found all values well with in range. Inform to managment and Dr. Indrajeet for same. Dr convensed and confirm to do all test on analyser. Analyser Handed over in a fine working condition. | Assisted Replacement | HW | Auto Check Module | QC Ampul | Installed
ABL835 FLEX | Tested scanner. Will not light ujp | Replaced defective bar code scanner. | D1F6E41F8A08BB25 | No | Tested scanner. Will not light ujp | Replaced defective bar code scanner. | Defect | HW | Basic Unit | Barcode reader, Basic Unit
ABL825 FLEX |  | B | D1F7744E8A0CC4AB | No |  |  | Cleaned | HW | pH/BG Module | Measuring chamber
ABL825 FLEX | Complete OXI test passed Lamp adjustment pass | Power down and restart. Errors cleared Analyzer is not live - Still testing Calibrations passed and QC passed. | D1F6E5398A16CA4E | No | Complete OXY test passed Lamp adjustment Pass | Power down and restart. Errors cleared Analyzer is not live - Still testing Calibrations passed and QC passed. | Adjustment | HW | Oxi module | Lamp unit
AQT90 FLEX | Changed all parts on machine that were outlined by the Techline investigated by Radiometer RMED. 6 month visit was due but main service visit carried out. All parts of main kit changed. Bank tests and LQC ran on board. Site to Re-calibrate and run LQC all parameters. Site to monitor. | Changed all parts on machine that were outlined by the Techline investigated by Radiometer RMED. 6 month visit was due but main service visit carried out. All parts of main kit changed. Bank tests and LQC ran on board. Site to Re-calibrate and run LQC all parameters. Site to monitor. | SWO-160920 | No | Changed all parts on machine that were outlined by the Techline investigated by Radiometer RMED. 6 month visit was due but main service visit carried out. All parts of main kit changed. Bank tests and LQC ran on board. Site to Re-calibrate and run LQC all parameters. Site to monitor. | Changed all parts on machine that were outlined by the Techline investigated by Radiometer RMED. 6 month visit was due but main service visit carried out. All parts of main kit changed. Bank tests and LQC ran on board. Site to Re-calibrate and run LQC all parameters. Site to monitor. | Unknown | Techline | Techline | Techline
AQT90 FLEX |  | B | D1F6CAD88A5518B6 | No |  |  | Defect | Consumables | Test Kit | Cartridge, PCT | Broken
ABL80 FLEX BASIC | 更换2张测试卡后修复 | 更换2张测试卡后修复 | SWO-161332 | No | Repair after replacing 2 test cards | Repair after replacing 2 test cards
ABL90 FLEX | Erneuerung der Sensorkassette und Solution pack | Erneuerung der Sensorkassette und Solution pack | SWO-161315 | Yes | Renewal of the sensor cartridge and solution pack | Renewal of the sensor cartridge and solution pack | Re classified RMED | ECKHI Email 24.jan-2117 | SP | FT | Leakage
ABL90 | Had analyzer shipped to my home. I turned analyzer on and installed SC and SP. It continued to give the ambient temp error. I calibrated the thermisters in service and the hemolyzer would not go above 31.12C. I also had total blockage in waste line. Flush that out. I removed hemolyzer and flushed and checked all connections. Analyzer would start initializing and then go to ambient temp error. I spoke with Andy and will ship this one back. | Will ship back to Andy Siwierka with tracking # 7782 1350 1539. He will ship out replacement. | D1F82B390046E1B2 | Yes | Had analyzer shipped to my home. I turned analyzer on and installed SC and SP. It continued to give the ambient temp error. I calibrated the thermisters in service and the hemolyzer would not go above 31.12C. I also had total blockage in waste line. Flush that out. I removed hemolyzer and flushed and checked all connections. Analyzer would start initializing and then go to ambient temp error. I spoke with Andy and will ship this one back. | Will ship back to Andy Siwierka with tracking # 7782 1350 1539. He will ship out replacement. | Re classified RMED | OOB
ABL90 FLEX | SC wurde routinemäßig am 17.01.17 7:33Uhr ersetzt ->QC Fehler bei Glu Lac bis 18.01.17 7:12 Uhr ->Gerät empfiehlt LP wechsle LP wurde ersetzt DataLogs & SD erstellt und an TS geschickt | SC wurde routinemäßig am 17.01.17 7:33Uhr ersetzt ->QC Fehler bei Glu Lac bis 18.01.17 7:12 Uhr ->Gerät empfiehlt LP wechsle LP wurde ersetzt DataLogs & SD erstellt und an TS geschickt | SWO-161248 | No | SC was routinely replaced on 17.01.17 7:33o ->QC error at Glu Lac until 18.01.17 7:12 o'clock ->Device recommends LP change LP has been replaced DataLogs & SD created and sent to TS | SC was routinely replaced on 17.01.17 7:33o ->QC error at Glu Lac until 18.01.17 7:12 o'clock ->Device recommends LP change LP has been replaced DataLogs & SD created and sent to TS | Assisted Replacement | SC | SC | Glu/Lac | Start up
ABL90 FLEX PLUS | Done trouble shooting in service mode to try and correct error but had to replace sensor cassette and error went away immediately. All calibrations ok and Qc's ok. | Done trouble shooting in service mode to try and correct error but had to replace sensor cassette and error went away immediately. All calibrations ok and Qc's ok. | SWO-161415 | No | Done trouble shooting in service mode to try and correct error but had to replace sensor cassette and error went away immediately. All calibrations ok and Qc's ok. | Done trouble shooting in service mode to try and correct error but had to replace sensor cassette and error went away immediately. All calibrations ok and Qc's ok. | Assisted Replacement | SC | SC | Cl
ABL90 FLEX | 更换902-980和903-197修复(By Ran Gao) | 更换902-980和903-197修复(By Ran Gao) | SWO-161055 | No | Replacement 902-980 and 903-197 repairs (By Ran Gao) | Replacement 902-980 and 903-197 repairs (By Ran Gao) | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | 工程师报错机器序列号，此单取消 （By Ran Gao） | 工程师报错机器序列号，此单取消 （By Ran Gao） | SWO-161071 | Yes | The engineer incorrectly reported the machine serial number and this order was cancelled (By Ran Gao) | The engineer incorrectly reported the machine serial number and this order was cancelled (By Ran Gao) | Re classified RMED | AQT issue
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-161755 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | pO2 | 375
AQT90 FLEX | 更换备件903-364后修复(By Ran Gao) | 更换备件903-364后修复(By Ran Gao) | SWO-161033 | Yes | Replacement of spare parts 903-364 after repair (By Ran Gao) | Replacement of spare parts 903-364 after repair (By Ran Gao) | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-161077 | No |  |  | Cleaned | HW | Storage Module | Process wheel | 161
ABL825 FLEX | Messkammer gereinigt, Ref Elektrode geprüft, Dichtungsring an der pH Elektrode erneuert. | Messkammer gereinigt, Ref Elektrode geprüft, Dichtungsring an der pH Elektrode erneuert. | SWO-161290 | No | Measuring chamber cleaned, ref electrode tested, sealing ring renewed at pH electrode. | Measuring chamber cleaned, ref electrode tested, sealing ring renewed at pH electrode. | Defect | Electrode | Electrode | pH | Sealing ring
AQT90 FLEX |  | B | D1F8266E004A0D56 | No
ABL80 FLEX CO-OX | Gerät mit gelber Ampel => LP nur noch 23 Zyklen Kal und QC ohne Fehler => Oximetrie nicht aktiviert Durchführung Service Test => Fehler bei Lampenstatus (LED Max Intensität zu hoch) und bei Hämolysatorstatus (Kavitation Dämpfen zu tief) Gerät zerlegt und Hemolyzer optisch überprüft => QC Austritt an oberer Schlauchverbindung und LS => Hemolyzer gereinigt und gespült Geräteneustart => Flüssigkeitssensorkalibrierung i.O. Oximetrieparameter aktiviert => anschließende Kal und QC i.O. Uhrzeit gestellt Analysator trennt und verbindet sich jede Minute fortlaufend mit Radiance | Gerät mit gelber Ampel => LP nur noch 23 Zyklen Kal und QC ohne Fehler => Oximetrie nicht aktiviert Durchführung Service Test => Fehler bei Lampenstatus (LED Max Intensität zu hoch) und bei Hämolysatorstatus (Kavitation Dämpfen zu tief) Gerät zerlegt und Hemolyzer optisch überprüft => QC Austritt an oberer Schlauchverbindung und LS => Hemolyzer gereinigt und gespült Geräteneustart => Flüssigkeitssensorkalibrierung i.O. Oximetrieparameter aktiviert => anschließende Kal und QC i.O. Uhrzeit gestellt Analysator trennt und verbindet sich jede Minute fortlaufend mit Radiance | SWO-161480 | No | Device with yellow traffic light => LP only 23 cycles Kal and QC without error => Oximetry not activated Execution Service Test => Error in lamp status (LED Max intensity too high) and hemolyser status (cavitation vapours too low) Disassembled device and optically checked Hemolyzer => QC outlet to upper hose connection and LS => Hemolyzer cleaned and rinsed Device restart => Liquid sensor calibration in the o. Oximetry parameter activated => subsequent Cal and QC in the following Time set Analyzer separates and connects to Radiance continuously every minute | Device with yellow traffic light => LP only 23 cycles Kal and QC without error => Oximetry not activated Execution Service Test => Error in lamp status (LED Max intensity too high) and hemolyser status (cavitation vapours too low) Disassembled device and optically checked Hemolyzer => QC outlet to upper hose connection and LS => Hemolyzer cleaned and rinsed Device restart => Liquid sensor calibration in the o. Oximetry parameter activated => subsequent Cal and QC in the following Time set Analyzer separates and connects to Radiance continuously every minute
ABL80 FLEX BASIC | connected zeroing cassette and zeroing done and installed solution pack and sensor cassette again its showing verification check failure .we need to replace sensor cassette connector board and signal data board for this isuee. | B | SWO-162943 | No | connected zeroing cassette and zeroing done and installed solution pack and sensor cassette again its showing verification check failure .we need to replace sensor cassette connector board and signal data board for this isuee.
ABL90 FLEX | neues LP mit anderer LOT geliefert und initialisieret, alle QC Level bestanden | neues LP mit anderer LOT geliefert und initialisieret, alle QC Level bestanden | SWO-161368 | No | new LP with other LOT supplied and initialized, all QC levels passed | new LP with other LOT supplied and initialized, all QC levels passed | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX | Lösungspack richtig arretiert, Kaltstartphase, Kalibrierung und alle drei QC,s IO System messbereit | Lösungspack richtig arretiert, Kaltstartphase, Kalibrierung und alle drei QC,s IO System messbereit | SWO-161592 | No | Solution pack properly locked, cold start phase, calibration and all three QC,s IO system ready for measurement | Solution pack properly locked, cold start phase, calibration and all three QC,s IO system ready for measurement | Assisted Replacement | SP | SP | pCO2
ABL90 FLEX | En ouvrant la fenêtre du SC je constate que la partie supérieur bouge un peu vers la gauche. Après démontage je vois que la partie supérieur du module Sensor Interface n'est pas bien fixe et donc ne se positionne pas toujours à la bonne place dans la Sensor Cassette. Après serrage de 2 visses , ajustage des contacts au dessus de la cassette, et remontage du module j'ai redémarré l'ABL et j'ai passé plusieurs calibrations (test de pO2 fait dans programme de service) sans erreur. Les 3 contrôles sont bien passées également.. | En ouvrant la fenêtre du SC je constate que la partie supérieur bouge un peu vers la gauche. Après démontage je vois que la partie supérieur du module Sensor Interface n'est pas bien fixe et donc ne se positionne pas toujours à la bonne place dans la Sensor Cassette. Après serrage de 2 visses , ajustage des contacts au dessus de la cassette, et remontage du module j'ai redémarré l'ABL et j'ai passé plusieurs calibrations (test de pO2 fait dans programme de service) sans erreur. Les 3 contrôles sont bien passées également.. | SWO-161510 | No | When I open the SC window I notice that the top part moves a little to the left. After dismantling I see that the upper part of the Sensor Interface module is not well fixed and therefore does not always position itself in the right place in the Sensor Cassette. After tightening 2 screws, adjusting the contacts above the cassette, and reassembling the module I restarted the ABL and I passed several calibrations (pO2 test done in service program) without error. All three checks are well passed as well. | When I open the SC window I notice that the top part moves a little to the left. After dismantling I see that the upper part of the Sensor Interface module is not well fixed and therefore does not always position itself in the right place in the Sensor Cassette. After tightening 2 screws, adjusting the contacts above the cassette, and reassembling the module I restarted the ABL and I passed several calibrations (pO2 test done in service program) without error. All three checks are well passed as well. | Adjustment | HW | Basic Unit | Sensor bay
ABL90 FLEX PLUS | Cleaned all electrical contact points. Reinitiliased analyzer. | Cleaned all electrical contact points. Reinitiliased analyzer. | SWO-161795 | No | Cleaned all electrical contact points. Reinitiliased analyzer. | Cleaned all electrical contact points. Reinitiliased analyzer. | Cleaned | HW | Basic Unit | Solution Pack PCB Id
ABL90 FLEX PLUS |  | B | D1F958057622F00D | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | New Oxi pinch valve fitted to the analyser. | B | SWO-161565 | No | New Oxi pinch valve fitted to the analyser. |  | Defect | HW | Oxi module | Pinch valve
ABL825 FLEX | Austausch Meßkammern neue Ventile und Ventilgehäuse am BG und EL Modul Dichtring pH El. erneuert Elektroden membraniert Einlaufverfahren | Austausch Meßkammern neue Ventile und Ventilgehäuse am BG und EL Modul Dichtring pH El. erneuert Elektroden membraniert Einlaufverfahren | SWO-161518 | No | Replacement of measuring chambers new valves and valve housings on the BG and EL modules Sealing ring pH El. Renewed Electrodes membraneed Entry procedure | Replacement of measuring chambers new valves and valve housings on the BG and EL modules Sealing ring pH El. Renewed Electrodes membraneed Entry procedure | Defect | Electrode | Electrode | pH | Sealing ring
ABL825 FLEX | Wiederherstellung der Layouts nicht möglich, Konfiguration von weiterem Gerät im Hause geholt und eingespielt, Layouts angepasst, fällige Membran ersetzt, mehrere Kal, QC, Patienten ID Variable eingerichtet | B | SWO-161606 | No | Restoration of layouts not possible, configuration of further device fetched in the house and recorded, layouts adapted, overdue membrane replaced, multiple Kal, QC, patient ID variable set up |  | Setup | SW | Patient Layout | Patient Layout
ABL825 FLEX | Bei Ankunft war Analysator Messbereit. Im Archiv war zu sehen dass die Kals seit geraumer Zeit abgebrochen wurden weil erst Spüllösung leer war und dann Kal 2 Lösung. Diese wurden jedoch in der Früh von Frau Schirp getauscht. Danach war die Kal wieder i.O. Frau Guth und Frau Schirp darüber informiert warum so viele Kals abgebrochen wurden. Trotzdem die Ventile aus PH BG und ELE MET Modul ausgebaut, wegen den ab und zu vorherrschenden Driften bei Lac. Beide Luftventile waren fest. Beide Ventile gangbar gemacht. Luftkanal gespült. LS Justage und Pumpenkal durchgeführt. Mehrere Kalibrierungen durchgeführt --> i.O. Keine fehlerhaften Driften mehr vorhanden. QCs für alle Level durchgeführt. Fragen von Personal ( Frau Guth, Schirp )zu Kalibrierung, Driften und Lösungstausch beantwortet und am ABL Menüpunkte zur Problemlösung aufgezeigt. ABL wieder Messbereit. Zwei Patientenmessungen durchführen lassen. --> diese waren i.O. | Bei Ankunft war Analysator Messbereit. Im Archiv war zu sehen dass die Kals seit geraumer Zeit abgebrochen wurden weil erst Spüllösung leer war und dann Kal 2 Lösung. Diese wurden jedoch in der Früh von Frau Schirp getauscht. Danach war die Kal wieder i.O. Frau Guth und Frau Schirp darüber informiert warum so viele Kals abgebrochen wurden. Trotzdem die Ventile aus PH BG und ELE MET Modul ausgebaut, wegen den ab und zu vorherrschenden Driften bei Lac. Beide Luftventile waren fest. Beide Ventile gangbar gemacht. Luftkanal gespült. LS Justage und Pumpenkal durchgeführt. Mehrere Kalibrierungen durchgeführt --> i.O. Keine fehlerhaften Driften mehr vorhanden. QCs für alle Level durchgeführt. Fragen von Personal ( Frau Guth, Schirp )zu Kalibrierung, Driften und Lösungstausch beantwortet und am ABL Menüpunkte zur Problemlösung aufgezeigt. ABL wieder Messbereit. Zwei Patientenmessungen durchführen lassen. --> diese waren i.O. | SWO-161509 | No | On arrival, analyzer was ready to measure. In the archive it could be seen that the kals were canceled for quite some time because first the flushing solution was empty and then Kal 2 solution. However, these were exchanged in the morning by Mrs Schirp. Afterwards, the Kal was again in the I.O. Mrs. Guth and Ms. Schirp informed about why so many kals were canceled. Nevertheless, the valves from PH BG and ELE MET module were removed, because of the from time to time prevailing drifts at Lac. Both air valves were fixed. Both valves made workable. Air duct flushed. LS adjustment and pump scale. Several calibrations performed --> in the end No more faulty drifts. QCs for all levels. Questions answered by staff ( Ms. Guth, Schirp ) about calibration, drifting and solution exchange and shown on the ABL menu items for problem solving. ABL ready to measure again. Have two patient measurements carried out. --> these were in the end. | On arrival, analyzer was ready to measure. In the archive it could be seen that the kals were canceled for quite some time because first the flushing solution was empty and then Kal 2 solution. However, these were exchanged in the morning by Mrs Schirp. Afterwards, the Kal was again in the I.O. Mrs. Guth and Ms. Schirp informed about why so many kals were canceled. Nevertheless, the valves from PH BG and ELE MET module were removed, because of the from time to time prevailing drifts at Lac. Both air valves were fixed. Both valves made workable. Air duct flushed. LS adjustment and pump scale. Several calibrations performed --> in the end No more faulty drifts. QCs for all levels. Questions answered by staff ( Ms. Guth, Schirp ) about calibration, drifting and solution exchange and shown on the ABL menu items for problem solving. ABL ready to measure again. Have two patient measurements carried out. --> these were in the end. | Cleaned | HW | pH/BG Module | Valve, air
ABL825 FLEX | Examined inle assembly. Found bent probe. | Replaced inlet probe, with customer stock. Instructed customer on inlet probe replacement. Verified operation of ABL825. | D1F95A51764008A6 | No | Examined inle assembly. Found bent probe. | Replaced inlet probe, with customer stock. Instructed customer on inlet probe replacement. Verified operation of ABL825. | Defect | HW | Inlet module | Probe bend
AQT90 FLEX | サービスでは取り出せないので、分解してカートリッジを取り出す。下まで分解できないので、近くを確認したが破損のツメは見つからない。動作確認 | B | SWO-161782 | No | Because it cannot be taken out by the service, it disassembles it and the cartridge is taken out. Because it is not possible to disassemble it to the bottom, the swallow of the damage is not found though it confirmed near. Check operation |  | Defect | Consumables | Test Kit | Cartridge | Broken
AQT90 FLEX | Tube Holder erneuert. Abgebrochenes Teil von Tube Holder aus gerät entfernt. Nadel ausgebaut und überprüft - ok. LQC durchgeführt. Fehler im Zuge der Wartung beseitigt. | Tube Holder erneuert. Abgebrochenes Teil von Tube Holder aus gerät entfernt. Nadel ausgebaut und überprüft - ok. LQC durchgeführt. Fehler im Zuge der Wartung beseitigt. | SWO-161500 | No | Tube Holder renewed. Aborted part of Tube Holder from device removed. Needle removed and checked - ok. LQC. Fixed errors during maintenance. | Tube Holder renewed. Aborted part of Tube Holder from device removed. Needle removed and checked - ok. LQC. Fixed errors during maintenance. | Defect | HW | Aspiration Module | Tube holder
AQT90 FLEX | System auf Cups überprüft, Motorspindeln gereinigt und geschmiert. Servicedump gesichert, Funktion z.Zt. OK. | In den Aktivitäten: 1067,34114-FTC 105 Too few Counts from LS2 und 5670- Rotateinletwheel 25 auf. | SWO-161880 | No | System checked for cups, motor spindles cleaned and lubricated. Service dump secured, function currently OK. | In the activities: 1067,34114-FTC 105 Too few counts from LS2 and 5670- Rotateinletwheel 25 on.
AQT90 FLEX | Gerät im Servicemenu überprüft. Cupsensor Threshold bei 18. servicedump gezogen und auswerten lassen. Cupsensor gereinigt und Patientenmessungen durchgeführt. Sollte Fehler wieder auftreten, muss der Sensor erneuert werden. E-Teil bestellt. | Gerät am 20.01.17 zur Reparatur abends abgeholt. Cup-Sensor erneuert und kalibriert. Blutmessungen und Myoglobin-Kalibration durchgeführt. Fehler trat nicht mehr auf. Gerät am 23.01.17 wieder beim Kunden aufgestellt und mit Qualitätakontrolle und weiteren Blutmessungen überprüft. | SWO-162081 | No | device in the service menu. Cupsensor Threshold at 18th servicedump and have it evaluated. cup sensor and patient measurements carried out. If errors occur again, the sensor must be replaced. E-part. | Device picked up on 20.01.17 for repair in the evening. Cup sensor renewed and calibrated. blood measurements and myoglobin calibration. Errors no longer occurred. Device placed again at the customer on 23.01.17 and checked with quality control and further blood measurements. | Revisit required | HW | Incubator module | Sensor, Cup
AQT90 FLEX | 更换903-364,重新定标TNI定标卡944-212（12162），ckmb定标卡944-216（12179），myo定标卡944-214（12338），D二聚体定标卡944-220（12124）后修复 （By Ran Gao） | 更换903-364,重新定标TNI定标卡944-212（12162），ckmb定标卡944-216（12179），myo定标卡944-214（12338），D二聚体定标卡944-220（12124）后修复 （By Ran Gao） | SWO-161379 | No | Replace 903-364, re-order TNI calibration card 944-212 (12162), ckmb calibration card 944-216 (12179), myo calibration card 944-214 (12338), D two polyp calibration card 944-220 (12124) after repair (By Ran Gao) | Replace 903-364, re-order TNI calibration card 944-212 (12162), ckmb calibration card 944-216 (12179), myo calibration card 944-214 (12338), D two polyp calibration card 944-220 (12124) after repair (By Ran Gao) | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | 更换636-571和910-298修复 by andrew | 更换636-571和910-298修复 by andrew | SWO-161679 | No | Replace 636-571 and 910-298 Repair by Andrew | Replace 636-571 and 910-298 Repair by Andrew | Defect | HW | Analyzer Control Module | FPC Cable
ABL90 FLEX | LP erneuert | LP erneuert | SWO-161823 | No | LP renewed | LP renewed | User Error | SP | SP | Empty
ABL90 FLEX | Patientenlayout stimmt nicht mehr - Falsche Vorgabe gesetzt. Das richtige Layout wieder als Standard eingestellt. | Kunde bittet um einen erneuten Besuch und Korrektur der Einstellungen. | SWO-162000 | Yes | Patient layout is no longer correct - wrong default set. The correct layout is set as the default again. | Customer asks for a revisit and correction of the settings. | Re classified RMED | Phone support
ABL90 FLEX | Da sich durch den Fehler der Touchscreen nicht bedienen lässt , testweise neuen Monitor eingebaut. Kein Erfolg! Sicherungen auf AU Board getestet. o.k. Steckerverbindungen getestet o.k. Gerät zur Reparatur in GS17 genommen. | Geräteaufstellung in OP nach Reparatur in GS17. | SWO-162695 | No | Since the error does not allow the touchscreen to be operated, new monitor built-in on a test basis. No success! Fuses tested on AU board. Ok. Plug connections tested o.k. Device taken for repair in GS17. | Equipment installation in op after repair in GS17. | Revisit required | analyzer removed | Power Supply | Power Supply | See UpSWO 25/1
ABL90 |  | B | D2024E4FB0540019 | No |  |  | Assisted Replacement | SP | SP | FT
ABL90 FLEX PLUS | Analysator überprüft - Einlass demontiert und Anschlussdichtung erneuert - Rückmontage und Neustart durchgeführt - Funktionsüberprüfung etc. ohne Beanstandung ! Bemerkung: Bei telefonischer Unterstützung am Wochenende schloss Einlass bei SP-Wechsel erst nach Reset des Analysators ! | Analysator überprüft - Einlass demontiert und Anschlussdichtung erneuert - Rückmontage und Neustart durchgeführt - Funktionsüberprüfung etc. ohne Beanstandung ! Bemerkung: Bei telefonischer Unterstützung am Wochenende schloss Einlass bei SP-Wechsel erst nach Reset des Analysators ! | SWO-161836 | No | Analyzer checked - inlet disassembled and connection seal renewed - reassembly and restart performed - function check etc. without complaint ! Note: With telephone support at the weekend, admission for SP change only closed after the reset of the analyzer ! | Analyzer checked - inlet disassembled and connection seal renewed - reassembly and restart performed - function check etc. without complaint ! Note: With telephone support at the weekend, admission for SP change only closed after the reset of the analyzer !
ABL90 FLEX PLUS | Aanpassing Roma.ini | Aanpassing Roma.ini | SWO-161975 | Yes | Adaptation Roma. ini | Adaptation Roma. ini | Re classified RMED | Changing Roma.ini on customer request
ABL90 FLEX PLUS | Kein Lösungspack einer anderen Charge vor Ort. | Kein Lösungspack einer anderen Charge vor Ort. | SWO-161993 | No | No solution pack of another batch on site. | No solution pack of another batch on site. | Assisted Replacement | SP | SP | pH
ABL90 FLEX PLUS | SP90 LOT YB-51 installiert, QC ok | Kein Lösungspack einer anderen Charge vor Ort. | SWO-161994 | Yes | SP90 LOT YB-51 installiert, QC ok | No solution pack of another batch on site. | Re classified RMED | Same day same problem
ABL90 FLEX PLUS | Aanpassing Roma.ini | Aanpassing Roma.ini | SWO-161972 | Yes | Adaptation Roma. ini | Adaptation Roma. ini | Re classified RMED | Changing Roma.ini on customer request
AQT90 FLEX |  | B | SWO-161722 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | Desmontado inlet wheel. Comprobados sensores y motores. | B | D1FD311CD84402E2 | No | Disassembled wheel inlet. Tested sensors and motors.
AQT90 FLEX | Sustituida aguja y sensor. Realizado ajuste de la aguja. | B | D1FD262ED85802DC | No | Replaced needle and sensor. Made needle adjustment.
ABL825 FLEX | AC Modul geprüft und demontiert, Flachbandkabel im AC Modul ausgetauscht, Funktionskontrolle ok | AC Modul geprüft und demontiert, Flachbandkabel im AC Modul ausgetauscht, Funktionskontrolle ok | SWO-162173 | No | AC module tested and disassembled, ribbon cable replaced in AC module, function control ok | AC module tested and disassembled, ribbon cable replaced in AC module, function control ok | Defect | HW | Auto Check Module | Flat cable
ABL815 FLEX | Fehlersuche, Reinigung, Einlaßdichtung entfernt, Schlauch angebracht und durchgespült, Abfallwege getestet und gespült, alles wieder montiert, LS Justage, Pumpenkal, Kal, QC | Kundin hat bereits die Einlassdichtung ersetzt, mit Kundin LS- Pumpenkal durchgeführt - ABL tropft weiterhin aus Einlass, alut Kundin ist die Einlasssonde nicht verbogen Kundin bittet um Service | SWO-162167 | No | Troubleshooting, cleaning, inlet seal removed, hose attached and flushed, waste paths tested and rinsed, everything reassembled, LS adjustment, pump scale, kal, QC | Customer has already replaced the inlet seal, performed with customer LS pump scale - ABL continues to drip out of inlet, alut customer is not bent the inlet probe Customer asks for service | Revisit required | HW | Inlet module | Leakage
ABL825 FLEX | Nach Rücksprache mit RDEU Service und Remote-Aufschaltung den Analysator aufgesucht, überprüft und ServiceDump sowie Datenextrakt gezogen und an EvaRa weitergeleitet zwecks Sendung an RMED und Auswertung. Zurzeit Ampel grün, keine Fehler. | Nach Rücksprache mit RDEU Service und Remote-Aufschaltung den Analysator aufgesucht, überprüft und ServiceDump sowie Datenextrakt gezogen und an EvaRa weitergeleitet zwecks Sendung an RMED und Auswertung. Zurzeit Ampel grün, keine Fehler. | SWO-161902 | No | After consultation with RDEU Service and remote connection, the analyzer is visited, checked and ServiceDump and data extract pulled and forwarded to EvaRa for transmission to RMED and evaluation. Currently traffic light green, no errors. | After consultation with RDEU Service and remote connection, the analyzer is visited, checked and ServiceDump and data extract pulled and forwarded to EvaRa for transmission to RMED and evaluation. Currently traffic light green, no errors.
ABL825 FLEX | Konfiguration angepasst | Konfiguration angepasst | SWO-162172 | No | Configuration adapted | Configuration adapted | Setup | SW | Configuration | Configuration | Shuld not be an upSWO
ABL815 FLEX | Installation und Konfiguration vorbereitet. Aufstellung in Abteilung MEI durch Jona Lünenbürger. | B | SWO-162690 | Yes | installation and configuration. Installation in Department MEI by Jona Lünenbürger. |  | Re classified RMED | HW | New Analyzer
AQT90 FLEX | removed the cartridge and inserted new one now working satisfactorily. also done PCT Calibration. | removed the cartridge and inserted new one now working satisfactorily. also done PCT Calibration. | SWO-161814 | No | removed the cartridge and inserted new one now working satisfactorily. also done PCT Calibration. | removed the cartridge and inserted new one now working satisfactorily. also done PCT Calibration. | Defect | Consumables | Test Kit | Cartridge, broken | 5689
AQT90 FLEX | Gerät am 20.01.17 zur Reparatur abends abgeholt. Cup-Sensor erneuert und kalibriert. Blutmessungen und Myoglobin-Kalibration durchgeführt. Fehler trat nicht mehr auf. Gerät am 23.01.17 wieder beim Kunden aufgestellt und mit Qualitätakontrolle und weiteren Blutmessungen überprüft. | Gerät am 20.01.17 zur Reparatur abends abgeholt. Cup-Sensor erneuert und kalibriert. Blutmessungen und Myoglobin-Kalibration durchgeführt. Fehler trat nicht mehr auf. Gerät am 23.01.17 wieder beim Kunden aufgestellt und mit Qualitätakontrolle und weiteren Blutmessungen überprüft. | SWO-162080 | No | Device picked up on 20.01.17 for repair in the evening. Cup sensor renewed and calibrated. blood measurements and myoglobin calibration. Errors no longer occurred. Device placed again at the customer on 23.01.17 and checked with quality control and further blood measurements. | Device picked up on 20.01.17 for repair in the evening. Cup sensor renewed and calibrated. blood measurements and myoglobin calibration. Errors no longer occurred. Device placed again at the customer on 23.01.17 and checked with quality control and further blood measurements. | Revisit required | HW | Incubator module | Sensor, Cup
AQT90 FLEX | Scanner gedemonteerd, kabelcontacten en lensjes schoongemaakt met alkohol. Serviceinstellingen gecontroleerd, OK. Nogmaals bepaling aangevraagd, scant direct buisje en barcode. Systeemreiniging en -controle met succes uitgevoerd. | B | SWO-161799 | No | Scanner disassembled, cable contacts and Lenes cleaned with Alkohol. Service settings checked, OK. Once again provision requested, scans direct tube and barcode. System cleaning and control performed successfully.
ABL90 FLEX | Mit dem Anwender (Frau Domberg) Ringversuche (RfB) durchgeführt und ausgewertet | Fr. Domberg bittet um eine Überprüfung. | SWO-162303 | No | Performed and evaluated with the user (Mrs. Domberg) ring tests (RfB) | Ms. Domberg asks for a review. | No fault found | Parameter | Parameter | Analyzer comparison
ABL90 FLEX | performed tHb calibration | performed tHb calibration | SWO-162286 | No | performed tHb calibration | performed tHb calibration | Adjustment | HW | Oxi module | tHb Cal
ABL90 FLEX | Replaced faulty Oxi valve tube. | Replaced faulty Oxi valve tube. | SWO-162206 | No | Replaced Faulty oxy valve tube. | Replaced Faulty oxy valve tube. | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX | Breites Flachbandkabel oberhalb des Sensorkassettenfaches wieder richtig eingebaut, Servicetest (Thermistorkalibration und Ansteuerung des Sensorkassettenfaches) --> beide Funktionen i.O., Startphase, System und Funktion i.O. | Breites Flachbandkabel oberhalb des Sensorkassettenfaches wieder richtig eingebaut, Servicetest (Thermistorkalibration und Ansteuerung des Sensorkassettenfaches) --> beide Funktionen i.O., Startphase, System und Funktion i.O. | SWO-162177 | No | Wide ribbon cable above the sensor cassette compartment reinstalled correctly, service test (thermistor calibration and control of the sensor cassette compartment) --> both functions i.O., start phase, system and function i.O. | Wide ribbon cable above the sensor cassette compartment reinstalled correctly, service test (thermistor calibration and control of the sensor cassette compartment) --> both functions i.O., start phase, system and function i.O. | Defect | HW | Sensor Compartment | FPC Cable | 166
ABL90 FLEX PLUS | Zusammen mit Herrn Lück die Netzwerkeinstellungen im Haus und am Gerät vorgenommen. QC Daten werden nun übertragen. Fehler im LP, Glu Level 1 fällt raus. Es ist kein anderes LP im Haus vorhanden. Neues LP ist bestellt. | Zusammen mit Herrn Lück die Netzwerkeinstellungen im Haus und am Gerät vorgenommen. QC Daten werden nun übertragen. Fehler im LP, Glu Level 1 fällt raus. Es ist kein anderes LP im Haus vorhanden. Neues LP ist bestellt. | SWO-162405 | No | Together with Mr. Lück, the network settings were made in the house and on the device. QC data is now transferred. Error in LP, Glu Level 1 falls out. There is no other LP in the house. New LP has been ordered. | Together with Mr. Lück, the network settings were made in the house and on the device. QC data is now transferred. Error in LP, Glu Level 1 falls out. There is no other LP in the house. New LP has been ordered. | Setup | SW | His/lis | His/lis
ABL835 FLEX | Comprobar tubos, cambiar la aguja toma muestras, calibrar y probar con muestras y controles. | Cambiar tubo inlet módulo pH. | D1FE56144E46945A | No | Check tubes, change the needle takes samples, calibrate and test with samples and controls. | Change tube inlet module pH. | Defect | HW | Basic Unit | Pre Heater tube
ABL837 FLEX |  | B | D1FE73E74E14C819 | No |  |  | Assisted Replacement | Electrode | Membrane | Cl
ABL820 FLEX | Customer had inlet liquid sensor error and oxi liquid sensor errors. | Showed customer how to flush breather valve on fluidic channel and flushed inlet gasket tubing. Performed service liquid sensor calibration. Verified instrument operation. | D2001AE7C43B75A0 | No | Customer had inlet liquid sensor error and oxi liquid sensor errors. | Showed customer how to flush breather valve on fluidic channel and flushed inlet gasket tubing. Performed service liquid sensor calibration. Verified instrument operation. | Cleaned | HW | Basic Unit | Fluidic pathway
ABL815 FLEX | Fehlersuche, QC abgebrochen dadurch alle Parameter gesperrt, Kalwerte durchgesehen - alle Parameter membraniert, mehrere Kal, QC | Kundin hat bereits die Einlassdichtung ersetzt, mit Kundin LS- Pumpenkal durchgeführt - ABL tropft weiterhin aus Einlass, alut Kundin ist die Einlasssonde nicht verbogen Kundin bittet um Service | SWO-162166 | No | Troubleshooting, QC aborted thereby locking all parameters, looking through cal values - all parameters membraned, several cal, QC | Customer has already replaced the inlet seal, performed with customer LS pump scale - ABL continues to drip out of inlet, alut customer is not bent the inlet probe Customer asks for service | Revisit required | HW | Inlet module | Leakage | 2'nd visit
ABL825 FLEX | Fehlersuche Proteinentfernung, Pumpenkalibrierung, LS-Justage Oximentrietest im Serviceprogramm durchgeführt Einlassdichtung erneuert Mehrmals 2-Pkt.-Kalibrierung und THb-Kalibrierung durchgeführt Fehler besteht bei QC-Messung ==> Ersatzteil muss bestellt werden | Fehlersuche Proteinentfernung, Pumpenkalibrierung, LS-Justage Oximentrietest im Serviceprogramm durchgeführt Einlassdichtung erneuert Mehrmals 2-Pkt.-Kalibrierung und THb-Kalibrierung durchgeführt Fehler besteht bei QC-Messung ==> Ersatzteil muss bestellt werden | SWO-162311 | No | Troubleshooting Protein removal, pump calibration, LS-adjustment Oximentrie test carried out in the service program Inlet seal renewed Several 2-pcs calibration and THb calibration performed Error exists with QC measurement ==> Spare part must be ordered | Troubleshooting Protein removal, pump calibration, LS-adjustment Oximentrie test carried out in the service program Inlet seal renewed Several 2-pcs calibration and THb calibration performed Error exists with QC measurement ==> Spare part must be ordered | Cleaned | HW | Fluidic pathway | Fluidic pathway
ABL827 FLEX | Tausch des AU Boards, I-Botton getauscht, Konfiguration, Roma.db und Roma.ini gesichert und nach Tausch wieder eingespielt, System Keys aufgespielt, Software aufgespielt, Update Type ausgeführt, System- und Funktionskontrolle in Ordnung | Tausch des AU Boards, I-Botton getauscht, Konfiguration, Roma.db und Roma.ini gesichert und nach Tausch wieder eingespielt, System Keys aufgespielt, Software aufgespielt, Update Type ausgeführt, System- und Funktionskontrolle in Ordnung | SWO-162425 | No | Exchange of the AU board, I-Botton exchanged, configuration, Roma.db and Roma.ini backed up and replayed after exchange, system keys played up, software played up, Update Type executed, system and function control in order | Exchange of the AU board, I-Botton exchanged, configuration, Roma.db and Roma.ini backed up and replayed after exchange, system keys played up, software played up, Update Type executed, system and function control in order | Defect | HW | Analyzer Control Module | Analyzer control PCB
ABL825 FLEX | Intermittent Oxi hardware problem on QC with error 0579. No errors when arrived. Ran Complete OXI test to isolate. Lamp unit test fails with Neon intensity too low, 34%. need to be >50%. | Removed OXI module and reseat all connections. Remove and inspect Neon bulb. No issues. Perform complete OXI test. Neon # above 50 %, 69%. Ran Calibrations. Performed tHb Calibration in service. All ok. Contact : Shawn Bassett PH: 715-393-2071 | D1FE52AB4E322F7C | No | Intermittent Oxi hardware problem on QC with error 0579. No errors when arrived. Ran Complete OXI test to isolate. Lamp unit test fails with Neon intensity too low, 34%. Need to be > 50%. | Removed OXI module and reseat all connections. Remove and inspect Neon bulb. No issues. Perform complete OXI test. Neon # above 50 %, 69%. Ran Calibrations. Performed tHb Calibration in service. All ok. Contact : Shawn Bassett PH: 715-393-2071 | Defect | HW | Oxi module | Lamp unit | Neon
AQT90 FLEX | replaced the optical unit, now working fine and test are running smoothly. | B | SWO-162277 | No | replaced the optical unit, now working fine and test are running smoothly. |  | Defect | HW | Process Module | Optical Unit | 129
AQT90 FLEX | aligned the barcode scanner and check complete wet section as well as the liquid transport. the needle wash unit got blockage due to which it is showing error , cleaned it properly. now analyser working fine. | aligned the barcode scanner and check complete wet section as well as the liquid transport. the needle wash unit got blockage due to which it is showing error , cleaned it properly. now analyser working fine. | SWO-162273 | No | aligned the barcode scanner and check complete wet section as well as the liquid transport. the needle wash unit got blockage due to which it is showing error , cleaned it properly. now analyser working fine. | aligned the barcode scanner and check complete wet section as well as the liquid transport. the needle wash unit got blockage due to which it is showing error , cleaned it properly. now analyser working fine. | Adjustment | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX BASIC |  | B | SWO-162336 | No
ABL80 FLEX BASIC | Jahreswartung wird vorgezogen (siehe SWO) | Jahreswartung wird vorgezogen (siehe SWO) | SWO-162423 | No | Annual maintenance is brought forward (see SWO) | Annual maintenance is brought forward (see SWO)
ABL90 FLEX | Erneuerung des Solutions pack | Erneuerung des Solutions pack | SWO-162559 | Yes | Renewal of the Solutions pack | Renewal of the Solutions pack | Re classified RMED | ECKHI Email 24.jan-2117
ABL90 FLEX | インレットプローブ及びガスケット交換。内部流路洗浄。各流体系テスト項目チェック。 Cal QC 測定：OK | B | SWO-162587 | No | Inlet probes and gasket replacements. Internal flow path cleaning. Check each flow system test item. Cal QC Measurement: OK |  | Replaced | HW | Inlet module | Probe clot
ABL90 | Blood samples with error 581 Spectrum mismatch. | Replaced hemolyzer Installed a New Pm kit Ran Thb cal | D1FFB319C4479EF0 | No | Blood samples with error 581 Spectrum mismatch. | Replaced hemolyzer Installed a New Pm kit Ran Thb cal | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | In new sensor cassette in Qc level B & C Lac level is above the range.Run fluid transport and take data log for further inspection. | In new sensor cassette in Qc level B & C Lac level is above the range.Run fluid transport and take data log for further inspection. | SWO-162858 | No | In new sensor cassette in Qc level B & C Lac level is above the range. Run fluid transport and take data log for further inspection. | In new sensor cassette in Qc level B & C Lac level is above the range. Run fluid transport and take data log for further inspection. | Assisted Replacement | SC | SC | Reference | 1361
ABL90 FLEX | Geräteaufstellung in OP nach Reparatur in GS17. | Geräteaufstellung in OP nach Reparatur in GS17. | SWO-162691 | No | Equipment installation in op after repair in GS17. | Equipment installation in op after repair in GS17. | Defect | HW | Power Supply | Power Supply
ABL90 | found that it states unable to pump solutions | back flush analyzer and cleaned inlet gasket. | D1FFDC00C4440177 | No | found that it states unable to pump solutions | back flush analyzer and cleaned inlet gasket. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Lösungspack aus GS17 erneuert. | Lösungspack aus GS17 erneuert. | SWO-162692 | No | Solution pack from GS17 renewed. | Solution pack from GS17 renewed. | Assisted Replacement | SP | SP | Leak current
ABL90 | Lot XV-15 SP caused calibration errors driving QC out of range. Instrument had sat in maintenance hold for several days without SC being cycled and needs replacement. | Replaced SC & SP. Instrument calibrated normally. QC is in range and instrument operates normally. | D1FFD13CC46384A5 | No | Lot XV-15 SP caused calibration errors driving QC out of range. Instrument had sat in maintenance hold for several days without SC being cycled and needs replacement. | Replaced SC & SP. Instrument calibrated normally. QC is in range and instrument operates normally. | Assisted Replacement | SP | SP | FT
ABL90 FLEX PLUS | ｲﾝﾚｯﾄ：PCB・ﾓｼﾞｭｰﾙ交換 | B | SWO-162316 | No | Inlet: PCB module replacement |  | Defect | HW | Auto Inlet module | Inlet Position Cal.
ABL90 FLEX PLUS | Customer had followed guided trouble shooting and replaced inlet gasket and solution pack. I replaced inlet gasket, still failed. Took inlet from adjacent analyser, which passed. Tried another gasket, which failed. No more gaskets on site. Travelled to Glenfield to collect a gasket from their spare analyser. Now passing vacuum test but failing cthb on drift. Analyser now fully functional. | Escalated to John for call out in the morning. | SWO-162481 | No | Customer had followed guided trouble shooting and replaced inlet gasket and solution pack. I replaced inlet gasket, still failed. Took inlet from adjacent analyser, which passed. Tried another gasket, which failed. No more gaskets on site. Travelled to Glenfield to collect a gasket from their spare analyser. Now passing vacuum test but failing cthb on drift. Analyser now fully functional. | Escalated to John for call out in the morning. | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D1FFA82BC40D41ED | No |  |  | Defect | HW | Basic Unit | Wetsection cotrol PCB
AQT90 FLEX |  | 光学单元903-364更换后，自检结果：固定偏差和闪光灯强度失败。因此要重新借光单903-364，此单先关单 by andrew | SWO-162385 | No |  | After the optical unit 903-364 is replaced, the self-test results: fixed deviation and flash strength failure. So to re-borrow the single 903-364, this single first close the single by andrew | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX |  | B | SWO-162330 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
ABL837 FLEX | Cleaned dust from unit and test ok. | Cleaned dust from unit and test ok. | SWO-162454 | No | Cleaned dust from unit and test ok. | Cleaned dust from unit and test ok. | Cleaned | HW | Auto Check Module | Barcode Reader
ABL815 FLEX | Ausbau des Autocheck Moduls, auseinander bauen. Motor von Karussell Träger gelöst. Autocheck Karussell mit Antriebsachse verklebt, gelöst und gründlich gereinigt. Autocheck Modul zusammengebaut, Funktion in Ordnung. | Ausbau des Autocheck Moduls, auseinander bauen. Motor von Karussell Träger gelöst. Autocheck Karussell mit Antriebsachse verklebt, gelöst und gründlich gereinigt. Autocheck Modul zusammengebaut, Funktion in Ordnung. | SWO-162563 | No | Expansion of the Autocheck module, disassemble. Motor detached from carousel carrier. Autocheck carousel with drive axle glued, dissolved and thoroughly cleaned. Autocheck module assembled, function fine. | Expansion of the Autocheck module, disassemble. Motor detached from carousel carrier. Autocheck carousel with drive axle glued, dissolved and thoroughly cleaned. Autocheck module assembled, function fine. | Cleaned | HW | Auto Check Module | caroussel
ABL825 FLEX | CLOG REQUIRES REMOVAL AND FLUSHHING | CLEANED NIPPLE AND REPLACED WASTE TUBING | D20165C73A54066A | No | CLOG REQUIRES REMOVAL AND FLUSHHING | CLEANED NIPPLE AND REPLACED WASTE TUBING | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Proteinentfernung durchgeführt - Messkammer gereinigt Lactat Elektrode erneuert Lac Vorverstärkernullpunkt kalibriert Luftweg im Fluidic Modul gespült Ventile im pH BG sowie El.Met. Modul überprüft - ok Gerät mehrfach kalibriert, Drift bei Lactat 0,0 QC Messungen durchgeführt | Proteinentfernung durchgeführt - Messkammer gereinigt Lactat Elektrode erneuert Lac Vorverstärkernullpunkt kalibriert Luftweg im Fluidic Modul gespült Ventile im pH BG sowie El.Met. Modul überprüft - ok Gerät mehrfach kalibriert, Drift bei Lactat 0,0 QC Messungen durchgeführt | SWO-162549 | No | Protein removal performed - measuring chamber cleaned Lactate electrode renewed Lac preamplifier zero point calibrated Airway flushed in fluidic module valves in pH BG and El.Met. Module checked - ok Device calibrated multiple times, drift at lactate 0.0 QC measurements performed | Protein removal performed - measuring chamber cleaned Lactate electrode renewed Lac preamplifier zero point calibrated Airway flushed in fluidic module valves in pH BG and El.Met. Module checked - ok Device calibrated multiple times, drift at lactate 0.0 QC measurements performed | Assisted Replacement | Electrode | Electrode | Lac
ABL815 FLEX | Original Inlet Modul zurück gebaut, Modulventil V1-V4 demontiert und überholt, Funktionskontrolle ok | Original Inlet Modul zurück gebaut, Modulventil V1-V4 demontiert und überholt, Funktionskontrolle ok | SWO-162776 | No | Original Inlet module built back, module valve V1-V4 disassembled and overhauled, function control ok | Original Inlet module built back, module valve V1-V4 disassembled and overhauled, function control ok | Unknown | HW | Inlet module | Leakage
AQT90 FLEX | 代替えのフロントパネルに入れ替え、持ち帰る。 | B | SWO-163128 | No | Replace with an alternate front panel and take it home. |  | Defect | HW | Cover and Chassis | Inlet Door
ABL80 | checked data logs for errors, checked QC log, Red qc upper range is 69 instrument reading 71-74mmHg | removed and replaced solution pack | D20152713A0AFD02 | No | checked data logs for errors, checked QC log, Red qc upper range is 69 instrument reading 71-74mmHg | removed and replaced solution pack
ABL90 FLEX | Solutionpack geliefert und eingesetzt, Kalibration, QC- Kontrolle | Solutionpack geliefert und eingesetzt, Kalibration, QC- Kontrolle | SWO-162705 | Yes | Solutionpack delivered and used, calibration, QC control | Solutionpack delivered and used, calibration, QC control | Re classified RMED | Logistical upSWO
ABL90 | Intermittent Spectrum 581 erros on blood samples. | Replaced the hemolyzer Performed Thb cal. | D201161A3A379EFA | No | Intermittent Spectrum 581 erros on blood samples. | Replaced the hemolyzer Performed Thb cal. | Defect | HW | Oxi module | Hemolyzer
ABL90 | The touch screen was out of cal | I ran the touch screen cal and all is working now | D200D5D43A3913C7 | No | The touch screen was out of cal | I ran the touch screen cal and all is working now | Adjustment | Touch screen | DMS_PC | Touch screen | Cal
ABL90 FLEX PLUS | Nettoyage de la section fluidique e.o Modification du setup e.o Calibration e.o contrôle de qualité e.o | Nettoyage de la section fluidique e.o Modification du setup e.o Calibration e.o contrôle de qualité e.o | SWO-162704 | No | Cleaning the fluidsection e.o Changing the setup e.o Calibration e.o quality control e.o | Cleaning the fluidsection e.o Changing the setup e.o Calibration e.o quality control e.o | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL800 BASIC |  | B | SWO-162843 | No |  |  | Revisit required | HW | Basic unit | Barcode reader, Basic Unit
ABL827 FLEX |  | B | SWO-162589 | No |  |  | Unknown | HW | FlexQ Module | Barcode Reader
ABL825 FLEX | Gerät geprüft, Messkammer geprüft und gereinigt, Funktionskontrolle ok | Gerät geprüft, Messkammer geprüft und gereinigt, Funktionskontrolle ok | SWO-162809 | No | Device tested, measuring chamber tested and cleaned, function control ok | Device tested, measuring chamber tested and cleaned, function control ok | Cleaned | HW | pH/BG Module | Measuring chamber
ABL825 FLEX | pCO2 Wert in AC5+ L3 zwei mal zu hoch gemessen mit zwischenzeitlicher Reinigung Kal 2, Messkammern überprüft und leicht blutige Ablagerungen vor Ref. entfernt. pCO2 Elektrode gereinigt und neu membraniert, Kontakte geputzt. Kalibrationen, AC5+ Messung => OK | pCO2 Wert in AC5+ L3 zwei mal zu hoch gemessen mit zwischenzeitlicher Reinigung Kal 2; Messkammern überprüft und leicht blutige Ablagerungen vor Ref. entfernt. pCO2 Elektrode gereinigt und neu membraniert, Kontakte geputzt. Kalibrationen, AC5+ Messung => OK | SWO-162659 | No | pCO2 value in AC5+ L3 measured twice too high with intermediate cleaning Kal 2; measuring chambers and slightly bloody deposits before ref. away. pCO2 electrode cleaned and re-membraneed, contacts cleaned. Calibrations, AC5+ measurement => OK | pCO2 value in AC5+ L3 measured twice too high with intermediate cleaning Kal 2; measuring chambers and slightly bloody deposits before ref. away. pCO2 electrode cleaned and re-membraneed, contacts cleaned. Calibrations, AC5+ measurement => OK | Cleaned | HW | Fluidic pathway | Fluidic pathway | pCO2
ABL835 FLEX | This happened after a reboot. All reboots after that would not resolve issue. | Back up all data and settings. Tried cleaning up rom.ini. Same issue. Removed Kira and reinstall. Same issue. Had to reinstall Win 7 and all other software. Worked with LIS person on interface. Also disinfect and remove old analyzers as this customer went Live with new analyzers. | D200F80B3A0B4FC6 | No | This happened after a reboot. All reboots after that would not resolve issue. | Back up all data and settings. Tried cleaning up rom.ini. Same issue. Removed Kira and reinstall. Same issue. Had to reinstall Win 7 and all other software. Worked with LIS person on interface. Also disinfect and remove old analyzers as this customer went Live with new analyzers. | Defect | SW | Win 7 | Re-Install
ABL80 FLEX BASIC | 因前一个SWO借件发货有问题 ，重开一个。 | 因前一个SWO借件发货有问题 ，重开一个。 | SWO-163161 | No | Re-open one due to a problem with the previous SWO debit shipment. | Re-open one due to a problem with the previous SWO debit shipment.
ABL80 CO-OX | Telephone call indicated symptoms of failed CF card. Arrived and observed sound at power up, but no display | Using trade in analyzer to isolate and troubleshoot. Isolated cause down to defective CPU. | D2026453B0034415 | No | Telephone call indicated symptoms of failed CF card. Arrived and observed sound at power up, but no display | Using trade in analyzer to isolate and troubleshoot. Isolated cause down to defective CPU.
ABL90 FLEX | Fehlersuche SP ist bereits seit 24 h installiert, mehrere Kal sind gelaufen, Statuswerte zu hoch Reinitialisierung alter SP,keine Besserung Austausch SP, Fehler bleibt, StatuswerteGlu, Lac zu hoch Wechsel SC SP und SC zur Reklamationsmeldungen an RDeu | Fehlersuche SP ist bereits seit 24 h installiert, mehrere Kal sind gelaufen, Statuswerte zu hoch Reinitialisierung alter SP,keine Besserung Austausch SP, Fehler bleibt, StatuswerteGlu, Lac zu hoch Wechsel SC SP und SC zur Reklamationsmeldungen an RDeu | SWO-163031 | No | Troubleshooting SP has been installed since 24 h, several Kal have run, Status values too high Reinitialization of old SP, no improvement Exchange SP, error remains, status valuesGlu, Lac too high Change SC SP and SC to report complaints to RDeu | Troubleshooting SP has been installed since 24 h, several Kal have run, Status values too high Reinitialization of old SP, no improvement Exchange SP, error remains, status valuesGlu, Lac too high Change SC SP and SC to report complaints to RDeu | Defect | SC | SC | Glu/Lac
ABL90 | While customer was replacing SC analyzer appears to have froze, also customer logon failed to work. Suspect customer rushed SC installation and did not wait for prompt for next steps causing analyzer to stall. Also customer tried to insatll an expired SC. | Walked customer through SC installation, pausing for each step to be completed, also had customer install old SC, still had cycles available, appears to be installing properly. Customer will call if problem returns. | D201913FB01C27A4 | No | While customer was replacing SC analyzer appears to have froze, also customer logon failed to work. Suspect customer rushed SC installation and did not wait for prompt for next steps causing analyzer to stall. Also customer tried to insatll an expired SC. | Walked customer through SC installation, pausing for each step to be completed, also had customer install old SC, still had cycles available, appears to be installing properly. Customer will call if problem returns. | User error | SC installation | Expired SC | Expired SC
ABL90 FLEX | Fehlersuche, Einlaßdichtung+ LP wurde bereits ersetzt, Abfallwege gespült - speziell Hemolysator, Einlaß gespült, Servicetests durchgeführt, Vakuumfehler behoben, Neustartphase abgewartet, QC | Fehlersuche, Einlaßdichtung+ LP wurde bereits ersetzt, Abfallwege gespült - speziell Hemolysator, Einlaß gespült, Servicetests durchgeführt, Vakuumfehler behoben, Neustartphase abgewartet, QC | SWO-163083 | No | Troubleshooting, inlet seal+ LP has already been replaced, waste paths flushed - especially hemolysator, inlet flushed, service tests performed, vacuum error fixed, restart phase awaited, QC | Troubleshooting, inlet seal+ LP has already been replaced, waste paths flushed - especially hemolysator, inlet flushed, service tests performed, vacuum error fixed, restart phase awaited, QC | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Apparently this user came in this morning to the analyzer displaying the error which occured overnight. No calls overnight from the overnight shift reporting any issues. | Asked the customer to power down, unplug for 5 Minutes+, then plug back in and attempt to restart. Problem resolved. | D201ED43B04BE665 | Yes | Apparently this user came in this morning to the analyzer displaying the error which occured overnight. No calls overnight from the overnight shift reporting any issues. | Asked the customer to power down, unplug for 5 Minutes+, then plug back in and attempt to restart. Problem resolved. | Re classified RMED | Phone support
ABL90 FLEX | LP geliefert und eingesetzt Startphase System IO | LP geliefert und eingesetzt Startphase System IO | SWO-163010 | Yes | LP supplied and used start-up phase system IO | LP supplied and used start-up phase system IO | Re classified RMED | Logistical upSWO
ABL90 FLEX | LP geliefert | LP geliefert | SWO-163012 | Yes | LP supplied | LP supplied | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Ran complete oximetry tests all passing. Replaced pinch valve tube which solved the issue. Passed calibration and QCs straight away. | Ran complete oximetry tests all passing. Replaced pinch valve tube which solved the issue. Passed calibration and QCs straight away. | SWO-163124 | No | Ran complete oximetry tests all passing. Replaced pinch valve tube which solved the issue. Passed calibration and QCs straight away. | Ran complete oximetry tests all passing. Replaced pinch valve tube which solved the issue. Passed calibration and QCs straight away. | Defect | HW | Oxi module | Pinch valve tube
ABL825 FLEX | Fehlersuche, Gas2 Regelung getestet, Drucktest negativ, Ventil durchgespült, Anschluß erneuert, Drucktest wiederholt - i.O. , Regelung getestet - i.O., Neustart, Kal, QC | Fehlersuche, Gas2 Regelung getestet, Drucktest negativ, Ventil durchgespült, Anschluß erneuert, Drucktest wiederholt - i.O. , Regelung getestet - i.O., Neustart, Kal, QC | SWO-163086 | No | Troubleshooting, Gas2 control tested, pressure test negative, valve flushed through, connection renewed, pressure test repeated - i.O., control tested - i.O., restart, Kal, QC | Troubleshooting, Gas2 control tested, pressure test negative, valve flushed through, connection renewed, pressure test repeated - i.O., control tested - i.O., restart, Kal, QC | Cleaned | HW | Basic Unit | Gas 2, valve
ABL827 FLEX |  | B | D20220F1B0283D08 | No |  |  | Defect | HW | Inlet module | Probe bend
ABL835 FLEX | Fehler aufgetreten nach HJW => neue Einlassdichtung benetzt, ohne Erfolg. Servicemenü aufgerufen und mehrfach Einlass kalibriert, o.E. Nadelantrieb manuell bewegt, o.E. Nadel überprüft, sie ist vollständig gerade und ohne Verschleißmarken. Weitere neue Einlassdichtung eingesetzt, o.E. Einlassnadel ersetzt, nun sofort Einlass-Kal bestanden => erste Einlassdichtung (aus W-Kit) eingesetzt, ebenfalls problemlos. Weiteres Verfahren s. HJW-Case | Fehler aufgetreten nach HJW => neue Einlassdichtung benetzt, ohne Erfolg. Servicemenü aufgerufen und mehrfach Einlass kalibriert, o.E. Nadelantrieb manuell bewegt, o.E. Nadel überprüft, sie ist vollständig gerade und ohne Verschleißmarken. Weitere neue Einlassdichtung eingesetzt, o.E. Einlassnadel ersetzt, nun sofort Einlass-Kal bestanden => erste Einlassdichtung (aus W-Kit) eingesetzt, ebenfalls problemlos. Weiteres Verfahren s. HJW-Case | SWO-163044 | No | Error occurred after HJW => new inlet seal wetted, to no avail. Service menu called and repeatedly calibrated inlet, o.E. needle drive manually moved, o.E. needle checked, it is completely straight and without wear marks. Further new inlet seal inserted, o.E. inlet needle replaced, now immediately inlet-scale passed => first inlet seal (made of W-Kit) inserted, also without problems. Further procedure see HJW-Case | Error occurred after HJW => new inlet seal wetted, to no avail. Service menu called and repeatedly calibrated inlet, o.E. needle drive manually moved, o.E. needle checked, it is completely straight and without wear marks. Further new inlet seal inserted, o.E. inlet needle replaced, now immediately inlet-scale passed => first inlet seal (made of W-Kit) inserted, also without problems. Further procedure see HJW-Case | Defect | HW | Inlet module | Needle
ABL825 FLEX | OXI-Modul ausgetauscht S N entnommenes OXI-Modul: 902-833R2034N002 THb-Kalibrierung | OXI-Modul ausgetauscht S N entnommenes OXI-Modul: 902-833R2034N002 THb-Kalibrierung | SWO-163056 | No | OXI module replaced S N removed OXI module: 902-833R2034N002 THb calibration | OXI module replaced S N removed OXI module: 902-833R2034N002 THb calibration | Defect | HW | Oxi module | Spectrometer | Spectrum mismatch
ABL825 FLEX | Daten überprüft, Messkammer gereinigt und neue Membranen bestellen lassen Gerät kalibriert und mit QC überprüft | Daten überprüft, Messkammer gereinigt und neue Membranen bestellen lassen Gerät kalibriert und mit QC überprüft | SWO-163117 | No | Data checked, measuring chamber cleaned and new membranes ordered Device calibrated and checked with QC | Data checked, measuring chamber cleaned and new membranes ordered Device calibrated and checked with QC | Assisted Replacement | Electrode | Membrane | Lac | Short lifetime
ABL90 FLEX PLUS | Replaced wet section control PCB. Calibrated inlet module. | Replaced wet section control PCB. Calibrated inlet module. | SWO-163259 | No | Replaced wet section control PCB. Calibrated inlet module. | Replaced wet section control PCB. Calibrated inlet module. | Defect | HW
ABL90 FLEX |  | B | SWO-163138 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-163139 | No | The replacement test card is OK. | The replacement test card is OK.
AQT90 FLEX | 更换903-119一个正常。 by andrew | 更换903-119一个正常。 by andrew | SWO-163165 | No | Replace the 903-119 one normal. by andrew | Replace the 903-119 one normal. by andrew | Defect | HW | Process Module | Cup wash unit
AQT90 FLEX | sustituida process Wheel y filtro LS2. Pasados QCs y test de blancos. Se abre techline | B | D203ACB12625FB7A | No | replaced process Wheel and LS2 filter. Past QCs and white tests. Techline opens |  | Defect | HW | Process Module | Process wheel
AQT90 FLEX | Cleaned Process Wheel, Tried to reinstalled Optical Unit, Cleaned Analyzer, Reinstalled Motor, Tried to Resolve, Problem Found with Optical Unit, process wheel and Motor, need spares for Testing, waiting for spares. | B | SWO-163133 | No | Cleaned Process Wheel, Tried to reinstalled Optical Unit, Cleaned Analyzer, Reinstalled Motor, Tried to Resolve, Problem Found with Optical Unit, process wheel and Motor, need spares for Testing, waiting for spares. |  | Revisit required | HW | Process Module | Optical Unit
ABL90 FLEX | ＳＰ交換 | B | SWO-163320 | No | SP Exchange |  | Assisted Replacement | SP | SP | SP
ABL80 FLEX CO-OX | Nullpunktkalibrierung mit Zeroing Kassette durchgeführt. Werte sind innerhalb Grenzwerte nach Manuel. Erneute Kalibration mit Qualitätskontrolle brauchte kein Erfolg. Leihgerät aus Geschäftsstelle mitgenommen zum Test, ob Betriebsmittel im Leihgerät sich gleich verhalten. LP im Anwender Gerät zu wenig Messzyklen. Anwender haben keinen Ersatz LP. Messzyklen im Leihgerät zu wenig um zu tauschen. Leihgerät wieder mitgenommen. Anwender wird den LP tauschen, wenn geliefert und die SC. Anwender meldet sich, wenn weiterhin Qualitätskontroll Grenzverletung vorhanden. | Nullpunktkalibrierung mit Zeroing Kassette durchgeführt. Werte sind innerhalb Grenzwerte nach Manuel. Erneute Kalibration mit Qualitätskontrolle brauchte kein Erfolg. Leihgerät aus Geschäftsstelle mitgenommen zum Test, ob Betriebsmittel im Leihgerät sich gleich verhalten. LP im Anwender Gerät zu wenig Messzyklen. Anwender haben keinen Ersatz LP. Messzyklen im Leihgerät zu wenig um zu tauschen. Leihgerät wieder mitgenommen. Anwender wird den LP tauschen, wenn geliefert und die SC. Anwender meldet sich, wenn weiterhin Qualitätskontroll Grenzverletung vorhanden. | SWO-163400 | No | Zero-point calibration performed with zeroing cassette. Values are within limits according to Manuel. Renewing calibration with quality control did not require success. Loaner taken from the office for testing whether equipment in the rental device behaves the same. LP in the user device too few measuring cycles. Users do not have a replacement LP. measuring cycles in the rental device too little to exchange. borrower. User will exchange the LP when delivered and the SC. User reports if quality control border vulnerability is still present. | Zero-point calibration performed with zeroing cassette. Values are within limits according to Manuel. Renewing calibration with quality control did not require success. Loaner taken from the office for testing whether equipment in the rental device behaves the same. LP in the user device too few measuring cycles. Users do not have a replacement LP. measuring cycles in the rental device too little to exchange. borrower. User will exchange the LP when delivered and the SC. User reports if quality control border vulnerability is still present.
ABL90 |  | B | D207299B88293C0D | Yes |  |  | Assisted Replacement | SP | SP | Leak current
ABL90 | 1177 flow selector errors. Customer replaced solution pack and got the same result. Got a couple of 1177 errors while checking after maintenance performed. Unit was moved and had discharged over the week end. | Replaced the flow selector detector PCB. Flushed analyzer. Cycled through fluid transport service programs. | D2062B9612223188 | Yes | 1177 flow selector errors. Customer replaced solution pack and got the same result. Got a couple of 1177 errors while checking after maintenance performed. Unit was moved and had discharged over the week end. | Replaced the flow selector detector PCB. Flushed analyzer. Cycled through fluid transport service programs. | Defect | HW | Basic Unit | PCB, Flow Selector
ABL90 FLEX PLUS |  | B | D207154A881122BF | Yes |  |  | Re classified RMED | Logistical upSWO
Radiance | Move Radiance PC from Server Room to the Lab | Move Radiance PC from Server Room to the Lab | SWO-166777 | Yes | Move Radiance PC from Server Room to the Lab | Move Radiance PC from Server Room to the Lab
ABL827 FLEX | sample errors (0328) and leaking at inlet capillary. | Replaced 2 inlet gaskets. Staff replaced inlet probe. Flushed all lines and flow paths. Ran liquid sensor cals and inlet cals. Replaced membrane valves in pH BG module. Replaced pump tubes (electrode). | D20665BD1227A4F3 | No | sample errors (0328) and leaking at inlet capillary. | Replaced 2 inlet gaskets. Staff replaced inlet probe. Flushed all lines and flow paths. Ran liquid sensor cals and inlet cals. Replaced membrane valves in pH BG module. Replaced pump tubes (electrode). | Defect | HW | Inlet module | Inlet Gasket
ABL815 FLEX | Monitor erneuert. Kalibriert und Töne konfiguriert. | Monitor erneuert. Kalibriert und Töne konfiguriert. | SWO-163367 | No | Monitor renewed. Calibrated and configured totones. | Monitor renewed. Calibrated and configured totones. | Defect | HW | DMS_PC | Touch screen
ABL827 FLEX | AutoCheck unit errors intermittently during auto QC runs. Found breaker arm not returning to full return position. | Reposition breaker motor assembly. Adjusted breaker arm position and carousal position in diagnostics also. Ran all checks. | D20648A61202ABB4 | No | AutoCheck unit errors intermittently during auto QC runs. Found breaker arm not returning to full return position. | Reposition breaker motor assembly. Adjusted breaker arm position and carousal position in diagnostics also. Ran all checks. | Adjustment | HW | Auto Check Module | Breaker Motor
AQT90 FLEX | Replaced Optical Unit with New One, Cleaned Process wheel, Reinstalled Process Unit, Problem Found In Stepper Motor while Movement for the Optical Process, Ordered Stepper Motor for Replacement. | Replaced Optical Unit with New One, Cleaned Process wheel, Reinstalled Process Unit, Problem Found In Stepper Motor while Movement for the Optical Process, Ordered Stepper Motor for Replacement. | SWO-164636 | No | Replaced Optical Unit with New One, Cleaned Process wheel, Reinstalled Process Unit, Problem Found In Stepper Motor while Movement for the Optical Process, Ordered Stepper Motor for Replacement. | Replaced Optical Unit with New One, Cleaned Process wheel, Reinstalled Process Unit, Problem Found In Stepper Motor while Movement for the Optical Process, Ordered Stepper Motor for Replacement. | Defect | HW | Process Module | Optical Unit | 131
ABL80 FLEX | Over flow check. Waste manifold line leak. Clean & check. I will replace it the next time a problem. Data log clean. | Over flow check. Waste manifold line leak. Clean & check. I will replace it the next time a problem. Data log clean. | SWO-163489 | No | Over flow check. Waste manifold line leak. Clean & check. I will replace it the next time a problem. Data log clean. | Over flow check. Waste manifold line leak. Clean & check. I will replace it the next time a problem. Data log clean.
ABL80 CO-OX | Troubleshot to bad CPU which was ordered | Replaced CPU assy. Setup analyzer to dept specifications, including ensuring connection to LIS establish and results passing. | D2078D0E885CC6A9 | No | Troubleshot to bad CPU which was ordered | Replaced CPU assy. Setup analyzer to dept specifications, including ensuring connection to LIS establish and results passing.
ABL9 | alle Daten auf ein USB Stick verschoben und Neustart durch geführt, System wieder messbereit Daten Ingo Sprenger zur Kontrolle geschickt | alle Daten auf ein USB Stick verschoben und Neustart durch geführt, System wieder messbereit Daten Ingo Sprenger zur Kontrolle geschickt | SWO-163691 | No | all data moved to a USB stick and rebooted, system ready to be measured again Data Sent to Ingo Sprenger for control | all data moved to a USB stick and rebooted, system ready to be measured again Data Sent to Ingo Sprenger for control
ABL90 FLEX | Sensor cassette replace. Gasket module clean. Solution pack connecter clean. Calibration & Control check. | Sensor cassette replace. Gasket module clean. Solution pack connecter clean. Calibration & Control check. | SWO-163490 | No | Sensor cassette replace. Gasket module clean. Solution pack connecter clean. Calibration & Control check. | Sensor cassette replace. Gasket module clean. Solution pack connecter clean. Calibration & Control check. | Assisted Replacement | SC | SC | SC
ABL90 FLEX PLUS | LP ersetzt - QC i.O. ein LP YB-50 XL zurückgenommen, da er sehr wahrscheinlich nicht funktionieren wird | LP ersetzt - QC i.O. ein LP YB-50 XL zurückgenommen, da er sehr wahrscheinlich nicht funktionieren wird | SWO-163607 | No | LP replaced - QC in the world an LP YB-50 XL is taken back as it will most likely not work | LP replaced - QC in the world an LP YB-50 XL is taken back as it will most likely not work | Assisted Replacement | SP | SP | pCO2 | YB5
ABL90 FLEX PLUS |  | B | D208A5ABFE12AD42 | No |  |  | Defect | HW | Auto Inlet module | Auto Inlet module
ABL90 FLEX PLUS | It looks like the wet section is not communicating with PC module, I will return with a new PC Module, Power Supply & Wet Section Control PCB to see which part is at fault. | B | SWO-163706 | No | It looks like the wet section is not communicating with PC module, I will return with a new PC Module, Power Supply & Wet Section Control PCB to see which part is at fault. |  | Revisit required | SW | Offline | Offline
ABL90 | CHEQUEO DE CONEXIONES Y CONFIGURACIÓN DEL EQUIPO. TRAS ABRIR INCIDENCIA EN INFORMÁTICA DE HOSPITAL, SE SOLUCIONA JUNTO CON PERSONAL DE IT | RECONFIGURACIÓN POR PARTE DE IT DE LA WLAN | D20753C988168D95 | No | CHECK OF CONNECTIONS AND EQUIPMENT SETTINGS. AFTER OPENING INCIDENT IN HOSPITAL INFORMATION TECHNOLOGY, IT IS SOLVED TOGETHER WITH IT STAFF | IT RECONFIGURATION OF THE WLAN | Setup | His/lis | His/lis | His/lis
ABL800 FLEX | replace the crea reagent. calibrate system multiple time.calibration pass & qc crea pass. | replace the crea reagent. calibrate system multiple time.calibration pass & qc crea pass. | SWO-164608 | No | replace the crea reagent. calibrate system multiple time.calibration pass & qc crea pass. | replace the crea reagent. calibrate system multiple time.calibration pass & qc crea pass. | Assisted Replacement | Consumables | Solution bottle | Crea
ABL810 FLEX | Comprobación del estado del analizador, telefónicamente cliente informa que el electrodo de pH esta roto. Se envia electrodo nuevo para su reposición. | B | D20EB7664C1479F9 | Yes | Checking the status of the analyzer, telephone client informs that the pH electrode is broken. New electrode is sent for replacement. |  | Re classified RMED | Telph support
ABL815 FLEX |  | Kunde bittet um einen erneuten Service. | SWO-163676 | Yes |  | Customer asks for a new service. | Re classified RMED | Telph support
ABL800 BASIC | オキシモジュール交換 | B | SWO-163746 | No | Oxy module replacement |  | Defect | HW | Oxi module | Lamp unit | Neon low
ABL825 FLEX | Ohne äußeren Einfluss vier QC Messungen mit OSM => Proteinentfernung, Kal und tHb-Kal => OSM. OxiCompleteTest im Servicemenü, alle Werte nahezu mittig und OK. Mehrere Kalibrationen und erneute tHb-Kal, OK. Endkontrolle: 3 AC5+ Messungen => Alle OK | Ohne äußeren Einfluss vier QC Messungen mit OSM => Proteinentfernung, Kal und tHb-Kal => OSM. OxiCompleteTest im Servicemenü, alle Werte nahezu mittig und OK. Mehrere Kalibrationen und erneute tHb-Kal, OK. Endkontrolle: 3 AC5+ Messungen => Alle OK | SWO-163789 | No | Without external influence four QC measurements with OSM => protein removal, kal and tHb-kal => OSM. OxiCompleteTest in the service menu, all values almost central and OK. Multiple calibrations and re-calibration tHb-Kal, OK. Final control: 3 AC5+ measurements => All OK | Without external influence four QC measurements with OSM => protein removal, kal and tHb-kal => OSM. OxiCompleteTest in the service menu, all values almost central and OK. Multiple calibrations and re-calibration tHb-Kal, OK. Final control: 3 AC5+ measurements => All OK | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | tHb
ABL825 FLEX | Reinigung, Proteinentfernung, 1.Kal. und tHb Kal. durchgeführt. tHb Test und QCs i.O. | Reinigung, Proteinentfernung, 1.Kal. und tHb Kal. durchgeführt. tHb Test und QCs i.O. | SWO-163689 | No | Reinigung, Proteinentfernung, 1.Kal. und tHb Kal. durchgeführt. tHb Test und QCs i.O. | Reinigung, Proteinentfernung, 1.Kal. und tHb Kal. durchgeführt. tHb Test und QCs i.O. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | tHb
ABL825 FLEX | Neon lamp intensity and neon voltage are out of range. | Replaced lamp unit per Radiometer specifications. Calibrated and ran QC. Instrument operates normally. | D20747F0884663D9 | No | Neon lamp intensity and neon voltage are out of range. | Replaced lamp unit per Radiometer specifications. Calibrated and ran QC. Instrument operates normally. | Defect | HW | Oxi module | Lamp unit
ABL835 FLEX | pH BG modul PCB skiftet | pH BG modul PCB skiftet | SWO-163289 | No | pH BG modul PCB skiftet | pH BG modul PCB skiftet | Defect | HW | pH/BG Module | PCB, ph/BG Module | Afledt fejl på Win
ABL80 CO-OX | sustitución de hemolizador completo, calibración de Cthb, prueba de varias muestras analizadas seguidas y comprobando la repetibilidad. | Sustitución hemolizador | D2088C7CFE372358 | No | complete hemolizer replacement, Cthb calibration, testing of several samples analyzed in a row and checking repeatability. | Hemolizer substitution
ABL90 FLEX |  | Kundin hat das Lösungspack erneuert und hat weiterhin Probleme mit dem Gerät. Kundin wünscht einen Service Einsatz. | SWO-164038 | No |  | Customer has renewed the solution pack and still has problems with the device. Customer wishes a service commitment. | Replaced | HW | Inlet module | Probe bend
ABL90 FLEX | Anwender hat LP bereits erneuert. Startphase im Serviceprogramm wiederholt, Funktion OK. Initialisierung durchgeführt, OK. | Anwender hat LP bereits erneuert. Startphase im Serviceprogramm wiederholt, Funktion OK. Initialisierung durchgeführt, OK. | SWO-164073 | No | User has already renewed LP. Repeated start-up phase in service program, function OK. Initialization performed, OK. | User has already renewed LP. Repeated start-up phase in service program, function OK. Initialization performed, OK. | No fault found | SP | SP | Vaccum error
ABL90 FLEX | LC IT connection set up. Local IT changed IT settings on system. Static IP address assigned to ABL90. Set up and tested. Connection set up and working. | LC IT connection set up. Local IT changed IT settings on system. Static IP address assigned to ABL90. Set up and tested. Connection set up and working. | SWO-166659 | No | LC IT connection set up. Local IT changed IT settings on system. Static IP address assigned to ABL90. Set up and tested. Connection set up and working. | LC IT connection set up. Local IT changed IT settings on system. Static IP address assigned to ABL90. Set up and tested. Connection set up and working. | Setup | His/lis | His/lis | His/lis
ABL90 FLEX | Replaced the hemolyser and ran a tHb calibration. Also while on site I cleaned up the old ABL827s from ED Resus & EAU and then took them back to Radiometer. | asked users to run THB and call back if issues continue no callback, case closed. 26 1 | SWO-163981 | No | Replaced the hemolyser and ran a tHb calibration. Also while on site I cleaned up the old ABL827s from ED Resus & EAU and then took them back to Radiometer. | asked users to run THB and call back if issues continue no callback, case closed. 26 1 | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX PLUS | Changed power cables and moved plug from UPS socket to normal power socket. | B | SWO-164196 | No | Changed power cables and moved plug from UPS socket to normal power socket. |  | Unknown | HW | SP | Leak current
ABL90 FLEX PLUS | Gerät geprüft. Wegen inkonsistente Datenbank Fehler Neustart durchgeführt. Angelieferten Lösungspack nach Startphase installiert. Neuer Lösungspack bringt nun pCO2 Kalfehler! Diverse Kal und QC durchgeführt. | Gerät geprüft. Wegen inkonsistente Datenbank Fehler Neustart durchgeführt. Angelieferten Lösungspack nach Startphase installiert. Neuer Lösungspack bringt nun pCO2 Kalfehler! Diverse Kal und QC durchgeführt. | SWO-164072 | No | device checked. Restart due to inconsistent database error. Delivered solution pack after start-up phase. New solution pack now brings pCO2 cal error! Various Kal and QC performed. | device checked. Restart due to inconsistent database error. Delivered solution pack after start-up phase. New solution pack now brings pCO2 cal error! Various Kal and QC performed. | Assisted Replacement | SP | SP | pCO2
ABL90 | SE SUSTITUYEN DIFERENTES PIEZAS Y TUBOS Y REACTIVOS DE DIFERENTES LOTES, PARA SUBSANAR ERRORES NO SIENDO EFECTIVIO. DESMONTAJE DE MODULO DE COOXIMETRIA PARA LIMPIEZA DE VALVULAS DE GANCHO Y HEMOLIZADOR, EN EL QUYE SE DETECTA UNA OBSTRUCCION, SE VUELVEN A REALIZAR PRUEBAS SIENDO CORRECTAS. CALIBRADO Y VERIFICADO CON Q.C. | OBSTRUCION HEMOLIZADOR | D208BE27FE4833B1 | No | DIFFERENT PARTS AND TUBES AND REACTIVES OF DIFFERENT LOTS ARE SUBSTITUTE, TO CHANGE ERRORS NOT BEING EFFECTIVE. DISASSEMBLY OF COOXIMETRIA MODULE FOR CLEANING OF GANCHO AND HEMOLIZER VALVE, IN WHICH AN OBSTRUCTION IS DETECTED, TESTS ARE RETURNED BEING CORRECT. CALIBRATED AND VERIFIED WITH Q.C. | OBSTRUCION HEMOLIZADOR | Cleaned | Fluidic pathway | Oxi module | Hemolyzer
ABL835 FLEX | Examined display assembly. Verified IP settings for NetOpt remote support connection. | Performed display calibration. Worked with IT systems analyst in setting correct IP settings for remote connectivity. Verified operation of ABL835. Unit fully operational. | D20AFCC7EA5D2C4A | No | Examined display assembly. Verified IP settings for NetOpt remote support connection. | Performed display calibration. Worked with IT systems analyst in setting correct IP settings for remote connectivity. Verified operation of ABL835. Unit fully operational. | Defect | HW | DMS_PC | Touch screen
ABL827 FLEX | Sample errors (328) and leaking at capillary | Replaced inlet gasket and probe . Replace y piece for tubings. Replaced pump tubings (clear from waste bottle and inlet contacts. Leaking stopped after clear tubings replaced. | D208E590FE46C1B5 | No | Sample errors (328) and leaking at capillary | Replaced inlet gasket and probe . Replace y piece for tubings. Replaced pump tubings (clear from waste bottle and inlet contacts. Leaking stopped after clear tubings replaced. | Defect | HW | Inlet module | Inlet Gasket
ABL800 BASIC | Found problem with pH BG module air Flat membrane problem, replaced the same restarted instrument calibrated results are OK, performed the sample found results are OK found instrument is functional. | Found problem with pH BG module air Flat membrane problem, replaced the same restarted instrument calibrated results are OK, performed the sample found results are OK found instrument is functional. | SWO-164729 | No | Found problem with pH BG module air Flat membrane problem, replaced the same restarted instrument calibrated results are OK, performed the sample found results are OK found instrument is functional. | Found problem with pH BG module air Flat membrane problem, replaced the same restarted instrument calibrated results are OK, performed the sample found results are OK found instrument is functional. | Defect | HW | pH/BG Module | Valve, air
ABL810 FLEX | z.Zt. Kein Fehler feststellbar, Prüfung Festplattenkapazität LW C: , Prüfung Konfiguration, Kontrolle Schnittstelle, Kontrolle Firewall, Reinigung Bildschirm, Probelauf Hinweis, dass bitte keine alkoholischen Desinfektions- und Reinigungsmittel für den Bildschirm verwendet werden !!! | z.Zt. Kein Fehler feststellbar, Prüfung Festplattenkapazität LW C: , Prüfung Konfiguration, Kontrolle Schnittstelle, Kontrolle Firewall, Reinigung Bildschirm, Probelauf Hinweis, dass bitte keine alkoholischen Desinfektions- und Reinigungsmittel für den Bildschirm verwendet werden !!! | SWO-163812 | No | Currently. No error detectable, checking disk capacity LW C: , check configuration, control interface, control firewall, cleaning screen, test run Please note that no alcoholic disinfectants and cleaning agents are used for the screen !!! | Currently. No error detectable, checking disk capacity LW C: , check configuration, control interface, control firewall, cleaning screen, test run Please note that no alcoholic disinfectants and cleaning agents are used for the screen !!! | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Nach Rücksprache mit RDEU und RMED den pH Dichtring ersetzt durch 834-618 (834-682). pH Messkammer sowie Elektrode gereinigt, Kalibration, Endkontrolle: AC5+ L3 Messung => Alle Werte OK, pCO2 nahezu mittig. Gerät bzw. L3 Messungen werden weiter beobachtet | Nach Rücksprache mit RDEU und RMED den pH Dichtring ersetzt durch 834-618 (834-682). pH Messkammer sowie Elektrode gereinigt, Kalibration, Endkontrolle: AC5+ L3 Messung => Alle Werte OK, pCO2 nahezu mittig. Gerät bzw. L3 Messungen werden weiter beobachtet | SWO-163921 | No | After consultation with RDEU and RMED, the pH sealing ring is replaced by 834-618 (834-682). pH measuring chamber and electrode cleaned, calibration, final inspection: AC5+ L3 Measurement => All values OK, pCO2 almost central. Device or L3 measurements are further observed | After consultation with RDEU and RMED, the pH sealing ring is replaced by 834-618 (834-682). pH measuring chamber and electrode cleaned, calibration, final inspection: AC5+ L3 Measurement => All values OK, pCO2 almost central. Device or L3 measurements are further observed | Defect | Electrode | Electrode | pH | Sealing ring
ABL825 FLEX | Erneut Probleme mit eingefrorenem Touchscreen laut Anwender, Herrn Otten, Monitoreingaben seien nicht mehr möglich. Beim Eintreffen aktuell keine Probleme feststellbar => neuen Monitor mit geänderter Firmware bestellt, wird nach Lieferung ersetzt. | Erneut Probleme mit eingefrorenem Touchscreen laut Anwender, Herrn Otten, Monitoreingaben seien nicht mehr möglich. Beim Eintreffen aktuell keine Probleme feststellbar => neuen Monitor mit geänderter Firmware bestellt; wird nach Lieferung ersetzt. | SWO-163978 | No | Once again, according to the user, Mr Otten, problems with a frozen touchscreen are no longer possible. No problems can be detected on arrival at the moment => new monitor with changed firmware ordered; will be replaced after delivery. | Once again, according to the user, Mr Otten, problems with a frozen touchscreen are no longer possible. No problems can be detected on arrival at the moment => new monitor with changed firmware ordered; will be replaced after delivery. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Einlassnadel nach oben geknickt => im Rahmen der Tätigkeiten der FAN915-318 gerichtet. | Einlassnadel nach oben geknickt => im Rahmen der Tätigkeiten der FAN915-318 gerichtet. | SWO-163989 | No | Inlet pin bent upwards => within the scope of the activities of fan915-318. | Inlet pin bent upwards => within the scope of the activities of fan915-318. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Fails OXI test for neon lamp | Replaced OXI PCB, complete alignment and calibration. Now passes OXI test | D2089D85FE36E966 | No | Fails OXI test for neon lamp | Replaced OXYGEN PCB, complete alignment and calibration. Now passes OXYGEN test | Defect | HW | Oxi module | PCB, Oxi
AQT90 FLEX | IP Adressen korrigiert. Speicherbefehl mit Batch Datei ausgeführt. Datenübertragung iO. | IP Adressen korrigiert. Speicherbefehl mit Batch Datei ausgeführt. Datenübertragung iO. | SWO-163950 | No | IP addresses corrected. Memory command executed with batch file. Data transfer iO. | IP addresses corrected. Memory command executed with batch file. Data transfer iO. | Setup | SW | His/lis | His/lis
AQT90 FLEX |  | B | D20B3ADCEA3DE2E8 | No |  |  | Defect | Consumables | Test Kit | Cartridge | Broken, nose
ABL80 FLEX CO-OX | Beim Eintreffen nur pH und pCO2 grün => Rest rot => Kal und QC aber i.O. => bis auf pH und pCO2 alle Parameter deaktiviert Weitere Überprüfung und Reparatur im Rahmen der fälligen Jahreswartung | Beim Eintreffen nur pH und pCO2 grün => Rest rot => Kal und QC aber i.O. => bis auf pH und pCO2 alle Parameter deaktiviert Weitere Überprüfung und Reparatur im Rahmen der fälligen Jahreswartung | SWO-164218 | No | On arrival only pH and pCO2 green => rest red => Kal and QC but i.O. => all parameters deactivated except for pH and pCO2 Further inspection and repair as part of the due annual maintenance | On arrival only pH and pCO2 green => rest red => Kal and QC but i.O. => all parameters deactivated except for pH and pCO2 Further inspection and repair as part of the due annual maintenance
ABL90 FLEX | Hemolyzer erneuert. | Hemolyzer erneuert. | SWO-164375 | No | Hemolyzer. | Hemolyzer. | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | Gerät im Service überprüft, Kunde hat in den letzten Tagen 2x SC90 und einmal LP90 getauscht, versucht LP90 an einem anderen ABL einzusetzen - ging nach Rücksprache Arzt nicht, Tausch des LP90, nach Startphase und QC alles wieder grün, System- und Funktionskontrolle in Ordnung | Gerät im Service überprüft, Kunde hat in den letzten Tagen 2x SC90 und einmal LP90 getauscht, versucht LP90 an einem anderen ABL einzusetzen - ging nach Rücksprache Arzt nicht, Tausch des LP90, nach Startphase und QC alles wieder grün, System- und Funktionskontrolle in Ordnung | SWO-164306 | No | Device checked in service, customer has exchanged 2x SC90 and once LP90 in the last days, tried to use LP90 on another ABL - did not go after consultation doctor, exchange of the LP90, after start phase and QC everything back green, system and function control in order | Device checked in service, customer has exchanged 2x SC90 and once LP90 in the last days, tried to use LP90 on another ABL - did not go after consultation doctor, exchange of the LP90, after start phase and QC everything back green, system and function control in order | Assisted Replacement | SP | SP | pCO2 and pO2
ABL90 |  | B | D20A091B74130AF6 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL90 | found that the flow value is coming up at 16000. | replaced the wet section control PCB, Sensor Interface PCB, and sensor cassette bay unit | D209F36974424041 | No | found that the flow value is coming up at 16000. | replaced the wet section control PCB, Sensor Interface PCB, and sensor cassette bay unit | Defect | HW | Basic Unit | Wetsection control PCB
ABL90 FLEX PLUS | Nieuwe powersupply 910-285R0022N469 gemonteerd. | Nieuwe powersupply 910-285R0022N469 gemonteerd. | SWO-164096 | No | New Powersupply 910-285R0022N469 mounted. | New Powersupply 910-285R0022N469 mounted. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX PLUS | As I was unable to get the analyser up and running it was decided to take this analyser back to Radiometer and return with a new one to replace it. | As I was unable to get the analyser up and running it was decided to take this analyser back to Radiometer and return with a new one to replace it. | SWO-164367 | No | As I was unable to get the analyser up and running it was decided to take this analyser back to Radiometer and return with a new one to replace it. | As I was unable to get the analyser up and running it was decided to take this analyser back to Radiometer and return with a new one to replace it. | Revisit required | analyzer removed | analyzer removed | analyzer removed
ABL90 | Limpieza del sistema con agua. Chequeo por servicio. Sustitucion de sp Cal, qc | B | D209E960742CCC70 | No | Cleaning the system with water. Check for service. Substitution of sp Cal, qc |  | Assisted Replacement | SP | SP | Leak current
AQT90 FLEX | 更换备件903-119后修复 by jesson | 更换备件903-119后修复 by jesson | SWO-164350 | No | Repair by jesson after replacement of spare parts 903-119 | Repair by jesson after replacement of spare parts 903-119 | Defect | HW | Process Module | Cup wash unit
ABL810 FLEX | Change Lactate and PCO2 membrane. | Change Lactate and PCO2 membrane. | SWO-164725 | No | Change Lactate and PCO2 membrane. | Change Lactate and PCO2 membrane. | Assisted Replacement | Electrode | Membrane | Lac | and pCO2
ABL837 FLEX | Replaced printer head and test ok. | Replaced printer head and test ok. | SWO-164201 | No | Replaced printer head and test ok. | Replaced printer head and test ok. | Defect | HW | Printer Module | Printer head
ABL800 BASIC | Removed the Auto check module, Cleaned the analyzer and Scanner restarted instrument found it was scanning the all the ampules found instrument is functional, | Removed the Auto check module, Cleaned the analyzer and Scanner restarted instrument found it was scanning the all the ampules found instrument is functional, | SWO-164730 | No | Removed the Auto check module, Cleaned the analyzer and Scanner restarted instrument found it was scanning the all the ampules found instrument is functional, | Removed the Auto check module, Cleaned the analyzer and Scanner restarted instrument found it was scanning the all the ampules found instrument is functional, | Cleaned | HW | Auto Check Module | Barcode reader, Autocheck Module
ABL835 FLEX |  | Bei Ankunft Bildschirm blau mit weißer Schrift. CPU Unit getauscht danach war Fehler verschwunden und ABL startet wieder ganz normal. MAC Adresse Herrn Luttech von der IT mitgeteilt und auf der Rückseite des ABLs angebracht. FlexQ Modul getauscht und kalibriert, da Kunde bemängelte dass die Sampler nur sporadisch eingelesen werden. Trotz Kalibrierung konnte das Problem nicht behoben werden. Frau Narozny hat mich gebeten die LAC und GLU Elektrode zu membranieren da diese fällig waren. Ventile aus ELE MET und PH BG Modul ausgebaut weil Staub im Inneren der Membrane zu sehen war. Mehrere Kalibrierungen durchgeführt. QCs für alle Level durchgeführt. --> i.O. Frau Narozny hat mit mitgeteilt das die QCs im Labor angekommen sind und i.O. sind. Somit ist der ABL auch an das Cobas System angebunden. ABL wieder messbereit | SWO-164144 | No |  | On arrival screen blue with white font. CPU Unit changed after then error had disappeared and ABL starts back to normal. MAC address informed Mr Luttech of the IT and attached to the back of the ABL. FlexQ module is exchanged and calibrated, as customer complained that the samplers are only read in sporadically. Despite calibration, the problem could not be resolved. Mrs Narozny asked me to membrane the LAC and GLU electrode because they were due. Valves removed from ELE MET and PH BG module because dust could be seen inside the membrane. Multiple calibrations performed. QCs for all levels. --> i.O. Ms Narozny said the QCs arrived at the lab and are in the running. Thus, the ABL is also connected to the Cobas system. ABL ready to be measured again | Defect | HW | DMS_PC | CPU
ABL825 FLEX |  | B | D209B1C1745A0A0E | No |  |  | Adjustment | HW | Oxi module | Lamp unit
ABL815 FLEX | Mehrere Proteinentfernungen, Reinigung, Kal sowie OxiCompleteTest durchgeführt, AC5+ Messung OK, erste Blutmessung erneut mit Fehler 581 => OxiModul ersetzt, Kal, OCT, tHb-Kal, AC5+ Messung, Blutmessung => weiterhin 581 => Originalmodul wieder eingebaut, abgeglichen, kalibriert, 3 Level QC kontrolliert, alle OK. | Mehrere Proteinentfernungen, Reinigung, Kal sowie OxiCompleteTest durchgeführt, AC5+ Messung OK, erste Blutmessung erneut mit Fehler 581 => OxiModul ersetzt, Kal, OCT, tHb-Kal, AC5+ Messung, Blutmessung => weiterhin 581 => Originalmodul wieder eingebaut, abgeglichen, kalibriert, 3 Level QC kontrolliert, alle OK. | SWO-164296 | No | Several protein removals, cleaning, cal and OxiCompleteTest performed, AC5+ measurement OK, first blood measurement replaced again with error 581 => OxiModul, Kal, OCT, tHb-Kal, AC5+ measurement, blood measurement => still 581 => original module re-installed, matched, calibrated, 3 level QC controlled, all OK. | Several protein removals, cleaning, cal and OxiCompleteTest performed, AC5+ measurement OK, first blood measurement replaced again with error 581 => OxiModul, Kal, OCT, tHb-Kal, AC5+ measurement, blood measurement => still 581 => original module re-installed, matched, calibrated, 3 level QC controlled, all OK. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Erneuerungsintervalle für Membranen auf Anweisung von TSP (BorWe) angelegt Glu, Lac und Ref. Membran monatlich alle anderen 3 monatlich | Erneuerungsintervalle für Membranen auf Anweisung von TSP (BorWe) angelegt Glu, Lac und Ref. Membran monatlich alle anderen 3 monatlich | SWO-164297 | No | Renewal intervals for membranes created on the instruction of TSP (BorWe) Glu, Lac and Ref. Membrane monthly all other 3 monthly | Renewal intervals for membranes created on the instruction of TSP (BorWe) Glu, Lac and Ref. Membrane monthly all other 3 monthly | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Display wurde durch René ersetzt wie abgemacht (Folgereparatur SWO-154647) | Display wurde durch René ersetzt wie abgemacht (Folgereparatur SWO-154647) | SWO-164089 | No | Display was replaced by René as done (Follow-up repair SWO-154647) | Display was replaced by René as done (Follow-up repair SWO-154647) | Defect | HW | DMS_PC | Touch screen
ABL837 FLEX |  | B | D209D9A87406AC90 | No |  |  | Defect | HW | Basic Unit | Analyzer control PCB
ABL820 FLEX | Confirmed problem. | Removed a clot from the waste line. | D209C131742D550C | No | Confirmed problem. | Removed a clot from the waste line. | Cleaned | Fluidic pathway | Fluidic manifold | Y-piece | Clot/Leakage
AQT90 FLEX |  | B | D209E0577405B204 | Yes
AQT90 FLEX | Ausbau der alten Nadel und Einbau einer neuen Nadel. Überprüfung der Beweglichkeit der Nadel im Needleblock, Funktion in Ordnung. Qualitätskontroll Messungen durchgeführt. | Ausbau der alten Nadel und Einbau einer neuen Nadel. Überprüfung der Beweglichkeit der Nadel im Needleblock, Funktion in Ordnung. Qualitätskontroll Messungen durchgeführt. | SWO-164322 | Yes | Removal of the old needle and installation of a new needle. Checking the mobility of the needle in the needle block, function in order. Quality control measurements carried out. | Removal of the old needle and installation of a new needle. Checking the mobility of the needle in the needle block, function in order. Quality control measurements carried out. | Re classified RMED
AQT90 FLEX | AQT zur Zeit betriebsbereit, SD gesichert. Abfallbehälter XV-58 | AQT zur Zeit betriebsbereit, SD gesichert. Abfallbehälter XV-58 | SWO-164365 | No | AQT currently operational, SD secured. Waste container XV-58 | AQT currently operational, SD secured. Waste container XV-58 | Defect | Consumables | SP | Cup Stack | 167
AQT90 FLEX | Verbindung Barcodescanner überprüft. Geräteneustart. Scannerfunktion wieder i.O. | Verbindung Barcodescanner überprüft. Geräteneustart. Scannerfunktion wieder i.O. | SWO-164136 | No | Connection barcode scanner checked. Device restart. Scanner function again in the | Connection barcode scanner checked. Device restart. Scanner function again in the | Defect | HW | Aspiration Module | Barcode reader, Inlet
ABL80 FLEX CO-OX | Überprüfung SC, SP beides wurde durch Kunden bereits erneuert, SP reinitialisiert nach Einlaufprozedur QC und Kal i.O. | Überprüfung SC, SP beides wurde durch Kunden bereits erneuert, SP reinitialisiert nach Einlaufprozedur QC und Kal i.O. | SWO-164398 | No | Review SC, SP both have already been renewed by customers, SP reinitialized after the inlet procedure QC and Kal i.O. | Review SC, SP both have already been renewed by customers, SP reinitialized after the inlet procedure QC and Kal i.O.
ABL80 OSM | Troubleshot to a bad CF card | I only had a CF card for CPU 2743. Ordered CF card for 945 CPU. | D20EB4F04C6210D2 | No | Troubleshot to a bad CF card | I only had a CF card for CPU 2743. Ordered CF card for 945 CPU.
ABL90 FLEX | Cleaned sensor contact pins. | Cleaned sensor contact pins. | SWO-164917 | No | Cleaned sensor contact pins. | Cleaned sensor contact pins. | Cleaned | HW | Basic Unit | SC Contacts
ABL90 | Verified. | Power supply fuses are good. Disconnected the battery so that the customer can use the analyzer over the weekend. Parts on order. No MTR Required. | D20B3CD1EA0C9264 | No | Verified. | Power supply fuses are good. Disconnected the battery so that the customer can use the analyzer over the weekend. Parts on order. No MTR Required. | Revisit required | HW | Power supply | Battery
ABL90 FLEX | Anwender haben Analysator neu gestartet bei Erscheinen der Meldung. Danach ist Analysator wieder ohne Meldung. Kein weiterer Fehler festgestellt. Datalogs und Service Dump gezogen zur weiteren Auswertung. | Anwender haben Analysator neu gestartet bei Erscheinen der Meldung. Danach ist Analysator wieder ohne Meldung. Kein weiterer Fehler festgestellt. Datalogs und Service Dump gezogen zur weiteren Auswertung. | SWO-164536 | No | Users have restarted Analyzer when the message appears. After that, the analyzer is again without a message. No further error detected. Datalogs and Service Dump drawn for further evaluation. | Users have restarted Analyzer when the message appears. After that, the analyzer is again without a message. No further error detected. Datalogs and Service Dump drawn for further evaluation. | Defect | SW | SW | SP exchange
ABL90 FLEX | Analysator zeigte beim Service die Meldung nicht. Überprüfung des Analysators, keine Meldung zum Lösungspack Fehler gefunden. Datalogs und Service Dump gezogen zur Auswertung. pO2 Kalibrierfehler, SC erneuert. | Analysator zeigte beim Service die Meldung nicht. Überprüfung des Analysators, keine Meldung zum Lösungspack Fehler gefunden. Datalogs und Service Dump gezogen zur Auswertung. pO2 Kalibrierfehler, SC erneuert. | SWO-164537 | No | The analyzer did not show the message to the service. Checking the analyzer, no message found about the solution pack error. Datalogs and Service Dump drawn for evaluation. pO2 calibration error, SC renewed. | The analyzer did not show the message to the service. Checking the analyzer, no message found about the solution pack error. Datalogs and Service Dump drawn for evaluation. pO2 calibration error, SC renewed. | Defect | SW | SW | SP exchange
ABL90 FLEX PLUS |  | B | D20B2233EA39CC49 | No |  |  | Adjustment | HW | Auto Inlet module | Inlet drive gear
ABL90 FLEX PLUS | Einlassdichtung erneuert, Fluid Transport geprüft, tHb Kal. durchgeführt. | Einlassdichtung erneuert, Fluid Transport geprüft, tHb Kal. durchgeführt. | SWO-164834 | No | Inlet seal renewed, fluid transport tested, tHb Kal. Performed. | Inlet seal renewed, fluid transport tested, tHb Kal. Performed. | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-168133 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-164352 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Lac
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-164819 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Glu/Lac
ABL825 FLEX |  | B | D20AB7F8EA32B97A | Yes
ABL825 FLEX | ABL äußerlich gereinigt Proteinentfernung durchgeführt Tubing nipple kit sowie Abfallpumpenschläuche erneuert Gerät dekontaminiert, Inlet Gasket aus Kundenbestand erneuert ABL kalibriert und überprüft | ABL äußerlich gereinigt Proteinentfernung durchgeführt Tubing nipple kit sowie Abfallpumpenschläuche erneuert Gerät dekontaminiert, Inlet Gasket aus Kundenbestand erneuert ABL kalibriert und überprüft | SWO-164453 | No | ABL externally cleaned Protein removal performed Tubing nipple kit as well as waste pump hoses renewed Device decontaminated, Inlet Gasket renewed from customer base ABL calibrated and checked | ABL externally cleaned Protein removal performed Tubing nipple kit as well as waste pump hoses renewed Device decontaminated, Inlet Gasket renewed from customer base ABL calibrated and checked | Cleaned | Fluidic pathway | Inlet module | Fluidic pathway
ABL835 FLEX | Proteinentfernung durchgeführt pH Elektrode gereinigt, Inletgasket aus Kundenbestand erneuert Schläuche zwischen Einlasssonde und pH BG Modul erneuert Gerät kalibriert und überprüft, Blutmessungen durchgeführt | Proteinentfernung durchgeführt pH Elektrode gereinigt, Inletgasket aus Kundenbestand erneuert Schläuche zwischen Einlasssonde und pH BG Modul erneuert Gerät kalibriert und überprüft, Blutmessungen durchgeführt | SWO-164454 | No | Protein removal performed pH electrode cleaned, inlet gasket renewed from customer base Hoses between inlet probe and pH BG module renewed Device calibrated and checked, blood measurements carried out | Protein removal performed pH electrode cleaned, inlet gasket renewed from customer base Hoses between inlet probe and pH BG module renewed Device calibrated and checked, blood measurements carried out | Cleaned | Fluidic pathway | pH/BG Module | Fluidic pathway
ABL825 FLEX | Überprüfung der Deckelsensoren über das Servicemenü: Sensoren arbeiten unzuverlässig, vor allem der untere. --> Erneuerung der entsprechenden Autocheck Interface PCB incl. Sensoren. Funktionskontrolle: I.O. -Garantie- | Überprüfung der Deckelsensoren über das Servicemenü: Sensoren arbeiten unzuverlässig, vor allem der untere. --> Erneuerung der entsprechenden Autocheck Interface PCB incl. Sensoren. Funktionskontrolle: I.O. -Garantie- | SWO-164571 | No | Checking the lid sensors via the service menu: Sensors work unreliably, especially the lower one. --> Renewal of the corresponding Autocheck Interface PCB incl. sensors. Functional control: I.O. -Guarantee | Checking the lid sensors via the service menu: Sensors work unreliably, especially the lower one. --> Renewal of the corresponding Autocheck Interface PCB incl. sensors. Functional control: I.O. -Guarantee | Defect | HW | Auto Check Module | PCB, Autocheck
ABL837 FLEX | Removed waste line and waste bottle connector and flushed as blocked removed blood clot. Replaced waste line and test ok. | Removed waste line and waste bottle connector and flushed as blocked removed blood clot. Replaced waste line and test ok. | SWO-164482 | No | Removed waste line and waste bottle connector and flushed as blocked removed blood clot. Replaced waste line and test ok. | Removed waste line and waste bottle connector and flushed as blocked removed blood clot. Replaced waste line and test ok. | Cleaned | Fluidic pathway | Basic Unit | Fluidic pathway
ABL815 FLEX | Mehrfach Inlet Kal in Service Programm durchgeführt. Ok. Nach Neustart wieder Inlet Kal Error. Einlassdichtung ersetzt. | B | SWO-164527 | No | Multiple Inlet Kal in Service Program performed. Ok. After rebooting again Inlet Kal Error. replaced inlet seal. |  | Defect | HW | Inlet module | Inlet gasket
AQT90 FLEX | Schlauch Nr 6 steckt nicht auf dem Manifold. Schlauch wieder angesteckt. Funktionstest iO. | Schlauch Nr 6 steckt nicht auf dem Manifold. Schlauch wieder angesteckt. Funktionstest iO. | SWO-164553 | No | Tube No 6 is not on the Manifold. hose replugged. Functional test iO. | Tube No 6 is not on the Manifold. hose replugged. Functional test iO. | Defect | Consumables | SP | Valve, blocked
AQT90 FLEX |  | B | D20AFF6DEA1C40A7 | No |  |  | Defect | HW | Analyzer Control Module | Cable, UIM
AQT90 FLEX | Umgebungstemperatur angeglichen, Probennnehmerkennung gereinigt und kalibriert, Probenmessungen o.k. | Umgebungstemperatur angeglichen, Probennnehmerkennung gereinigt und kalibriert, Probenmessungen o.k. | SWO-164474 | No | Ambient temperature, sample detection cleaned and calibrated, sample measurements ok. | Ambient temperature, sample detection cleaned and calibrated, sample measurements ok. | Setup | HW | Analyzer | Ambient Temp | 12
AQT90 FLEX |  | B | D20ADBCFEA26746F | No |  |  | Defect | HW | Process Module | Tubes
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-164641 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-165225 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-164637 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX |  | 光单依然存在质量问题，报错1031，需要重新再借一个903-364，此单先关闭处理 by jesson | SWO-164634 | No |  | Light list still has quality problems, error 1031, need to borrow another 903-364, this single first close processing by jesson | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-164646 | No |  |  | Defect | HW | Process Module | PCB, Inlet
ABL80 FLEX BASIC | 更换备件902-972后修复(by tim) | 更换备件902-972后修复(by tim) | SWO-164631 | No | Repair after replacement of spare parts 902-972 (by tim) | Repair after replacement of spare parts 902-972 (by tim)
ABL90 FLEX | Lösungspack erneuert und Ersatz bereitgestellt. Funktionskontrolle in Ordnung. | Lösungspack erneuert und Ersatz bereitgestellt. Funktionskontrolle in Ordnung. | SWO-164738 | No | solution pack renewed and replacement provided. Functional control in order. | solution pack renewed and replacement provided. Functional control in order. | Assisted Replacement | SP | SP | pO2
ABL90 FLEX | 更换903-183后定标通过，组装时发现903-158上的4个弹片弹起的程度明显不一致，更换后修复。by jesson | 更换903-183后定标通过，组装时发现903-158上的4个弹片弹起的程度明显不一致，更换后修复。by jesson | SWO-164401 | No | After the replacement 903-183 after the calibration pass, assembly found that the 903-158 on the four shrapnel up the degree of significantinin, after replacement repair. by jesson | After the replacement 903-183 after the calibration pass, assembly found that the 903-158 on the four shrapnel up the degree of significantinin, after replacement repair. by jesson | Defect | HW | Basic Unit | Sensor Interface PCB
AQT90 FLEX | 更换636-550后修复 by jesson | 更换636-550后修复 by jesson | SWO-164615 | No | Repair by jesson after replacement 636-550 | Repair by jesson after replacement 636-550 | Defect | HW | Process Module | Process Unit sensor | 5784
AQT90 FLEX |  | B | SWO-164712 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
ABL80 FLEX BASIC |  | B | SWO-164880 | No
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-164824 | No | Replace the test card | Replace the test card
ABL80 FLEX Basic version |  | B | D20EDA304C3E9A86 | Yes
ABL80 CO-OX | Checked lines for crimps. Checked waste pump. Found problem to be defective solution pack | Replaced solution pack. | D20EDABD4C2770A0 | Yes | Checked lines for crimps. Checked waste pump. Found problem to be defective solution pack | Replaced solution pack.
ABL80 CO-OX | Found analyzer inoperable upon arriving, customer knew that it was clotted from the blood that was run. | Opened analyzer and found all tubings with blood clots. Attempted to flush tubings and pumps but could not get clean enough as unit sat over weekend. Installed PM kit into analyzer and flushed hemolyzer with hypochlorite solution. Tested analyzer and found operations to be normal, checked for any leaks before installing covers, found ok. | D21B45EAE8565126 | No | Found analyzer inoperable upon arriving, customer knew that it was clotted from the blood that was run. | Opened analyzer and found all tubings with blood clots. Attempted to flush tubings and pumps but could not get clean enough as unit sat over weekend. Installed PM kit into analyzer and flushed hemolyzer with hypochlorite solution. Tested analyzer and found operations to be normal, checked for any leaks before installing covers, found ok.
ABL80 OSM | Troubleshot to a bad CF card | Replaced the CF card, and this resolved the problem. | D20EB8344C1210D5 | No | Troubleshot to a bad CF card | Replaced the CF card, and this resolved the problem.
ABL90 | Reconnected the battery and verifed the problem. After the battery was replaced the ABL90 still would not power on. | Replaced the battery. Replaced the power supply after the battery did not solve the problem. | D20F401D4C07A7AC | No | Reconnected the battery and verifed the problem. After the battery was replaced the ABL90 still would not power on. | Replaced the battery. Replaced the power supply after the battery did not solve the problem. | Unknown | HW | Power Supply | Power Supply
ABL90 FLEX | Mehrere Kalibrierungen und QC-MESSUNGEN wiederholt, System und Funktion i.O. | Mehrere Kalibrierungen und QC-MESSUNGEN wiederholt, System und Funktion i.O. | SWO-165015 | No | Multiple calibrations and QC MEASUREMENTS repeated, system and function in the | Multiple calibrations and QC MEASUREMENTS repeated, system and function in the | Assisted Replacement | SC | SC | Glu/Lac | Start up
ABL90 FLEX | Kreißsaal hat ein LP erhalten, Intensiv hat zwei LPs erhalten, Notaufnahme hat ein PL erhalten. In de Notaufnahme noch mehrere Fragen zum ABL beantwortet und die Fehlerbehebung durchgesprochen. | Kreißsaal hat ein LP erhalten, Intensiv hat zwei LPs erhalten, Notaufnahme hat ein PL erhalten. In de Notaufnahme noch mehrere Fragen zum ABL beantwortet und die Fehlerbehebung durchgesprochen. | SWO-165196 | Yes | Kreißsaal has received an LP, Intensively received two LPs, Emergency room has received a PL. In the emergency room several questions about the ABL answered and the troubleshooting was answered. | Kreißsaal has received an LP, Intensively received two LPs, Emergency room has received a PL. In the emergency room several questions about the ABL answered and the troubleshooting was answered. | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Startphase QC und Kalibration. | Fabian: Auch nach allen durchgeführten Aktionen und erneutem Wechsel des SP sowie der Einlassdichtung keine Änderung, weiterhin Leckage-Meldung mit Intervention erforderlich. Fabian wünscht einen Serviceeinsatz morgen so früh wie möglich. | SWO-164924 | No | Start phase QC and calibration. | Fabian: Even after all actions taken and the CHANGE of the SP and the inlet seal no change, still leakage message with intervention required. Fabian wants a service assignment tomorrow as early as possible. | Revisit required | HW | Leakage | Leakage
ABL90 FLEX PLUS | WE Dienst | WE Dienst | SWO-165159 | No | WE Service | WE Service | Revisit required | analyzer removed | analyzer removed | analyzer removed
ABL90 FLEX PLUS | Ich habe auf der Station B2 Ca. aktiviert beide ABL´s ein kalibriert und QC´s gemessen. Danach haben wir vier Vergleichsmessungen durchgeführt. Das Ergebnis bei Ca. lag bei beiden ABL´s bei vier Messungen zwischen 1,28 und 1,29. Das Ca. auf der St. B2 habe ich nach den Vergleichsmessungen deaktiviert. | Ich habe auf der Station B2 Ca. aktiviert beide ABL´s ein kalibriert und QC´s gemessen. Danach haben wir vier Vergleichsmessungen durchgeführt. Das Ergebnis bei Ca. lag bei beiden ABL´s bei vier Messungen zwischen 1,28 und 1,29. Das Ca. auf der St. B2 habe ich nach den Vergleichsmessungen deaktiviert. | SWO-165167 | No | I have activated on station B2 Ca. both ABL's one calibrated and MEASURED QC's. After that, we carried out four comparative measurements. The result for Approx. was four measurements between 1.28 and 1.29 for both ABLs. The approx. on the St. B2 I deactivated after the comparison measurements. | I have activated on station B2 Ca. both ABL's one calibrated and MEASURED QC's. After that, we carried out four comparative measurements. The result for Approx. was four measurements between 1.28 and 1.29 for both ABLs. The approx. on the St. B2 I deactivated after the comparison measurements. | No fault found | Performance | Performance | Ca
ABL90 FLEX PLUS |  | B | D20EE0C04C50B98C | Yes
ABL90 FLEX PLUS | Zodoende de barcodescanner unit heb vervangen vervolgens heb gecontroleerd. Tevens de beveiligingsopties en maken van de back-up, in beide ABL90 toestellen heb uitgelegd aangepast. | Zodoende de barcodescanner unit heb vervangen vervolgens heb gecontroleerd. Tevens de beveiligingsopties en maken van de back-up, in beide ABL90 toestellen heb uitgelegd aangepast. | SWO-165075 | No | So the barcode scanner unit got replaced then checked. Also changed the security options and create the backup, in both ABL90 devices explained. | So the barcode scanner unit got replaced then checked. Also changed the security options and create the backup, in both ABL90 devices explained. | User error | Barcode Quality | Barcode Quality | Barcode Quality
TCM CombiM |  | B | SWO-164883 | No
AQT90 FLEX | 更换903-364修复(by zhenjie) | 更换903-364修复(by zhenjie) | SWO-164938 | No | Replacement 903-364 repair (by zhenjie) | Replacement 903-364 repair (by zhenjie) | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX |  | B | SWO-164785 | No |  |  | Defect | HW | Process Module | Process Unit sensor | 5784
AQT90 FLEX | Limpiado céluda de hematocrito. Pasados QCs y muestsras para comprobar el estado del equipo. | B | D20EA9BA4C13C2DF | No | Cleaned with hematocrit. Passed QCs and standards to check the status of the equipment. |  | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1223
ABL800 BASIC | Reset the Oxy-metry module & Recalibrate system. Run oxy test & passed. calibrate system & qc pass. | Reset the Oxy-metry module & Recalibrate system. Run oxy test & passed. calibrate system & qc pass. | SWO-166257 | No | Reset the Oxy-metry module & Recalibrate system. Run oxy test & passed. calibrate system & qc pass. | Reset the Oxy-metry module & Recalibrate system. Run oxy test & passed. calibrate system & qc pass. | Adjustment | HW | Oxi module | tHb Cal
ABL820 FLEX | Patti said that everything was out. It was givomng a liquid sensor error. I asked her to check the data logs and found that she was getting pH BG module detected rinse when not expected. This is the tiop of the air vavle broken off. That is what I found. Patti had to leave so I will sign verbal. | Replaced pH BG module air valve. | D20ED6E34C08B6BF | No | Patti said that everything was out. It was givomng a liquid sensor error. I asked her to check the data logs and found that she was getting pH BG module detected rinse when not expected. This is the tiop of the air vavle broken off. That is what I found. Patti had to leave so I will sign verbal. | Replaced pH BG module air valve. | Defect | HW | pH/BG Module | Valve, air
ABL800 BASIC | Halogen Lamp 調整 3400 | OXI PCB(903-654)交換 | SWO-165102 | No | Halogen Lamp Adjustment 3400 | OXI PCB (903-654) Exchange | Adjustment | HW | Oxi module | Lamp unit | Halogen
ABL827 FLEX | Confirmed problem upon arrival. | Removed modules and checked the air valves. Updated the Zero current (Crea was 0.000). Sens is still too high. Scanned and installed all new Cal 1, Cal 2 and Cleaning bottles. | D20F2A504C50169D | No | Confirmed problem upon arrival. | Removed modules and checked the air valves. Updated the Zero current (Crea was 0.000). Sens is still too high. Scanned and installed all new Cal 1, Cal 2 and Cleaning bottles. | Adjustment | HW | Met II | Crea | Zero Current
ABL825 FLEX | keine Lactatmembranen auf der Station deshalb keine Membranierung möglich, Glukoseelektrode membraniert und Gerät kalibriert, | keine Lactatmembranen auf der Station deshalb keine Membranierung möglich, Glukoseelektrode membraniert und Gerät kalibriert, | SWO-164971 | No | no lactate membranes on the station therefore no membraneing is possible, glucose electrode membranes and device calibrated, | no lactate membranes on the station therefore no membraneing is possible, glucose electrode membranes and device calibrated, | Assisted Replacement | Electrode | Membrane | Glu | potentiel revisit (Lac)
ABL810 FLEX | Einl.dichtung gereinigt, Gerät kalibriert, div. Konfigurationseinstellungen durchgeführt, | Einl.dichtung gereinigt, Gerät kalibriert, div. Konfigurationseinstellungen durchgeführt, | SWO-164944 | No | Sealing cleaned, device calibrated, various configuration settings performed, | Sealing cleaned, device calibrated, various configuration settings performed, | Adjustment | HW | Inlet module | Probe cal
ABL825 FLEX | Neon # too low on 2 4, 0316 and 0258. Also tHb on Calibration is reading 613. Also, when adjusting lamp, value is not stable. | Replaced Lamp Unit. Perform complete Oxi alignment. Cal tHb # now 572. Since customer has to restart QC validation, Brought customer one box each of QC as a courtesy. | D20EB0AA4C5A987F | No | Neon # too low on 2 4, 0316 and 0258. Also tHb on Calibration is reading 613. Also, when adjusting lamp, value is not stable. | Replaced Lamp Unit. Perform complete Oxi alignment. Cal tHb # now 572. Since customer has to restart QC validation, Brought customer one box each of QC as a courtesy. | Defect | HW | Oxi module | Lamp unit
AQT90 FLEX | Replaced and Re Installed New Optical Unit and Flat Cable with New one,perform system cleaning, Problem Resolve. | Replaced and Re Installed New Optical Unit and Flat Cable with New one,perform system cleaning, Problem Resolve. | SWO-166269 | No | Replaced and Re Installed New Optical Unit and Flat Cable with New one,perform system cleaning, Problem Resolve. | Replaced and Re Installed New Optical Unit and Flat Cable with New one,perform system cleaning, Problem Resolve. | Defect | HW | Process Module | Optical Unit
ABL80 FLEX BASIC |  | B | SWO-165118 | No
ABL80 CO-OX | Found defective cooling on EMI computer. | Replaced fan and tested analyzer. | D2102C23C2284343 | No | Found defective cooling on EMI computer. | Replaced fan and tested analyzer.
ABL80 FLEX BASIC | Found that Sensor cassette got leaked the solution, do to that found T M error, replaced the SC but same error present, found connector problem, need to replace with new SBC ASAP. failed SC details 300 30, Lot 230569, Sn, 1520247. | Taken the sensor connector board from standby analyzer replaced the same, performed the zeroing, installed the new SC found instrument is functional. | SWO-166197 | No | Found that Sensor cassette got leaked the solution, do to that found T M error, replaced the SC but same error present, found connector problem, need to replace with new SBC ASAP. failed SC details 300 30, Lot 230569, Sn, 1520247. | Taken the sensor connector board from standby analyzer replaced the same, performed the zeroing, installed the new SC found instrument is functional.
ABL80 FLEX BASIC | Taken the sensor connector board from standby analyzer replaced the same, performed the zeroing, installed the new SC found instrument is functional. | Taken the sensor connector board from standby analyzer replaced the same, performed the zeroing, installed the new SC found instrument is functional. | SWO-166199 | Yes | Taken the sensor connector board from standby analyzer replaced the same, performed the zeroing, installed the new SC found instrument is functional. | Taken the sensor connector board from standby analyzer replaced the same, performed the zeroing, installed the new SC found instrument is functional. | Re classified RMED | Double entry
ABL90 | Customer noticed this at the end result of LIS report. Looking further Radiance was also showing a variance between what was reported. Upon further inspection results are coming from the analyzer like that. Machine setup verified. Sample input information verified as the same. | No correction found at this point. Sending data off to open CQC case. | D20FF93CC23D9D49 | Yes | Customer noticed this at the end result of LIS report. Looking further Radiance was also showing a variance between what was reported. Upon further inspection results are coming from the analyzer like that. Machine setup verified. Sample input information verified as the same. | No correction found at this point. Sending data off to open CQC case.
ABL90 FLEX | Datenspeicher überprüft, Kal und QC wiederholt, z. Zt. OK. | Datenspeicher überprüft, Kal und QC wiederholt, z. Zt. OK. | SWO-165391 | No | Data store checked, Kal and QC repeated, currently OK. | Data store checked, Kal and QC repeated, currently OK. | Assisted Replacement | SC | SC | Glu/Lac | Start up
ABL90 FLEX | 4 Lösungspacks geliefert | 4 Lösungspacks geliefert | SWO-165534 | Yes | 4 solution packs delivered | 4 solution packs delivered | Re classified RMED | Logistical upSWO
ABL90 FLEX | Gerät mit warmen Wasser 10- 15 mal durchgespült. Einlasssonde durch Kundeneigene ersetzt. Kal und QC sind OK. Absprache mit Kunden getroffen, dass ein ABL90 LP als Ersatz geliefert wird. Aktuelle Lot ist die YH50. | DRINGEND! Kd. benötigt schnell den Besuch eines KDB. | SWO-165536 | No | Device rinsed with warm water 10-15 times. inlet probe replaced by customer-owned. Kal and QC are OK. agreement with customers that an ABL90 LP will be delivered as a replacement. The current lot is the YH50. | urgent! Kd. needs a visit to a KDB quickly. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | LP zum Kunden gebracht, da kein neuer vorrätig ist bzw. Bestellung nicht kam | LP ersetzt - QC i.O. ein LP YB-50 XL zurückgenommen, da er sehr wahrscheinlich nicht funktionieren wird | SWO-165519 | Yes | LP brought to the customer as no new stock is available or order did not arrive | LP replaced - QC in the world an LP YB-50 XL is taken back as it will most likely not work | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS |  | B | D21007B4C22C7672 | Yes |  |  | Re classified RMED | Ewen mail | Thursday, February 9, 217 12:56 AM | Reference
ABL90 FLEX PLUS |  | B | D20FE80CC2057665 | Yes |  |  | Re classified RMED | Ewen mail | Thursday, February 9, 217 12:56 AM | Glu | LEH Mail
AQT90 FLEX | 清除测试卡角内的碎片，更换新测试卡后修复 | 清除测试卡角内的碎片，更换新测试卡后修复 | SWO-165469 | No | Remove debris from the test card angle and repair it after replacing the new test card | Remove debris from the test card angle and repair it after replacing the new test card | Defect | Consumables | Test Kit | Cartridge | 5784 (Broken)
ABL800 FLEX | Replace the membrane & recalibrate system. | Replace the membrane & recalibrate system. | SWO-166256 | No | Replace the membrane & recalibrate system. | Replace the membrane & recalibrate system. | Assisted Replacement | Electrode | Membrane | Glu
ABL825 FLEX | pCO2-Elektrode erneuert (alt: DK-02) Diverse Pumpenschläuche erneuert | pCO2-Elektrode erneuert (alt: DK-02) Diverse Pumpenschläuche erneuert | SWO-165414 | No | pCO2 electrode renewed (old: DK-02) Various pump hoses renewed | pCO2 electrode renewed (old: DK-02) Various pump hoses renewed | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | replaced the printer head with new printer head now printer working. | B | SWO-166268 | No | replaced the printer head with new printer head now printer working. |  | Defect | HW | Printer Module | Printer head
ABL815 FLEX | Alle Fliesswege erneuert , Probe Messungen beobachtet. Kein Fehler feststellbar | Kunde bittet dringend um Einsatz, schaltet Gerät jetzt ab, Schlauch wurde bereits aufgesetzt (über den Kal.Lösungen)und springt direkt wieder ab | SWO-165289 | No | All flow paths renewed, sample measurements observed. No error detectable | Customer urgently asks for use, turns off device now, hose has already been put on (via the Kal.Solutions) and jumps off again | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL837 FLEX | Found sensor dirty | Cleaned with alcohol sensor | D20FE02AC24D73CD | No | Found sensor dirty | Cleaned with alcohol sensor | Cleaned | HW | Auto Check Module | Sensor
ABL825 FLEX | Oxi Modul geprüft, AC's wiederholt ok, SW geprüft, Viren Scan negativ, Funktionskontrolle ok | Oxi Modul geprüft, AC's wiederholt ok, SW geprüft, Viren Scan negativ, Funktionskontrolle ok | SWO-165374 | No | Oxi module checked, AC's repeated ok, SW checked, virus scan negative, function control ok | Oxi module checked, AC's repeated ok, SW checked, virus scan negative, function control ok | Cleaned | Fluidic pathway | Oxi module | Oxi module
ABL837 FLEX |  | B | D222BF81AC127EB4 | No |  |  | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX |  | B | D210019EC2052542 | No |  |  | Defect | HW | Incubator module | PCB, Incubator | 12
AQT90 FLEX | Fehlerursache beschränkt sich auf Verstopfung der Abfallstrecke und Needle Wash Unit. Reinigungsprozeduren der Pumpen 1&2 und Pumpen 4&5 durchgeführt. Pumpenkopf P5 ausgetauscht. Aufgrund des Fehler 1999 wurde das Antriebsrad des Incubatorwheels erneuert. Initialisierungsphase. Test im Servicemenü ohne Beanstandung. Blankmessungen durchgeführt und mittels Tool analysiert. Alle 5 Testmessungen unter 1600, i.O QC in allen Parametern ohne Beanstandung Sporadische Fehlermeldung BNP Sensor kaputt oder nicht vorhanden. Sensor wird bestellt und nach Erhalt eingebaut. | Kd. benötigt zur Behebung der Störung den Besuch eines KDB. | SWO-165407 | No | The cause of the error is limited to blockage of the waste section and Needle Wash Unit. cleaning procedures of pumps 1&2 and pumps 4&5. Pump head P5. Due to the fault in 1999, the drive wheel of the Incubatorwheel was renewed. Initialization phase. Test in the service menu without complaint. blank measurements and analyzed using a tool. All 5 test measurements under 1600, i.O QC in all parameters without complaint Sporadic error message BNP sensor broken or non-existent. Sensor is ordered and installed upon receipt. | Kd. requires a visit to a KDB to resolve the fault. | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1999
ABL80 OSM | CONFIRMED ISSUE WHEN ONSITE. | PER ASM JAIME APONTE, DOWNLOAD SYSTEM LOG FILES TO FILE TECHLINE CASE, REPLACE TUBINGS AND PARTS IN PM KITS. | D2117DF13841E546 | No | CONFIRMED ISSUE WHEN ONSITE. | PER ASM JAIME APONTE, DOWNLOAD SYSTEM LOG FILES TO FILE TECHLINE CASE, REPLACE TUBINGS AND PARTS IN PM KITS.
ABL80 FLEX BASIC | 更换备用机上的903-001后修复(by tim) | 更换备用机上的903-001后修复(by tim) | SWO-165493 | No | Repair after replacing 903-001 on the standby (by tim) | Repair after replacing 903-001 on the standby (by tim)
ABL80 CO-OX | Spoke with Bobbie and found that the CA failed to calibrate days after SC being replaced. Arrived onsite and verifed the failure. | Replaced the Hemolyzer, Analog Board and Sensor connector Board. Replaced SP and SC and all parameters green once SC replaced. Monitored system cycles for two weeks without failure. | D2117BA53808375A | Yes | Spoke with Bobbie and found that the CA failed to calibrate days after SC being replaced. Arrived onsite and verifed the failure. | Replaced the Hemolyzer, Analog Board and Sensor connector Board. Replaced SP and SC and all parameters green once SC replaced. Monitored system cycles for two weeks without failure.
ABL80 FLEX CO-OX | Kundengerät zur Überprüfung in GS Aufstellung Reparaturaustauschgerät LP und SC installiert tHb-Kal i.O. Pat-Datenbank gelöscht | Ein Vorortservice wird gewünscht. | SWO-165622 | No | Customer device for verification in GS Installation of repair replacement device LP and SC installed tHb-Kal i.O. Pat database deleted | A suburban service is desired.
ABL80 OSM | Ran service test to check hemolyzer. Failed Hemolyzer Fres, Hemolyzer peak voltage, Hemolyzer frequency, Cavitation absorb. | Replaced hemolyzer 906-032. Ran tHb calibration. Ran multiple test samples to verify. | D2115E983862B7BD | No | Ran service test to check hemolyzer. Failed Hemolyzer Fres, Hemolyzer peak voltage, Hemolyzer frequency, Cavitation absorb. | Replaced hemolyzer 906-032. Ran tHb calibration. Ran multiple test samples to verify.
ABL80 FLEX BASIC | 更换902-970后修复(by tim) | 更换902-970后修复(by tim) | SWO-165224 | No | Repair after replacement 902-970 (by tim) | Repair after replacement 902-970 (by tim)
ABL90 FLEX | Cleaned sensor pins on interface PCB on power up now has temp error top plate out of range. Unable to resolve issue parts required. | Replaced sensor interface PCB and test ok Calibration and QC ok. | SWO-165708 | No | Cleaned sensor pins on interface PCB on power up now has temp error top plate out of range. Unable to resolve issue parts required. | Replaced sensor interface PCB and test ok Calibration and QC ok. | Revisit required | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX PLUS |  | B | D2113BD0380DD6D8 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-168131 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | chip not read
AQT90 FLEX |  | B | SWO-165564 | No |  |  | Cleaned | HW | Aspiration Module | Cavo Pump | 1999
AQT90 FLEX |  | B | SWO-165548 | No |  |  | Cleaned | HW | Aspiration Module | Needle Wash Unit | 3412
ABL837 FLEX | Analyzer has been powered down and was likely down overnight. AutoCheck cover door is observed to open and close continuously. A grinding sound is also heard. Upon further investigation, fluid is observed to have entered into the AutoCheck module. | Replaced one of the AutoCheck interface pcb's due to fluid damage. Verified AutoCheck performance. Also replaced reference, glucose and lactate membranes due to failed parameters, which were likely caused by the analyzer being powered down overnight. Analyzer is within acceptable operating specifications. | D21156D5381C251B | No | Analyzer has been powered down and was likely down overnight. AutoCheck cover door is observed to open and close continuously. A grinding sound is also heard. Upon further investigation, fluid is observed to have entered into the AutoCheck module. | Replaced one of the AutoCheck interface pcb's due to fluid damage. Verified AutoCheck performance. Also replaced reference, glucose and lactate membranes due to failed parameters, which were likely caused by the analyzer being powered down overnight. Analyzer is within acceptable operating specifications. | Defect | HW | Auto Check Module | PCB, Autocheck | Fluid damage
ABL825 FLEX | Regelmäßig hohe Driften bei Hb Kal => per Remotezugriff diverse Aktivitäten durchgeführt, Drift nun gegen 0. Bei fortwährenden Problemen muss ggf. noch das Ocimodul ersetzt werden! | Regelmäßig hohe Driften bei Hb Kal => per Remotezugriff diverse Aktivitäten durchgeführt, Drift nun gegen 0. Bei fortwährenden Problemen muss ggf. noch das Ocimodul ersetzt werden! | SWO-165600 | No | Regularly high drifts at Hb Kal => performed via remote access various activities, drift now against 0. In case of continuous problems, the Ocimodul may need to be replaced! | Regularly high drifts at Hb Kal => performed via remote access various activities, drift now against 0. In case of continuous problems, the Ocimodul may need to be replaced! | Defect | HW | Oxi module | Oxi module
ABL827 FLEX | discussed instrument data: drew up the plan to clean and remembrane all electrodes. instrument study for 2 weeks no maintenance to be performed .noted salts in reference chamber and crescent shaped debris removed from cl- chamber | i remembraned all electrodes after cleanin the measuring chamber wells | D2114C1B382F3DD3 | No | discussed instrument data: drew up the plan to clean and remembrane all electrodes. instrument study for 2 weeks no maintenance to be performed .noted salts in reference chamber and crescent shaped debris removed from cl- chamber | i remembraned all electrodes after cleanin the measuring chamber wells | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL827 FLEX | Beim Eintreffen Fehlermeldung 316,258,247,248 --> ABL aus ein --> Fehler 0258,0316 --> Oxi-complete-Test durchgeführt --> fehlerhaft Ersatzteile werden bestellt Station und MedTech informiert | Beim Eintreffen Fehlermeldung 316,258,247,248 --> ABL aus ein --> Fehler 0258,0316 --> Oxi-complete-Test durchgeführt --> fehlerhaft Ersatzteile werden bestellt Station und MedTech informiert | SWO-165571 | No | On arrival error message 316,258,247,248 --> ABL from a --> Error 0258,0316 --> Oxi-complete test performed --> faulty Spare parts are ordered Station and MedTech informed | On arrival error message 316,258,247,248 --> ABL from a --> Error 0258,0316 --> Oxi-complete test performed --> faulty Spare parts are ordered Station and MedTech informed | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX | no action required user error. Performed the system clean procedure. | no action required user error. Performed the system clean procedure. | SWO-166064 | No | no action required user error. Performed the system clean procedure. | no action required user error. Performed the system clean procedure. | User error | User error | User error | User error
AQT90 FLEX | Analysator geöffnet. BNP Temperatursensor an Cup Wash Unit ausgetauscht. Initialisierungsphase. Servicetests QC ohne Beanstandung Eingabe der Anwender zur QC Prüfung | Analysator geöffnet. BNP Temperatursensor an Cup Wash Unit ausgetauscht. Initialisierungsphase. Servicetests QC ohne Beanstandung Eingabe der Anwender zur QC Prüfung | SWO-165701 | No | Analyzer open. BNP temperature sensor to Cup Wash Unit. Initialization phase. Service tests QC without complaint Entering the users for QC examination | Analyzer open. BNP temperature sensor to Cup Wash Unit. Initialization phase. Service tests QC without complaint Entering the users for QC examination | Defect | HW | Process Module | BNP Sensor
TCM400-4 | Cal error occurred on one of the sensor during sensor calibration. When inspect and replaced membrane, noticed one layer of the membrane still stuck on the sensor surface. | Removed extra membrane in sensor surface and remembrane sensor. | D2118764383301E5 | Yes | Cal error occurred on one of the sensor during sensor calibration. When inspect and replaced membrane, noticed one layer of the membrane still stuck on the sensor surface. | Removed extra membrane in sensor surface and remembrane sensor. | Cleaned | HW | Electrode | Membrane
TCM/Tosca | Ran a performance test and was able to recreate a high pCO2 measurement. | Replaced TOSCA Module | D21167F83839D9BB | No | Ran a performance test and was able to recreate a high pCO2 measurement. | Replaced TOSCA Module | Defect | HW | Tosca module
ABL80 FLEX BASIC | 更换924-738一个正常。(bt tim) | 更换924-738一个正常。(bt tim) | SWO-165837 | No | Replace 924-738 one normal. (bt tim) | Replace 924-738 one normal. (bt tim)
ABL90 FLEX | SP90 geliefert | SP90 geliefert | SWO-166022 | Yes | SP90 delivered | SP90 delivered | Re classified RMED | Logistical upSWO
ABL90 FLEX | Replaced sensor interface PCB and test ok Calibration and QC ok. | Replaced sensor interface PCB and test ok Calibration and QC ok. | SWO-166005 | No | Replaced sensor interface PCB and test ok Calibration and QC ok. | Replaced sensor interface PCB and test ok Calibration and QC ok. | Defect | HW | Basic Unit | Sensor Interface PCB
TCM CombiM |  | B | SWO-165861 | No
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-166729 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Glu | 377
ABL90 | REVISIÓN DE HISTÓRICOS. LIMPIEZA BAHIA DE ELECTRODOS. REALIZAR TEST TÉCNICOS. REEMPLAZAR CASSETTE DE ELECTRODOS DEFECTUOSO. REALIZAR TEST TÉCNICOS. CALIBRACIONES Y CONTROLES. | REVISIÓN DE HISTÓRICOS. LIMPIEZA BAHIA DE ELECTRODOS. REALIZAR TEST TÉCNICOS. REEMPLAZAR CASSETTE DE ELECTRODOS DEFECTUOSO. REALIZAR TEST TÉCNICOS. | D212AD33AE274CC8 | No | HISTORICAL REVIEW. CLEANING BAHIA DE ELECTRODOS. PERFORM TECHNICAL TESTS. REPLACEMENT DEFECT ELECTRODE CASSETTE. PERFORM TECHNICAL TESTS. CALIBRATIONS AND CONTROLS. | HISTORICAL REVIEW. CLEANING BAHIA DE ELECTRODOS. PERFORM TECHNICAL TESTS. REPLACEMENT DEFECT ELECTRODE CASSETTE. PERFORM TECHNICAL TESTS. | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX |  | B | SWO-165898 | No |  |  | Not Cat | Not cat. | Not cat. | Not cat. | 5684
ABL835 FLEX | Comprobar electrodo de referencia quitar un anillo sellador sobrante, limpiar y tensar los contactos eléctricos del electrodo. Revisión normal del equipo con limpieza de la cámara de medida y salida del desperdicio, comprobación del módulo oxi, realización de operaciones de mantenimiento y prueba con controles. | Quitar un anillo sobrante al electrodo de referencia. | D212B3BBAE26AD2A | No | Check reference electrode remove a leftover sealing ring, clean and tighten the electrical contacts of the electrode. Normal equipment check with measuring chamber cleaning and waste output, oxy module check, maintenance operations and testing with controls. | Remove a leftover ring from the reference electrode. | User error | Electrode | Electrode | Extra Sealing
ABL835 FLEX | ｌａｃ電極交換 Ｃａｌ２ ＱＣ ＯＫ | ｌａｃ電極交換 Ｃａｌ２ ＱＣ ＯＫ | SWO-167356 | No | Lac electrode replacement Cal2 QC OK | Lac electrode replacement Cal2 QC OK | Assisted Replacement | Electrode | Electrode | Lac | 376
ABL825 FLEX | Fehlersuche: Erneut Spectrum Mismatch Fehler bei QC Messungen => Oxi-Modul ersetzt, zweimal Kal1, OxiCompleteTest durchgeführt. Neustart, Kalibrationen, tHb-Kal, Funktionstest, QC Messungen OK | Fehlersuche: Erneut Spectrum Mismatch Fehler bei QC Messungen => Oxi-Modul ersetzt, zweimal Kal1, OxiCompleteTest durchgeführt. Neustart, Kalibrationen, tHb-Kal, Funktionstest, QC Messungen OK | SWO-165964 | No | Troubleshooting: Again Spectrum Mismatch Error in QC Measurements => Oxi module replaced, twice Kal1, OxiCompleteTest performed. Restart, calibrations, tHb-Kal, function test, QC measurements OK | Troubleshooting: Again Spectrum Mismatch Error in QC Measurements => Oxi module replaced, twice Kal1, OxiCompleteTest performed. Restart, calibrations, tHb-Kal, function test, QC measurements OK | Defect | HW | Oxi module | Spectrometer
ABL825 FLEX | Aufschaltung per Remotezugriff: Bei allen QC Leveln Na immer wieder oberhalb Grenze(n). Elektrode wurde bereits membraniert, keine Änderung. Ca sporadisch (selten) ebenfalls betroffen. Einsatz ab am Montag erforderlich! | Aufschaltung per Remotezugriff: Bei allen QC Leveln Na immer wieder oberhalb Grenze(n). Elektrode wurde bereits membraniert, keine Änderung. Ca sporadisch (selten) ebenfalls betroffen. Einsatz ab am Montag erforderlich! | SWO-165986 | Yes | Remote access: At all QC levels, Na repeatedly above limit(s). Electrode has already been membraneed, no change. Ca sporadically (rarely) also affected. Use required from Monday! | Remote access: At all QC levels, Na repeatedly above limit(s). Electrode has already been membraneed, no change. Ca sporadically (rarely) also affected. Use required from Monday! | Re classified RMED | Remote Support
ABL837 FLEX |  | B | D222CEC1AC317EBF | No |  |  | Defect | HW | DMS_PC | Power Supply
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-166092 | No | Repair after replacing test card | Repair after replacing test card
ABL80 CO-OX | Checked all data logs for any previous issues and previous issues and found the level 1 QC fails PO2 when ran the first time. | Performed zero cassette and ran three levels of QC. All tests are successful. | D213AF9B2459C92A | No | Checked all data logs for any previous issues and previous issues and found the level 1 QC fails PO2 when ran the first time. | Performed zero cassette and ran three levels of QC. All tests are successful.
ABL90 |  | B | D213CFF02401E17C | No |  |  | Adjustment | HW | Oxi module | tHb Cal
ABL90 | Remote connection is active. Contacted customer, accessed analyzer, view data logs, accessed service diagnostics, determined the fluid flow issues most likely caused by a faulty new SC just installed. | Suggested customer change Sensor Cassette, follow up indicates issue was resolved by replacement SC. | D216AE59101B21C2 | Yes | Remote connection is active. Contacted customer, accessed analyzer, view data logs, accessed service diagnostics, determined the fluid flow issues most likely caused by a faulty new SC just installed. | Suggested customer change Sensor Cassette, follow up indicates issue was resolved by replacement SC.
ABL90 | Reviewed data logs, calibration failing, observed error 1406 (not horizontal) ocuring from 2 8 to 2 9 intermittently, no errors after 2 9 @ 8:15AM. Sensor response errors began 2 9 @ 00:52. SC had been replace approximately a day earlier. Customer tried a second SC with the same issues. Learned while onsite another RT lead flushed the analyzer on 2 9 in the AM and calibration began passing. Suspect a possible unknown material passing through analyzer may have caused reported problem. | Upon my arrival another RT lead had flushed analyzer on 2 9, calibration began passing, only 1 level of external QC had failed from the previous day (2 9) but had not been rerun. I ran the QC, it passed. As a precautionary measure I removed and examined the sensor cassette bay and interface assembly, confirmed alignment, removed and reseat all connections. I reinstalled both the first and second SCs that had failed, both passed calibration and QC. Suspect possible unknown material, possibly from a sample was introduced into the analyzer causing the response and impedance error. All tests OK. | D213EF612444E088 | No | Reviewed data logs, calibration failing, observed error 1406 (not horizontal) ocuring from 2 8 to 2 9 intermittently, no errors after 2 9 @ 8:15AM. Sensor response errors began 2 9 @ 00:52. SC had been replace approximately a day earlier. Customer tried a second SC with the same issues. Learned while onsite another RT lead flushed the analyzer on 2 9 in the AM and calibration began passing. Suspect a possible unknown material passing through analyzer may have caused reported problem. | Upon my arrival another RT lead had flushed analyzer on 2 9, calibration began passing, only 1 level of external QC had failed from the previous day (2 9) but had not been rerun. I ran the QC, it passed. As a precautionary measure I removed and examined the sensor cassette bay and interface assembly, confirmed alignment, removed and reseat all connections. I reinstalled both the first and second SCs that had failed, both passed calibration and QC. Suspect possible unknown material, possibly from a sample was introduced into the analyzer causing the response and impedance error. All tests OK. | No fault found | No fault found | No fault found | No fault found | Cleaned Fluidic P.
ABL90 | error1307 | deleted temp file on c drive | D213DC26243F9D13 | No | error1307 | deleted temp file on c drive | Adjustment | SW | DMS_PC | Disc Space | temp files deleted
ABL90 | error 1307 | deleted temp file on c drive | D213DE31242C9D17 | No | error 1307 | deleted temp file on c drive | Adjustment | SW | DMS_PC | Disc Space | temp files deleted
ABL90 FLEX PLUS | Removed cover and adjusted inlet connection. Swapped back and removed spare analyser. Shut down spare analyser and returned to Glenfield workshop. Analyser sensor cassette had expired, replaced and ensured analyser passed all tests. | B | SWO-166208 | No | Removed cover and adjusted inlet connection. Swapped back and removed spare analyser. Shut down spare analyser and returned to Glenfield workshop. Analyser sensor cassette had expired, replaced and ensured analyser passed all tests. |  | Adjustment | HW | Auto Inlet module | Inlet Position Cal.
AQT90 FLEX |  | B | SWO-166109 | No |  |  | Defect | HW | Cover and Chassis | Inlet Door | 4584
AQT90 FLEX |  | B | SWO-166099 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
AQT90 FLEX | 更换备用机上的903-175后修复(by tim) | 更换备用机上的903-175后修复(by tim) | SWO-165789 | No | Post-replacement 903-175 on standby (by tim) | Post-replacement 903-175 on standby (by tim) | Defect | HW | Analyzer Control Module | Power Supply
ABL827 FLEX |  | B | D213F5D1245DBDAC | No |  |  | Defect | HW | DMS_PC | Power Supply
ABL820 FLEX | checked upper valves and replaced flat valve membrane and measuring chamber. | B | D21410A3245282E3 | No | checked upper valves and replaced flat valve membrane and measuring chamber. |  | Defect | HW | pH/BG Module | Valve, air
ABL827 FLEX |  | B | D2175EAF8622BC68 | No |  |  | Defect | HW | Inlet module | Inlet Gasket | Leaking
ABL825 FLEX |  | B | D213EFFB242D10DD | No |  |  | Defect | HW | Oxi module | Lamp unit
ABL815 FLEX | found solution nipple broken | replaced fluidic block; replaced fluidic block; removed and replaced all valves, moter, and pcb | D213D1C524281F52 | No | found solution nipple broken | replaced fluidic block; replaced fluidic block; removed and replaced all valves, moter, and pcb | Defect | HW | Inlet module | Nipple broken
ABL800 BASIC | 各ﾎﾟﾝﾌﾟﾁｭｰﾌﾞ、エアライン洗浄にて良好。 | 各ﾎﾟﾝﾌﾟﾁｭｰﾌﾞ、エアライン洗浄にて良好。 | SWO-166394 | No | Good for each pump tube and airline cleaning. | Good for each pump tube and airline cleaning. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL800 BASIC | ｲﾝﾚｯﾄﾌﾟﾛｰﾌﾞ、ｲﾝﾚｯﾄｶﾞｽｹｯﾄ交換。FlexQﾓｼﾞｭｰﾙの位置調整。 | ｲﾝﾚｯﾄﾌﾟﾛｰﾌﾞ、ｲﾝﾚｯﾄｶﾞｽｹｯﾄ交換。FlexQﾓｼﾞｭｰﾙの位置調整。 | SWO-166392 | No | Inlet probe, inlet gasket replacement. FlexQ module positioning. | Inlet probe, inlet gasket replacement. FlexQ module positioning. | Defect | HW | Inlet module | Inlet Gasket
AQT90 FLEX | CKMB Results failure, Analyzer Showing error 1029 Optical Sensitivity low, Perform System cleaning and Run sample, Problem Remains, order for Optical Unit for testion. | B | SWO-166271 | No | CKMB Results failure, Analyzer Showing error 1029 Optical Sensitivity low, Perform System cleaning and Run sample, Problem Remains, order for Optical Unit for testion. |  | Revisit required | HW | Process Module | Optical Unit
AQT90 FLEX | Eine TnI Testkassette mit abgebrochenem Mitnehmer geborgen. Storagewheel ausgebaut und erfolglos nach dem Bruchstück gesucht. Evtl. War die Kassette vor dem Einlegen durch Sturz beschädigt. Gerät zusammengebaut, Funktionskontrolle | Eine TnI Testkassette mit abgebrochenem Mitnehmer geborgen. Storagewheel ausgebaut und erfolglos nach dem Bruchstück gesucht. Evtl. War die Kassette vor dem Einlegen durch Sturz beschädigt. Gerät zusammengebaut, Funktionskontrolle | SWO-166296 | No | A TnI test cassette with aborted taker recovered. Storagewheel expanded and searched unsuccessfully for the fragment. Possibly the cassette was damaged by fall before inserting it. Device assembled, function control | A TnI test cassette with aborted taker recovered. Storagewheel expanded and searched unsuccessfully for the fragment. Possibly the cassette was damaged by fall before inserting it. Device assembled, function control | Defect | Consumables | Test Kit | Cartridge | Broken, nose
HbA1c 501 Analyzer | 代替機貸出。 | 代替機貸出。 | SWO-166249 | Yes | Alternative machine rental. | Alternative machine rental.
ABL90 FLEX PLUS |  | B | D216B827105782BA | Yes |  |  | Re classified RMED | Telp support /PER
ABL90 FLEX PLUS | Nettoyage de la section fluidique (suspicion caillot de sang dans la section fluidique) e.o Calibration e.o contrôle de qualité e.o | Nettoyage de la section fluidique (suspicion caillot de sang dans la section fluidique) e.o Calibration e.o contrôle de qualité e.o | SWO-166325 | No | Cleaning the fluid section (suspicion blood clot in the fluid section) e.o Calibration e.o quality control e.o | Cleaning the fluid section (suspicion blood clot in the fluid section) e.o Calibration e.o quality control e.o | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-168143 | No | Repair after replacing test card | Repair after replacing test card
ABL80 FLEX BASIC |  | B | SWO-166499 | No
ABL80 CO-OX | Error 2106 possible system blockage | Solution tubing to manifold dislodged. Reconnected tubing | D2178B27865373D8 | Yes | Error 2106 possible system blockage | Solution tubing to manifold dislodged. Reconnected tubing
ABL80 FLEX Basic version |  | B | D2175743864AA02F | No
ABL90 FLEX | Den Scanner getauscht. Danach Kal,QC durchgeführt. Scanner Funktion geprüft ok. | Den Scanner getauscht. Danach Kal,QC durchgeführt. Scanner Funktion geprüft ok. | SWO-166542 | No | Replace the scanner. Then Kal,QC performed. Scanner function checked ok. | Replace the scanner. Then Kal,QC performed. Scanner function checked ok. | Defect | HW | Basic Unit | Barcode reader
ABL90 FLEX | パーツキット交換 | B | SWO-166677 | No | Parts kit replacement |  | Defect | HW | Oxi module | Pinch valve
ABL90 | Instrument was asking customer to replace solution pack multiple times. Customer has already replaced the inlet gasket holder. Found bent inlet probe. | Adjust inlet probe. Verfied instrument operation. | D21F163B4A5C552B | No | Instrument was asking customer to replace solution pack multiple times. Customer has already replaced the inlet gasket holder. Found bent inlet probe. | Adjust inlet probe. Verfied instrument operation. | Replaced | HW | Inlet module | Probe bend
ABL90 FLEX | Replaced tubing using non stock spare, fitted original sc from spare analyser | Replaced tubing using non stock spare, fitted original sc from spare analyser | SWO-166594 | No | Replaced tubing using non stock spare, fitted original sc from spare analyser | Replaced tubing using non stock spare, fitted original sc from spare analyser | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX | neu eingesetzter LP Lot YB-53 meldet Chipdaten können nicht gelesen werden, Keine LP mehr vorrätig, bei eintreffen ist Gerät eingeschaltet und meldet einen inkonsistenten Datenbank Fehler -> LP eingesetzt , Startphase abgewartet, pCO2 & Glu Lac Kal Fehler -> SC sollte ersetzt werden ( fällig am 15.02) keine neue vorrätig -.>Schw. Kathrin kommt heute nachmittag | Schw. Kathrin hat in einem anderen Kühlschrank noch 2 SC, wird diese ersetzten und bestellt gleichzeitig noch einen LP | SWO-166553 | No | newly inserted LP Lot YB-53 reports chip data cannot be read, No more LP in stock, when it arrives, device is switched on and reports an inconsistent database error -> LP inserted, start phase waited, pCO2 & Glu Lac Kal error -> SC should be replaced (due on 15.02) no new in stock -.>Schw. Kathrin is coming this afternoon | Schw. Kathrin still has 2 SC in another refrigerator, will replace it and order an LP at the same time | Assisted Replacement | SP | SP | chip not read
ABL90 FLEX PLUS |  | B | D217B2C48635D009 | No |  |  | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-166726 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Glu/Lac
AQT90 FLEX | 更换备件903-175后修复(By Ran) | 更换备件903-175后修复(By Ran) | SWO-166274 | No | Replacement of spare parts 903-175 after repair (By Ran) | Replacement of spare parts 903-175 after repair (By Ran) | Defect | HW | Analyzer Control Module | Power Supply
ABL825 FLEX | SE INTENTA RESTITUIR EL SISTEMA REINICIANDO EL EQUIPO EN VARIAS OCASIONES NO SIENDO ESTO EFEECTIVO. SE SUSTITUYE PLACA DE CONTROL OXI, SE EFECTUAN AJUSTES Y CALIBRACION DE tHb. CALIBRADO Y VERIFICADO CON Q.C. | ERRORES DE HARDWARE COOX | D2178699861093BA | No | ATTEMPTS TO REST THE SYSTEM BY RESTARTING THE EQUIPMENT IN SEVERAL OCCASIONS NOT BEING THIS EFEECTIVE. OXI CONTROL PLACE IS SUBSTITUTE, tHb ADJUSTMENTS AND CALIBRATION ARE EFFECTED. CALIBRATED AND VERIFIED WITH Q.C. | COOX HARDWARE ERRORS | Defect | HW | Oxi module | PCB, OXI
ABL827 FLEX |  | B | D2179B968638C041 | No |  |  | Defect | HW | Inlet module | Nipple broken
ABL835 FLEX | Messkammer gesäubert, mehrere Blutmessungen durchgeführt, System und Funktion i.O. | Messkammer gesäubert, mehrere Blutmessungen durchgeführt, System und Funktion i.O. | SWO-166694 | No | Measuring chamber cleaned, several blood measurements carried out, system and function in the | Measuring chamber cleaned, several blood measurements carried out, system and function in the | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Lufteinlass komplett verstopft => freigespült. Weitere Tätigkeiten im Rahmen der HJW (gesonderter SWO). | Lufteinlass komplett verstopft => freigespült. Weitere Tätigkeiten im Rahmen der HJW (gesonderter SWO). | SWO-166549 | No | Air inlet completely clogged => rinsed free. Other activities under the HJW (separate SWO). | Air inlet completely clogged => rinsed free. Other activities under the HJW (separate SWO). | Cleaned | Fluidic pathway | Fluidic pathway | Air Intake
ABL825 FLEX | Alle in Frage kommenden Komponenten überprüft, Elektrolyt-Elektroden membraniert, Messkammern gereinigt. Ventiltest im Servicemenü, OK. Neustart, Kalibrationen, Endkontrolle: QC Messung, OK. Bei weiteren Ausreißern geht ein Datenextrakt zur Überprüfung an RDSEU | Alle in Frage kommenden Komponenten überprüft, Elektrolyt-Elektroden membraniert, Messkammern gereinigt. Ventiltest im Servicemenü, OK. Neustart, Kalibrationen, Endkontrolle: QC Messung, OK. Bei weiteren Ausreißern geht ein Datenextrakt zur Überprüfung an RDSEU | SWO-166615 | No | All eligible components checked, electrolyte electrodes membraneed, measuring chambers cleaned. Valve test in the service menu, OK. Restart, calibrations, final control: QC measurement, OK. For further outliers, a data extract goes to RDSEU for review | All eligible components checked, electrolyte electrodes membraneed, measuring chambers cleaned. Valve test in the service menu, OK. Restart, calibrations, final control: QC measurement, OK. For further outliers, a data extract goes to RDSEU for review | Assisted Replacement | Electrode | Membrane | Na+
ABL825 FLEX | Found seal rings missing from pO2 electrode and Glucose electrode. | Replace missing seal rings. | D217A4F186619B7E | No | Found seal rings missing from pO2 electrode and Glucose electrode. | Replace missing seal rings. | User error | Electrode | Electrode | Sealing ring
ABL827 FLEX | checked flow of solutions, checked waste lines and tubing, checked electrode chamber, | removed modules and cleaned spills, replaced pump tubing and connector tubing, replaced broken pH measuring chamber, removed debris from waste bottle nipple, | D218105D860E4858 | No | checked flow of solutions, checked waste lines and tubing, checked electrode chamber, | removed modules and cleaned spills, replaced pump tubing and connector tubing, replaced broken pH measuring chamber, removed debris from waste bottle nipple, | Defect | HW | pH/BG Module | Measuring chamber
ABL825 FLEX | AC Modul demontiert und gereinigt, Funktionskontrolle ok | AC Modul demontiert und gereinigt, Funktionskontrolle ok | SWO-166899 | No | AC module disassembled and cleaned, function control ok | AC module disassembled and cleaned, function control ok | Cleaned | HW | Auto Check Module | AutoCheck Module
ABL810 FLEX | Einlasssonde erneuert. | Einl.dichtung gereinigt, Gerät kalibriert, div. Konfigurationseinstellungen durchgeführt, | SWO-166662 | No | Inlet probe renewed. | Sealing cleaned, device calibrated, various configuration settings performed, | Defect | HW | Inlet module | Probe bend
ABL827 FLEX | Oximodul erneuert --> Einlasssonde verbogen, gerichtet Zusatz der Lösungen war nach Wechsel nicht aktiviert--> Fehler Crea --> berichtigt Finalkey eingespielt | Oximodul erneuert --> Einlasssonde verbogen, gerichtet Zusatz der Lösungen war nach Wechsel nicht aktiviert--> Fehler Crea --> berichtigt Finalkey eingespielt | SWO-166458 | No | Oximodul renewed --> inlet probe bent, directed Addition of solutions was not activated after change--> Error Crea --> corrected Finalkey recorded | Oximodul renewed --> inlet probe bent, directed Addition of solutions was not activated after change--> Error Crea --> corrected Finalkey recorded | Defect | HW | Oxi module | Oxi module
ABL80 FLEX BASIC | 更换测试卡和试剂包后修复 | 更换测试卡和试剂包后修复 | SWO-167296 | No | Repair after replacing test cards and reagent kits | Repair after replacing test cards and reagent kits
ABL80 FLEX BASIC |  | 更换924-738一个正常。(bt tim) | SWO-166688 | No |  | Replace 924-738 one normal. (bt tim)
ABL80 FLEX Basic version |  | B | D218C02FFC60C408 | No
ABL90 | While looking at analyzer Chuck and I found inconsistencies in the setup of this analyzer and the one next to it. | We adjusted the setting to match the other analyzer. Results seem to cross to radiance and MR system. Problem seems resolved for time being. | D218BA24FC1F167C | Yes | While looking at analyzer Chuck and I found inconsistencies in the setup of this analyzer and the one next to it. | We adjusted the setting to match the other analyzer. Results seem to cross to radiance and MR system. Problem seems resolved for time being.
ABL90 FLEX | CBC sample data compare test. User Pre analytical error. User training. | CBC sample data compare test. User Pre analytical error. User training. | SWO-167043 | No | CBC sample data compare test. User Pre analytical error. User training. | CBC sample data compare test. User Pre analytical error. User training. | User error | Sample | Oxi module | tHb Cal
ABL90 FLEX | LP installiert und Startphase abgewartet, neue Barcodetype zur Benutzer ID getestet, QC abgewartet | LP installiert und Startphase abgewartet, neue Barcodetype zur Benutzer ID getestet, QC abgewartet | SWO-167060 | Yes | LP installed and start-up phase awaited, new barcode type tested to user ID, QC awaited | LP installed and start-up phase awaited, new barcode type tested to user ID, QC awaited | Re classified RMED | Logistical upSWO
ABL90 FLEX | Kundin erklärt wie sie das Gerät manuell spült, Kundin geraten auch zusätzlich die Einlassdichtung zu erneuern. Sie wird sich wieder melden falls weitere Hilfen ötig ist. | Kundin erklärt wie sie das Gerät manuell spült, Kundin geraten auch zusätzlich die Einlassdichtung zu erneuern. Sie wird sich wieder melden falls weitere Hilfen ötig ist. | SWO-166973 | Yes | Customer explains how she rinses the device manually, customer also advises to renew the inlet seal. She will come back if further aid is available. | Customer explains how she rinses the device manually, customer also advises to renew the inlet seal. She will come back if further aid is available. | Re classified RMED | Phone support
ABL90 | Unit would be in home screen then would reboot, customer had turned unit off and restarted the analyzer this morning | Unit was in operation when i arrived, i reloaded Roma and tested operation of analyzer | D218EA81FC199F8B | No | Unit would be in home screen then would reboot, customer had turned unit off and restarted the analyzer this morning | Unit was in operation when i arrived, i reloaded Roma and tested operation of analyzer | Defect | SW | Roma.ini | Reboot
ABL90 FLEX PLUS | Diverse Tubings und Einlassdichtung erneuert. System und SC gespült. Kein Erfolg. SC Lot. 573 eingesetzt. Initalisierungsfehler braucht sehr lange bis alle Parameter o.k. sind. | Diverse Tubings und Einlassdichtung erneuert. System und SC gespült. Kein Erfolg. SC Lot. 573 eingesetzt. Initalisierungsfehler braucht sehr lange bis alle Parameter o.k. sind. | SWO-167051 | No | Various tubings and inlet seal renewed. system and SC. No success. SC Lot. 573. Initalization errors take a long time until all parameters are OK. | Various tubings and inlet seal renewed. system and SC. No success. SC Lot. 573. Initalization errors take a long time until all parameters are OK. | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS | The power supply was replaced with a new one which has had the issue corrected. | The power supply was replaced with a new one which has had the issue corrected. | SWO-167012 | No | The power supply was replaced with a new one which has had the issue corrected. | The power supply was replaced with a new one which has had the issue corrected. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX |  | B | SWO-166734 | No |  |  | Assisted Replacement | SC | SC | pO2
ABL825 FLEX | Software XP DMS910 und Analysator SW 6.16 neu installiert, Konfiguration geladen, Funktionskontrolle ok | Software XP DMS910 und Analysator SW 6.16 neu installiert, Konfiguration geladen, Funktionskontrolle ok | SWO-167258 | No | Software XP DMS910 and analyzer SW 6.16 reinstalled, configuration loaded, function control ok | Software XP DMS910 and analyzer SW 6.16 reinstalled, configuration loaded, function control ok | Defect | SW | Re install | Re install
ABL825 FLEX | Customer could not scan badges to log on to instrument. Recycled power with no change. Could not find operators in system. Not sure how they got deleted and they were missing. Customer is using Win7 operating system. | Reinstalled Win7 software and reinstalled settings from another instrument to restore the deleted operators so the customer could scan and logon to instrument. Verified instrument operation | D21B9C4AE804DD93 | No | Customer could not scan badges to log on to instrument. Recycled power with no change. Could not find operators in system. Not sure how they got deleted and they were missing. Customer is using Win7 operating system. | Reinstalled Win7 software and reinstalled settings from another instrument to restore the deleted operators so the customer could scan and logon to instrument. Verified instrument operation | Setup | SW | Win7 | Re install
ABL827 FLEX |  | B | D218CBCEFC27C91C | No |  |  | Cleaned | Electrode | Electrode | pCO2
ABL815 FLEX | Lac nach Anweisung RMED membraniert, Servicedump gezogen, Kal u. QC i.O. | Lac nach Anweisung RMED membraniert, Servicedump gezogen, Kal u. QC i.O. | SWO-166878 | No | Lac according to instruction RMED membraneed, service dump drawn, Kal and QC i.O. | Lac according to instruction RMED membraneed, service dump drawn, Kal and QC i.O. | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | checked data logs for errors, checked lamp unit and hemolyzer, | performed complete oxi test, test failed neon intensity and neon position, removed Oxi module and replaced Oxi board and spectrometer, adjusted lamp and performed complete oxi test, ran tHb cal | D21948BDFC434862 | No | checked data logs for errors, checked lamp unit and hemolyzer, | performed complete oxi test, test failed neon intensity and neon position, removed Oxi module and replaced Oxi board and spectrometer, adjusted lamp and performed complete oxi test, ran tHb cal | Defect | HW | Oxi module | PCB, Oxi
AQT90 FLEX |  | B | SWO-166893 | No |  |  | Defect | HW | Incubator module | PCB, Incubator
AQT90 FLEX |  | B | D21900A7FC45472D | No |  |  | Cleaned | HW | Aspiration Module | Sensor, Vacutainer
ABL80 CO-OX | I checked the logs and saw a few N but then it came in on the next run | I replaced the connector board and set the zero | D219F43C724D4B7C | No | I checked the logs and saw a few N but then it came in on the next run | I replaced the connector board and set the zero
ABL90 FLEX | ELD wurde erneuert und die Luftblasen sind verschwunden und somit die Fehlermeldung. | ELD wurde erneuert und die Luftblasen sind verschwunden und somit die Fehlermeldung. | SWO-167064 | No | ELD has been renewed and the air bubbles have disappeared and thus the error message. | ELD has been renewed and the air bubbles have disappeared and thus the error message. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Analysator zerlegt und Fehlersuche durchgeführt. Fließstrecke untersucht. Kein Fehler gefunden Analysator zusammengesetzt und wieder hochgefahren. Fehlermeldung war nicht mehr vorhanden. | Analysator zerlegt und Fehlersuche durchgeführt. Fließstrecke untersucht. Kein Fehler gefunden Analysator zusammengesetzt und wieder hochgefahren. Fehlermeldung war nicht mehr vorhanden. | SWO-167145 | Yes | Analyzer disassembled and troubleshooting performed. flow line. No error found analyzer assembled and restarted. Error message no longer existed. | Analyzer disassembled and troubleshooting performed. flow line. No error found analyzer assembled and restarted. Error message no longer existed. | Re classified RMED | Same day upSWO
ABL90 FLEX PLUS | Einlassdichtung erneuert. Fehler beseitigt. Einlaufverfahren durchgeführt. | Einlassdichtung erneuert. Fehler beseitigt. Einlaufverfahren durchgeführt. | SWO-167734 | No | Inlet seal renewed. Fixed a bug. run-in procedures. | Inlet seal renewed. Fixed a bug. run-in procedures. | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D21A32E5720C6890 | No |  |  | Defect | HW | Basic Unit | Wet sec control PCB
ABL835 FLEX | モジュール引上げ。 | B | SWO-168072 | No | Module pull up. |  | Defect | HW | Oxi module | Oxi module | tHb
ABL835 FLEX | MEASURING CHAMBER TEMİZLENDİ. AYARLAR KONTROL EDİLDİ. SORUNLU OLAN pH ELEKTRODU DEĞİŞTİRİLDİ. SORUN GİDERİLDİ. | MEASURING CHAMBER TEMİZLENDİ. AYARLAR KONTROL EDİLDİ. SORUNLU OLAN pH ELEKTRODU DEĞİŞTİRİLDİ. SORUN GİDERİLDİ. | SWO-167566 | Yes | MEASURING CHAMBER IS SECURED. THE SETTINGS HAVE BEEN CHECKED. THE PROBLEMATIC pH ELECTRODE HAS BEEN REPLACED. THE PROBLEM HAS BEEN SENT. | MEASURING CHAMBER IS SECURED. THE SETTINGS HAVE BEEN CHECKED. THE PROBLEMATIC pH ELECTRODE HAS BEEN REPLACED. THE PROBLEM HAS BEEN SENT.
ABL825 FLEX |  | B | D21A00B8722B32D6 | Yes |  |  | Re classified RMED | PM
ABL825 FLEX | inlet probe clip broken | replace inlet holder and clip | D21A1D4F72247888 | Yes | inlet probe clip broken | replace inlet holder and clip
ABL825 FLEX | FlexQ-Modul erneuert. | FlexQ-Modul erneuert. | SWO-167178 | No | FlexQ module is renewed. | FlexQ module is renewed. | Defect | HW | FlexQ Module | FlexQ Module
ABL837 FLEX | A complete software reinstallation was done to see if this resolves the problem. If the problems still exist after this then the PC module will be replaced. Also while on-site reinstalled the Live Connect certificate on SCBU, ED Majors, | I asked Mori to swap the display with one from one of the ABL827s decommissioned yesterday and then monitor the analyser to see if the problem persists. | SWO-167484 | No | A complete software reinstallation was done to see if this resolves the problem. If the problems still exist after this then the PC module will be replaced. Also while on-site reinstalled the Live Connect certificate on SCBU, ED Majors, | I asked Mori to swap the display with one from one of the ABL827s decommissioned yesterday and then monitor the analyser to see if the problem persists. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Checked the power supply, monitor and CPU. | Replaced the CPU unit. Analyzer started up with no issues or errors. Performed a temp. shut down and restarted with no issues. | D21A041F72467C8E | No | Checked the power supply, monitor and CPU. | Replaced the CPU unit. Analyzer started up with no issues or errors. Performed a temp. shut down and restarted with no issues. | Defect | HW | DMS_PC | CPU
AQT90 FLEX | System überprüft und Kundengespräch durchgeführt. Kein Fehler feststellbar. | System überprüft und Kundengespräch durchgeführt. Kein Fehler feststellbar. | SWO-167242 | No | system and customer call performed. No error detectable. | system and customer call performed. No error detectable. | Defect | SW | Analyzer Control Module | Off/On
ABL80 FLEX BASIC | 清洁扫描口后修复,but don't change the parts(By Ran) | 清洁扫描口后修复,but don't change the parts(By Ran) | SWO-166770 | No | Repair after cleaning the scan, but don't change the parts (By Ran) | Repair after cleaning the scan, but don't change the parts (By Ran)
ABL80 FLEX CO-OX | Fehlersuche, Befestigung Lösungsschlauch, Installation Lösungspack, Funktionskontrolle | FSE-Einsatz erforderlich | SWO-167409 | No | Troubleshooting, fixing solution hose, installation solution pack, function control | FSE deployment required
ABL90 FLEX | Removed archived data logs to free disk space. | Removed archived data logs to free disk space. | SWO-167670 | No | Removed archived data logs to free disk space. | Removed archived data logs to free disk space. | Adjustment | SW | DMS_PC | Disc Space | temp files deleted
ABL90 FLEX | Datenspeicher ausgewertet, Fehler treten teilweise nach Probenaspiration gescheitert auf, mit den Anwendern im Labor über Präanalytik gesprochen. SC aus Kundenbestand ersetzt. Reklamation an RDEU gemailt | Datenspeicher ausgewertet, Fehler treten teilweise nach Probenaspiration gescheitert auf, mit den Anwendern im Labor über Präanalytik gesprochen. SC aus Kundenbestand ersetzt. Reklamation an RDEU gemailt | SWO-167743 | No | Data storage evaluated, errors occur partly after sample aspiration failed, talked to the users in the laboratory about preanalysis. SC replaced from customer base. Complaint to RDEU emailed | Data storage evaluated, errors occur partly after sample aspiration failed, talked to the users in the laboratory about preanalysis. SC replaced from customer base. Complaint to RDEU emailed | Unknown | Unknown | Unknown | Unknown
ABL90 | SE EFECTUAN PRUEBAS CON DIFERENTES REACTIVOS Y LIMPIEZA DE TODO EL SISTEMA NO RESOLVIENDO EL PROBLEMA, SE DESMONTA PARCIALMENTE EL EQUIPO Y SE SUSTITUYE TUBO DE HEMOLIZADOR BAHIA, TRAS DIFERENTES PRUEBAS Y AJUSTES QUEDA TODO SOLUCIONADO. CALIBRADO Y VERFIFICADO CON Q.C. | B | D21B3A94E83BB106 | No | TESTS ARE EFFECTED WITH DIFFERENT REACTIVOS AND CLEANING OF THE ENTIRE SYSTEM DO NOT RESOLVE THE PROBLEM, THE EQUIPMENT IS PARTIALLY DISASSEMBLED AND BAHIA HEMOLIZER TUBE IS SUBSTITUTE, AFTER DIFFERENT TESTS AND ADJUSTMENTS IS ALL SOLVED. CALIBRATED AND VERFIFIED WITH Q.C. |  | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-167784 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Reference | 1361
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-167405 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Temperature | 166
ABL800 BASIC | Replaced reference electrode. Cleaned measuring chamber and electrode latch. | Removed new reference electrode. Installed customer unit. | SWO-169509 | No | Replaced reference electrode. Cleaned measuring chamber and electrode latch. | Removed new reference electrode. Installed customer unit. | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX |  | B | D21C7D635E5E5E56 | No |  |  | Assisted Replacement | Electrode | Electrode | Ca
ABL835 FLEX | Examined display cable and assembly. | Reloaded windows 7 and application software. Verified operation of ABL835. | D21B5758E84CEF5A | No | Examined display cable and assembly. | Reloaded windows 7 and application software. Verified operation of ABL835. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX |  | B | D21B5166E8171ED7 | No |  |  | Defect | HW | Basic Unit | Gas module
AQT90 FLEX |  | B | D21B4ADFE85EE920 | Yes
AQT90 FLEX | Barcodescanner im Servicemenu nicht ansprechbar. Main-Board bestellt. | Mainboard und Scannereinheit erneuert. Software neu installiert. Dateien zurückgespielt. Qualitätskontrollen durchgeführt. Gerät überprüft. | SWO-168255 | No | barcode scanner in the service menu is not responsive. Main-Board. | mainboard and scanner unit. Software reinstalled. Files played back. quality controls carried out. device checked. | Revisit required | HW | Analyzer Control Module | PCB, Main Board
ABL80 FLEX CO-OX | Fehlersuche, Installation neue Sensorkassette, Funktionskontrolle | FSE-Einsatz erforderlich | SWO-167807 | No | Troubleshooting, installation of new sensor cartridge, function control | FSE deployment required
ABL90 FLEX | Erneuerung der Sensorkassette gegen konditionierte Ware. Die ausgetauschte Ware wird getestet | B | SWO-167953 | No | Renewal of the sensor cassette against conditioned goods. The exchanged goods are tested |  | Assisted Replacement | SC | SC | Lac | 131
ABL90 FLEX | Unable to resolve during this visit. | Unable to resolve during this visit. | SWO-167960 | No | Unable to resolve during this visit. | Unable to resolve during this visit. | Revisit required | HW | Power Supply | Power Supply
ABL90 FLEX | Netzteil erneuert. | Netzteil erneuert. | SWO-168189 | No | power supply is renewed. | power supply is renewed. | Defect | HW | Power Supply | Power Supply
ABL90 | Customer had resolved issue upon arrival. Examined solution pack loading mechanism. | Reinstalled solution pack. Unable to replicate scenario. Verified operation of ABL90. Instructed user on loading of solution pack. | D21C9FFE5E134583 | Yes | Customer had resolved issue upon arrival. Examined solution pack loading mechanism. | Reinstalled solution pack. Unable to replicate scenario. Verified operation of ABL90. Instructed user on loading of solution pack.
ABL90 FLEX | Jw Kit eingebaut, Gerät forderte anschließend keine neuen LP an. OXI Bügelventil demontiert und gereinigt, Sensor Kassette erneuert. Gerät ausgeliefert. | Jw Kit eingebaut, Gerät forderte anschließend keine neuen LP an. OXI Bügelventil demontiert und gereinigt, Sensor Kassette erneuert. Gerät ausgeliefert. | SWO-167914 | No | Jw Kit installed, device then did not requested a new LP. OXI ironing valve disassembled and cleaned, sensor cassette renewed. device. | Jw Kit installed, device then did not requested a new LP. OXI ironing valve disassembled and cleaned, sensor cassette renewed. device. | Defect | HW | Oxi module | Pinch valve
ABL90 | Staff was on phone with tech support and the issue could not be fixed.. | Arrived on site with note on analyzer that states the printer is not working . Opened printer door to verify the correct paper and correct installation of paper. Closed door and tested with no errors or issues. The printer worked as it should. Staff will continue to monitor. | D21C63715E190CB7 | No | Staff was on phone with tech support and the issue could not be fixed.. | Arrived on site with note on analyzer that states the printer is not working . Opened printer door to verify the correct paper and correct installation of paper. Closed door and tested with no errors or issues. The printer worked as it should. Staff will continue to monitor. | No fault found | HW | Printer Module | Printer Module
ABL90 FLEX PLUS | Removed interface PCB and removed and cleaned contact pins. Refitted and test ok Glucose passing calibration with good sensitivity and status. Run QC all passed ok. | Removed interface PCB and removed and cleaned contact pins. Refitted and test ok Glucose passing calibration with good sensitivity and status. Run QC all passed ok. | SWO-167963 | No | Removed interface PCB and removed and cleaned contact pins. Refitted and test ok Glucose passing calibration with good sensitivity and status. Run QC all passed ok. | Removed interface PCB and removed and cleaned contact pins. Refitted and test ok Glucose passing calibration with good sensitivity and status. Run QC all passed ok. | Cleaned | HW | Sensor Compartment | Contacts
ABL90 FLEX PLUS |  | B | D21C7D9C5E2C1CEC | No |  |  | Defect | HW | CF Card | CF Card
ABL90 FLEX PLUS | Kurzeinweisung, neue Bestellisten verteilt, Einlsassnadel ausgerichtet, Zählerstände und Lagerbestände aller Geräte im Haus überprüft, insg nur 3 SP90 im Haus, Funktionskontrolle | Kurzeinweisung, neue Bestellisten verteilt, Einlsassnadel ausgerichtet, Zählerstände und Lagerbestände aller Geräte im Haus überprüft, insg nur 3 SP90 im Haus, Funktionskontrolle | SWO-167906 | No | Short briefing, new orderists distributed, single-seat needle aligned, meter readings and stock of all appliances in the house checked, in total only 3 SP90 in the house, function control | Short briefing, new orderists distributed, single-seat needle aligned, meter readings and stock of all appliances in the house checked, in total only 3 SP90 in the house, function control | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-168124 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Glu/Lac
ABL90 FLEX |  | B | SWO-167700 | No |  |  | Defect | HW | Basic Unit | Cassette retainer | 529
AQT90 FLEX |  | B | SWO-167712 | No |  |  | Revisit required | HW | Aspiration Module | PCB, Vacu vision
AQT90 FLEX |  | B | SWO-167726 | No |  |  | Defect | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX | 自行更换备用机件910-299后修复 | B | SWO-167413 | No | Repair after replacing spare parts 910-299 |  | Defect | HW | Analyzer Control Module | DMS_PC
ABL800 BASIC | Removed new reference electrode. Installed customer unit. | Removed new reference electrode. Installed customer unit. | SWO-169516 | Yes | Removed new reference electrode. Installed customer unit. | Removed new reference electrode. Installed customer unit. | Re classified RMED | Courtesy/Follow up
ABL835 FLEX | Susbtituição da sonda , descontaminação do circuito, limpeza das camaras de medida, substituição das membranas, verificação do estado de funcionamento, estabilização e passagem do QC. | B | D21C9EAE5E3ECDB4 | Yes | Replacement of the probe, decontamination of the circuit, cleaning of the measurement chambers, replacing the membranes, checking the state of operation, stabilization and passage of the QC. |  | Re classified RMED | PM
AQT90 FLEX | Analysator geöffnet. Probenröhrchen im Analysator. Nadel stark verbogen. Austausch der Nadeleinheit. Prüfung im Servicemenü. Inlet Door öffnet auf der linken Seite zeitweise nicht. Fehlende Sicherungsringe an Motorwelle. Übergangsweise mit Schlauchstück fest fixiert. Sicherungsringe werden bestellt und eingebaut. Initialisierungsphase Rücksetzen des Nadelzählers QC durch Labor | Analysator geöffnet. Probenröhrchen im Analysator. Nadel stark verbogen. Austausch der Nadeleinheit. Prüfung im Servicemenü. Inlet Door öffnet auf der linken Seite zeitweise nicht. Fehlende Sicherungsringe an Motorwelle. Übergangsweise mit Schlauchstück fest fixiert. Sicherungsringe werden bestellt und eingebaut. Initialisierungsphase Rücksetzen des Nadelzählers QC durch Labor | SWO-167910 | No | Analyzer open. Sample tubes in the analyzer. Needle strongly bent. Replacement of the needle unit. Check in the service menu. Inlet Door does not open temporarily on the left. Missing locking rings on motor shaft. Temporarily fixed with hose piece. Fuse rings are ordered and installed. Initialization phase Reset the needle counter QC by Laboratory | Analyzer open. Sample tubes in the analyzer. Needle strongly bent. Replacement of the needle unit. Check in the service menu. Inlet Door does not open temporarily on the left. Missing locking rings on motor shaft. Temporarily fixed with hose piece. Fuse rings are ordered and installed. Initialization phase Reset the needle counter QC by Laboratory | Defect | HW | Aspiration Module | Bend needle
AQT90 FLEX | Mainboard und Scannereinheit erneuert. Software neu installiert. Dateien zurückgespielt. Qualitätskontrollen durchgeführt. Gerät überprüft. | Mainboard und Scannereinheit erneuert. Software neu installiert. Dateien zurückgespielt. Qualitätskontrollen durchgeführt. Gerät überprüft. | SWO-168254 | No | mainboard and scanner unit. Software reinstalled. Files played back. quality controls carried out. device checked. | mainboard and scanner unit. Software reinstalled. Files played back. quality controls carried out. device checked. | Defect | HW | Analyzer Control Module | PCB, Main Board | 998
AQT90 FLEX | Das besagte Lösungspack nicht in den Daten gefunden. In den Erneuerungen vom letzten Jahr ein XR-52 LP gefunden, Verfallsdatum 14.12.16. Kunde hat einen verfallenen LP installiert. | Das besagte Lösungspack nicht in den Daten gefunden. In den Erneuerungen vom letzten Jahr ein XR-52 LP gefunden, Verfallsdatum 14.12.16. Kunde hat einen verfallenen LP installiert. | SWO-167983 | No | The said solution pack not found in the data. In the renewals of last year found an XR-52 LP, expiry date 14.12.16. Customer has installed a forfeited LP. | The said solution pack not found in the data. In the renewals of last year found an XR-52 LP, expiry date 14.12.16. Customer has installed a forfeited LP. | User error | Consumables | Consumables | SP expired
HbA1c 501 Analyzer | 代替機貸出。 | 代替機貸出。 | SWO-167971 | Yes | Alternative machine rental. | Alternative machine rental.
AQT90 FLEX |  | B | SWO-167985 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL800 BASIC | In calibration Na+,K+ shows calibration unstable.In system message shows El met upper liquid sensor error.Check and clean liquid sensor,calibrate which is Ok except Na+calibration is Ok.But in sample Glu & Lac result shows very low.Check measuring valve and found the problem in air valve membrane which is need to be replace to resolve the problem. | In calibration Na+,K+ shows calibration unstable.In system message shows El met upper liquid sensor error.Check and clean liquid sensor,calibrate which is Ok except Na+calibration is Ok.But in sample Glu & Lac result shows very low.Check measuring valve and found the problem in air valve membrane which is need to be replace to resolve the problem. | SWO-170050 | No | In calibration Na+,K+ shows calibration unstable. In system message shows El met upper liquid sensor error. Check and clean liquid sensor,calibrate which is Ok except Na+calibration is Ok.But in sample Glu & Lac result shows very low. Check measuring valve and found the problem in air valve membrane which is need to be replace to resolve the problem. | In calibration Na+,K+ shows calibration unstable. In system message shows El met upper liquid sensor error. Check and clean liquid sensor,calibrate which is Ok except Na+calibration is Ok.But in sample Glu & Lac result shows very low. Check measuring valve and found the problem in air valve membrane which is need to be replace to resolve the problem. | Revisit required | HW | El/Met Module | Valve, air
ABL80 FLEX BASIC |  | B | SWO-168019 | No
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-168094 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX PLUS | Drucküberprüfung Vakuum Fehler, Druck Überprüfung wird nicht bestanden. Fließwege gereinigt - keine Verstopfung. LP entfernt und das Gerät neu gestartet - Fehler besteht weiterhin. Fließwege auseinander gebaut, separat gespült und gereinigt. Drucktests im Serviceprogramm durchgeführt - Fehler besteht weiterhin. Einlassdichtung verschlissen - zieht Luft mit rein. Einlassdichtung nicht vor Ort und nicht im Auto. Reparatur nicht möglich. Ersatzteil wird morgen organisiert und eingebaut. | B | SWO-168129 | No | Pressure verification vacuum error, pressure check is not passed. Flow paths cleaned - no constipation. REMOVED LP and the device restarted - error still exists. Flow paths disassembled, rinsed separately and cleaned. Print tests performed in the service program - errors persist. Inlet seal worn out - draws air with it. Inlet seal not on site and not in the car. Repair not possible. Spare part will be organized and installed tomorrow. |  | Revisit required | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX |  | B | SWO-168018 | No |  |  | Revisit required | HW | Basic Unit | Sensor Interface PCB | 964
ABL90 FLEX |  | B | SWO-168026 | No |  |  | Defect | HW | Basic Unit | Wet sec control PCB
AQT90 FLEX | 更换905-953和903-186修复(By Ran) For this SWO, we didn't find 903-186 in 905-953 Kit. | 更换905-953和903-186修复(By Ran) For this SWO, we didn't find 903-186 in 905-953 Kit. | SWO-167981 | No | 更换905-953和903-186修复(By Ran) For this SWO, we didn't find 903-186 in 905-953 Kit. | 更换905-953和903-186修复(By Ran) For this SWO, we didn't find 903-186 in 905-953 Kit. | Defect | HW | Aspiration Module | PCB, Vacu vision
ABL827 FLEX | Réinstallation complète du software sytème et XP (main board 910) e.o Réinstallation de la section fluidique (analyzer control PCB) e.o Introduction complète manuelle de la configuration e.o Calibration e.o contrôle de qualité e.o | Réinstallation complète du software sytème et XP (main board 910) e.o Réinstallation de la section fluidique (analyzer control PCB) e.o Introduction complète manuelle de la configuration e.o Calibration e.o contrôle de qualité e.o | SWO-169549 | No | Complete system and XP (main board 910) e.o software reinstallation Reinstallation of the fluidic section (analyzer control PCB) e.o Full manual introduction of the e.o configuration Calibration e.o quality control e.o | Complete system and XP (main board 910) e.o software reinstallation Reinstallation of the fluidic section (analyzer control PCB) e.o Full manual introduction of the e.o configuration Calibration e.o quality control e.o | Defect | HW | Basic Unit | Analyzer control PCB
ABL80 FLEX BASIC |  | B | SWO-168104 | No
ABL80 CO-OX | Verified SC Smartchip Corrupt error. | Replaced with new SC80 CO-OX. | D2207ED3C013F0E4 | Yes | Verified SC Smartchip Corrupt error. | Replaced with new SC80 CO-OX.
ABL80 FLEX BASIC | Activated Ca parameter. Given two point calibration. Observed that sensor board was faulty. Need to replace new sensor Board. | Activated Ca parameter. Given two point calibration. Observed that sensor board was faulty. Need to replace new sensor Board. | SWO-169888 | No | Activated Ca parameter. Given two point calibration. Observed that sensor board was faulty. Need to replace new sensor Board. | Activated Ca parameter. Given two point calibration. Observed that sensor board was faulty. Need to replace new sensor Board.
ABL80 CO-OX |  | FOLLOWED SERVICE TIP 624208893 TO EDIT CHANGES IN ASTM.INI FILES. | D22080F0C0412180 | No |  | FOLLOWED SERVICE TIP 624208893 TO EDIT CHANGES IN ASTM. INI FILES.
ABL90 FLEX | Lösungspack erneuert. Danach alles QC in Ordnung. Gerät überprüft. | Lösungspack erneuert. Danach alles QC in Ordnung. Gerät überprüft. | SWO-169305 | No | Solution pack renewed. After that, all QC fine. device checked. | Solution pack renewed. After that, all QC fine. device checked. | Assisted Replacement | SP | SP | pO2
ABL90 FLEX | Replaced CPU and fitted original CF card. Powered on unit and CPU is working correctly with the customer's original CF card. | Unable to resolve during this visit. | SWO-169304 | No | Replaced CPU and fitted original CF card. Powered on unit and CPU is working correctly with the customer's original CF card. | Unable to resolve during this visit. | Defect | HW | DMS_PC | CPU
ABL90 FLEX | Sensorkassette mitgebracht und installiert. Lösungspack zur Reserve ebenfalls vor Ort gelassen. Die Kundin reklamiert, dass nicht klar genug auf die kurze Haltbarkeit der Sensorkassetten und Lösungspacks hingewiesen wurde und erwartet daher kulanterweise kostenlosen Ersatz für 1 ST Sensorkassette und 1 ST Lösungspacks. Die Haltbarkeiten der beiden Produkte wurden besprochen. | Sensorkassette mitgebracht und installiert. Lösungspack zur Reserve ebenfalls vor Ort gelassen. Die Kundin reklamiert, dass nicht klar genug auf die kurze Haltbarkeit der Sensorkassetten und Lösungspacks hingewiesen wurde und erwartet daher kulanterweise kostenlosen Ersatz für 1 ST Sensorkassette und 1 ST Lösungspacks. Die Haltbarkeiten der beiden Produkte wurden besprochen. | SWO-168325 | No | sensor cartridge and installed. Solution pack to the reserve also left on site. The customer complains that the short shelf life of the sensor cartridges and solution packs has not been clearly enough and therefore expects free replacements for 1 ST sensor cassette and 1 ST solution packs. The durability of the two products was discussed. | sensor cartridge and installed. Solution pack to the reserve also left on site. The customer complains that the short shelf life of the sensor cartridges and solution packs has not been clearly enough and therefore expects free replacements for 1 ST sensor cassette and 1 ST solution packs. The durability of the two products was discussed. | User error | SC Expired | SC Expired | SC Expired
ABL90 FLEX PLUS | Oxi- Schlauchventil ausgetauscht.eine kompletteJahreswartung durchgeführt. | Zusammen den ABL ohne LP neu gestartet. LP wieder installiert, Startphase von 55Min. läuft. | SWO-169372 | No | Oxi hose valve replaced.a complete annual maintenance carried out. | Restarted together the ABL without LP. LP reinstalled, start-up phase of 55min. | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX PLUS | Einlassdichtung undicht - erneuert. Gerät ohne LP. Es gestartet. Sensorkassette muss wieder einlaufen aufgrund der langen Standzeit. Empfindlichkeit von pCO2 kommt nicht in den Sollbereich. Sensorkassette erneuert. Neue Sensorkassette muss einlaufen, ich melde mich telefonisch und frage den Status ab. | B | SWO-168279 | No | Inlet seal leaks - renewed. Device without LP. It started. Sensor cartridge must re-enter due to the long service life. Sensitivity of pCO2 does not reach the target range. Sensor cartridge is renewed. New sensor cartridge has to come in, I log in by phone and check the status. |  | Defect | HW | Auto Inlet module | Inlet Gasket | Leaking
ABL90 FLEX PLUS |  | B | D2205171C051BCC9 | No |  |  | Assisted Replacement | SP | SP | pO2
ABL90 | Error en la base de datos se realizan reinicios, se reinstala la versión del programa borrando la base de datos, chequeo de funcionamiento general, instalación completa de consumibles, calibración general, test conntroles de calidad he introdución de muestra, funcionando correctamente. | Borrado de datos y carga de programa. | D22041B6C01FC151 | Yes | Database errors are performed restarts, program version is reinstalled by erasing database, general operation check, complete installation of consumables, general calibration, quality conntroles test I sample introduction, working properly. | Data erasure and program load.
ABL90 FLEX |  | B | SWO-168053 | No |  |  | Revisit required | HW | Basic Unit | Sensor Interface PCB | 964
ABL90 FLEX |  | 借的903-183缺货取消，改从广州仓库发货。需要重开SWO by andrew | SWO-168079 | No |  | Borrowed 903-183 out of stock cancellation, instead from the Warehouse in Guangzhou. Need to reopen SWO by Andrew | Revisit required | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-168083 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | pO2
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-168082 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Lac | 21
TCM5 BASIC | Removed sensor 92-67029 to workshop for investigation Demo monitor returned to workshop and customer monitor returned | Removed sensor 92-67029 to workshop for investigation Demo monitor returned to workshop and customer monitor returned | SWO-168247 | No | Removed sensor 92-67029 to workshop for investigation Demo monitor returned to workshop and customer monitor returned | Removed sensor 92-67029 to workshop for investigation Demo monitor returned to workshop and customer monitor returned
AQT90 FLEX |  | B | SWO-168063 | No |  |  | Adjustment | Consumables | Test Kit | CKMB | Calibration
AQT90 FLEX |  | 更换636-550后修复 by jesson | SWO-168158 | No |  | Repair by jesson after replacement 636-550 | Defect | HW | Process Module | Process Unit sensor | 5257
AQT90 FLEX |  | B | SWO-168109 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-168126 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL825 FLEX |  | B | D2207C9DC029C2CD | No |  |  | Defect | HW | Inlet module | Inlet module
ABL800 BASIC | Replace the air valve membrane by another instrument for temporary solution.Run LS adjust & Pump Calibration and run 2-point calibration which is Ok.then run OC which is also OK and run some sample which result is satisfactory. | In calibration Na+,K+ shows calibration unstable.In system message shows El met upper liquid sensor error.Check and clean liquid sensor,calibrate which is Ok except Na+calibration is Ok.But in sample Glu & Lac result shows very low.Check measuring valve and found the problem in air valve membrane which is need to be replace to resolve the problem. | SWO-170054 | No | Replace the air valve membrane by another instrument for temporary solution. Run LS adjust & Pump Calibration and run 2-point calibration which is Ok.then run OC which is also OK and run some sample which result is satisfactory. | In calibration Na+,K+ shows calibration unstable. In system message shows El met upper liquid sensor error. Check and clean liquid sensor,calibrate which is Ok except Na+calibration is Ok.But in sample Glu & Lac result shows very low. Check measuring valve and found the problem in air valve membrane which is need to be replace to resolve the problem. | Defect | HW | El/Met Module | Valve, air
ABL825 FLEX | Sporadische Aussetzer bzw. Grenzverletzungen bei MetHb und HHb => Proteinentfernung, Kal, OxiCompleteTest, zweimal hintereinander Hemolyzer Phase Regulation Failure => Hemolyzerunit ersetzt, komplett abgeglichen, Service-tHb-Kal, Kal, Funktionstest und QC Messung, OK. | Sporadische Aussetzer bzw. Grenzverletzungen bei MetHb und HHb => Proteinentfernung, Kal, OxiCompleteTest, zweimal hintereinander Hemolyzer Phase Regulation Failure => Hemolyzerunit ersetzt, komplett abgeglichen, Service-tHb-Kal, Kal, Funktionstest und QC Messung, OK. | SWO-169349 | No | Sporadic dropouts or boundary injuries in MetHb and HHb => Protein Removal, Kal, OxiCompleteTest, twice in a row Hemolyzer Phase Regulation Failure => Hemolyzerunit replaced, completely matched, Service-tHb-Kal, Kal, Function Test and QC Measurement, OK. | Sporadic dropouts or boundary injuries in MetHb and HHb => Protein Removal, Kal, OxiCompleteTest, twice in a row Hemolyzer Phase Regulation Failure => Hemolyzerunit replaced, completely matched, Service-tHb-Kal, Kal, Function Test and QC Measurement, OK. | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Einlassnadel durchgängig nach oben verbogen => gerade gerichtet. | Einlassnadel durchgängig nach oben verbogen => gerade gerichtet. | SWO-169355 | No | Inlet pin bent consistently upwards => straight enlongated. | Inlet pin bent consistently upwards => straight enlongated. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Grüne LED in FlexQ Position 3 ohne Funktion. Ersatzteil bestellt, wird eingebaut sofort nach Lieferung. Drei Dichtringe von FAN915-318 nachgetragen. | Grüne LED in FlexQ Position 3 ohne Funktion. Ersatzteil bestellt; wird eingebaut sofort nach Lieferung. Drei Dichtringe von FAN915-318 nachgetragen. | SWO-169357 | No | Green LED in FlexQ position 3 without function. spare part ordered; will be installed immediately after delivery. Three sealing rings of FAN915-318. | Green LED in FlexQ position 3 without function. spare part ordered; will be installed immediately after delivery. Three sealing rings of FAN915-318. | Revisit required | HW | FlexQ Module | Position-LED | Known issue AMF
ABL837 FLEX |  | B | D220E9AFC0332789 | No |  |  | Defect | HW | Inlet module | Probe bend
ABL827 FLEX | Upon arrival, verified above issue with unit Install. Found unit Exhibiting Neon Lamp low Intensity. Verified Failure in Oximetry service test. | Replacec above parts listed. Performed all service Oximetry alignment tests. All tests passed. | D2217879362FF541 | No | Upon arrival, verified above issue with unit Install. Found unit Exhibiting Neon Lamp low Intensity. Verified Failure in Oximetry service test. | Replacec above parts listed. Performed all service Oximetry alignment tests. All tests passed. | Defect | HW | Oxi module | Spectrometer
AQT90 FLEX | Fehlerhafte Parameterkassette mit abgebrochener Kunststoffführung entnommen und ersetzt. Justagen und Kalibrationen durchgeführt. Systemüberprüfung und Funktionstests in Ordnung. | Fehlerhafte Parameterkassette mit abgebrochener Kunststoffführung entnommen und ersetzt. Justagen und Kalibrationen durchgeführt. Systemüberprüfung und Funktionstests in Ordnung. | SWO-168169 | No | Incorrect parameter cassette with broken plastic guide removed and replaced. adjustments and calibrations. System verification and functional tests are in order. | Incorrect parameter cassette with broken plastic guide removed and replaced. adjustments and calibrations. System verification and functional tests are in order. | Defect | HW | Storage Module | Cartridge guide
AQT90 FLEX | Defecte cartridge verwijderd. | Defecte cartridge verwijderd. | SWO-168317 | No | Defective cartridge removed. | Defective cartridge removed. | Defect | Consumables | Test Kit | Cartridge | 5689 (Broken)
ABL9 | Gerät überprüft, Patch eingespielt, Datum und Uhrzeit korrigiert, verfallenes Verbrauchsmaterial ersetzt, QC außerhalb Grenzen, da das Lot keine ABL9 Grenzwerte hat | B | SWO-169649 | No | device checked, patch checked, date and time corrected, expired consumable sumer, QC outside limits as the lot has no ABL9 limits
ABL90 |  | B | D221DF83364ED610 | Yes |  |  | Re classified RMED | PM
ABL90 FLEX | Einlassdichtung erneuert. Gerät im Servicemenü überprüft. Pumpenfehler war erst nach einer Liquidsensorkalibrierung wieder in Ordnung. Alle Test fehlerfrei, danach Neustart und Kalibrierung. Anwender hat die falschen Kapillaren bestellt, die nicht für Elektrolytmessungen geeignet sind. Gerät überprüft | Einlassdichtung erneuert. Gerät im Servicemenü überprüft. Pumpenfehler war erst nach einer Liquidsensorkalibrierung wieder in Ordnung. Alle Test fehlerfrei, danach Neustart und Kalibrierung. Anwender hat die falschen Kapillaren bestellt, die nicht für Elektrolytmessungen geeignet sind. Gerät überprüft | SWO-170025 | No | Inlet seal renewed. device in the service menu. Pump fault was only correct after a liquid sensor calibration. All test error-free, then restart and calibration. User has ordered the wrong capillaries that are not suitable for electrolyte measurements. Device checked | Inlet seal renewed. device in the service menu. Pump fault was only correct after a liquid sensor calibration. All test error-free, then restart and calibration. User has ordered the wrong capillaries that are not suitable for electrolyte measurements. Device checked | Assisted Replacement | SP | SP | FT
ABL90 FLEX | IP Adresse neu konfiguriert und auf Sd Karte gespeichert | IP Adresse neu konfiguriert und auf Sd Karte gespeichert | SWO-169540 | No | IP address reconfigured and stored on Sd card | IP address reconfigured and stored on Sd card | Setup | His/lis | His/lis | His/lis
ABL90 FLEX PLUS | NetOP Einstellungen geprüft, Software Windows und V3.1 MR8 neu installiert, Funktionskontrolle fehlerhaft, Kontaktaufnahme mit der IT Abteilung Herr Molau, es wurde kein Port für den Analysator angelegt, Funktionskontrolle ok | NetOP Einstellungen geprüft, Software Windows und V3.1 MR8 neu installiert, Funktionskontrolle fehlerhaft, Kontaktaufnahme mit der IT Abteilung Herr Molau, es wurde kein Port für den Analysator angelegt, Funktionskontrolle ok | SWO-170245 | No | NetOP settings checked, software Windows and V3.1 MR8 reinstalled, function control faulty, contact with the IT department Mr. Molau, no port was created for the analyzer, function control ok | NetOP settings checked, software Windows and V3.1 MR8 reinstalled, function control faulty, contact with the IT department Mr. Molau, no port was created for the analyzer, function control ok | Setup | Remote Support | Remote Support | Remote Support
AQT90 FLEX |  | B | SWO-169408 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-169433 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
AQT90 FLEX |  | B | SWO-169486 | Yes |  |  | Re classified RMED | PCT not released | PCT not released | PCT not released | upSWO date 21.2
ABL827 FLEX |  | B | D221BFB8365F5FE4 | Yes |  |  | Re classified RMED | FAN
ABL815 FLEX | Lac gem. Anweisung gereinigt und neu membraniert, akz. Kal abgewartet, Kal u. QC i.O. Servicedump | Lac gem. Anweisung gereinigt und neu membraniert, akz. Kal abgewartet, Kal u. QC i.O. Servicedump | SWO-169879 | No | Lac according to instructions cleaned and re-membraneed, akz. Kal waited, Kal & QC i.O. Service dump | Lac according to instructions cleaned and re-membraneed, akz. Kal waited, Kal & QC i.O. Service dump | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | QC Flachbandkabel 14 polig erneuert, kein Erfolg, gleiches Verhalten. Vermutlich Motor oder Motoransteuerung defekt. Kabel zurückgetauscht. Vorübergehend AQC Modul aus Demogerät GS17 eingebaut bis Kundenmodul repariert werden kann. | QC Flachbandkabel 14 polig erneuert, kein Erfolg, gleiches Verhalten. Vermutlich Motor oder Motoransteuerung defekt. Kabel zurückgetauscht. Vorübergehend AQC Modul aus Demogerät GS17 eingebaut bis Kundenmodul repariert werden kann. | SWO-170676 | No | QC ribbon cable 14 poles renewed, no success, same behavior. Probably engine or motor control defective. cable replaced. Temporarily AQC module from demo device GS17 installed until customer module can be repaired. | QC ribbon cable 14 poles renewed, no success, same behavior. Probably engine or motor control defective. cable replaced. Temporarily AQC module from demo device GS17 installed until customer module can be repaired. | Revisit required | HW | Auto Check Module | Auto Check Module
ABL800 BASIC | cleaned measuring chamber. run protien removal. given two point calibration. run control. | cleaned measuring chamber. run protien removal. given two point calibration. run control. | SWO-170657 | No | cleaned measuring chamber. run protien removal. given two point calibration. run control. | cleaned measuring chamber. run protien removal. given two point calibration. run control. | Cleaned | Fluidic pathway | Fluidic pathway | Ca
ABL837 FLEX |  | B | D2231B35AC14F975 | No |  |  | Defect | HW | Basic Unit | Analyzer control PCB
ABL825 FLEX | Customer has tried a couple of Electrode without any good results Also, different membranes | I installed a known good module pcb and calibrated unit until all errors were gone Customer will monitor the problem. | D221AA8B3635308F | No | Customer has tried a couple of Electrode without any good results Also, different membranes | I installed a known good module pcb and calibrated unit until all errors were gone Customer will monitor the problem. | Defect | HW | El/Met Module | PCB, El/Met
ABL80 FLEX CO-OX | Hemolyser (R0596N0002) vervangen. tHb kalibratie en Co-ox test OK CalQC : OK Er werden eveneens een 8-tal stalen bepaald Voorlopig afwachten of de problemen hiermee opgelost werden. | Hemolyser (R0596N0002) vervangen. tHb kalibratie en Co-ox test OK CalQC : OK Er werden eveneens een 8-tal stalen bepaald Voorlopig afwachten of de problemen hiermee opgelost werden. | SWO-170303 | No | Hemolyser (R0596N0002) replaced. tHb Calibration and Co-ox test OK CalQC: OK 8 samples were also determined For now, wait and see if the problems were resolved. | Hemolyser (R0596N0002) replaced. tHb Calibration and Co-ox test OK CalQC: OK 8 samples were also determined For now, wait and see if the problems were resolved.
ABL80 FLEX | Cleaned SC and Placed for Cooling and Reinstalled, Na and Hct Parameter Not calibrate after Trying Re installation for Multiple times but problem remains, Replaced SC with New one, Problem resolved. | Cleaned SC and Placed for Cooling and Reinstalled, Na and Hct Parameter Not calibrate after Trying Re installation for Multiple times but problem remains, Replaced SC with New one, Problem resolved. | SWO-170656 | No | Cleaned SC and Placed for Cooling and Reinstalled, Na and Hct Parameter Not calibrate after Trying Re installation for Multiple times but problem remains, Replaced SC with New one, Problem resolved. | Cleaned SC and Placed for Cooling and Reinstalled, Na and Hct Parameter Not calibrate after Trying Re installation for Multiple times but problem remains, Replaced SC with New one, Problem resolved.
ABL80 CO-OX | Checked the data logs and verified the problem. | Replaced the hemolyzer. | D222C1EDAC16BB08 | No | Checked the data logs and verified the problem. | Replaced the hemolyzer.
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-169816 | No | Repair after replacing test card | Repair after replacing test card
AQT90 FLEX | Desmontagem do equipamento, verificação de fuga na P4, substituição do tubos , verificação do funcionamento, tudo OK | B | D2231B4EAC0A9521 | No | Disassembly of the equipment, leakage check on P4, replacing the tubes, checking the operation, everything OK |  | Defect | HW | Aspiration Module | Tubes
ABL825 FLEX | Einlassnadel erneut und offenbar ohne äußere Einwirkung bananenkrumm. Nadel gerichtet. | Einlassnadel erneut und offenbar ohne äußere Einwirkung bananenkrumm. Nadel gerichtet. | SWO-169945 | No | Inlet needle again and apparently without external influence banana-krumm. Needle directed. | Inlet needle again and apparently without external influence banana-krumm. Needle directed. | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Oberer Gummi-Dichtring des K Messplatzes offenbar ab Werk nicht in Position, dadurch Blockage der Elektrode. Messkammer zerlegt und Ring neu positioniert. | Oberer Gummi-Dichtring des K Messplatzes offenbar ab Werk nicht in Position, dadurch Blockage der Elektrode. Messkammer zerlegt und Ring neu positioniert. | SWO-169947 | No | Upper rubber sealing ring of the K measuring station apparently not in position ex works, thus blocking the electrode. measuring chamber disassembled and ring repositioned. | Upper rubber sealing ring of the K measuring station apparently not in position ex works, thus blocking the electrode. measuring chamber disassembled and ring repositioned. | User error | Electrode | Electrode | K+
ABL825 FLEX | Fehlersuche: Fließwege überprüft, Nadel gerichtet (s. a. SWO), Ca sowie Referenz membraniert, pH Dichtung ersetzt (s.a. SWO), mechanische Blockade im K Messplatz beseitigt (s.a. SWO), alle Messkammern gereinigt. Nach Neustart und Kal sowie tHb-Kal Lac Empf. unterhalb Grenze und pCO2 an Grenze => beide membraniert. Mehrere Kalibrationen, Werte stabil und Driften gegen 0, QC Messung L3 => pH und E`lyte satt außerhalb Bereich, weitere Kal 2 und Folgemessung => keine Änderung. Referenz erneut membraniert, Kal und QC Messung => Alle Werte nun fast mittig. | Fehlersuche: Fließwege überprüft, Nadel gerichtet (s. a. SWO), Ca sowie Referenz membraniert, pH Dichtung ersetzt (s.a. SWO), mechanische Blockade im K Messplatz beseitigt (s.a. SWO), alle Messkammern gereinigt. Nach Neustart und Kal sowie tHb-Kal Lac Empf. unterhalb Grenze und pCO2 an Grenze => beide membraniert. Mehrere Kalibrationen, Werte stabil und Driften gegen 0, QC Messung L3 => pH und E`lyte satt außerhalb Bereich, weitere Kal 2 und Folgemessung => keine Änderung. Referenz erneut membraniert, Kal und QC Messung => Alle Werte nun fast mittig. | SWO-169954 | No | Troubleshooting: Flow paths checked, needle directed (see a. SWO), Ca and reference membraneed, pH seal replaced (see a. SWO), mechanical blockage in the K measuring station eliminated (see s.a. SWO), all measuring chambers cleaned. After restart and Kal as well as tHb-Kal Lac Empf. below limit and pCO2 at limit => both membraneed. Several calibrations, values stable and drift against 0, QC measurement L3 => pH and e'lyte enough outside range, further scale 2 and follow-up measurement => no change. Reference re-logged, Kal and QC measurement => All values now almost central. | Troubleshooting: Flow paths checked, needle directed (see a. SWO), Ca and reference membraneed, pH seal replaced (see a. SWO), mechanical blockage in the K measuring station eliminated (see s.a. SWO), all measuring chambers cleaned. After restart and Kal as well as tHb-Kal Lac Empf. below limit and pCO2 at limit => both membraneed. Several calibrations, values stable and drift against 0, QC measurement L3 => pH and e'lyte enough outside range, further scale 2 and follow-up measurement => no change. Reference re-logged, Kal and QC measurement => All values now almost central. | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX | Start der Kalibrier- Reinigungsroutine um 06:00 Uhr, erste Kal gegen 06:15 Uhr immer mit Na Driftgrenzverletzung. Datenextrakt gezogen zwecks Auswertung durch RDEU RMED | Start der Kalibrier- Reinigungsroutine um 06:00 Uhr; erste Kal gegen 06:15 Uhr immer mit Na Driftgrenzverletzung. Datenextrakt gezogen zwecks Auswertung durch RDEU RMED | SWO-170056 | No | Start of the calibration cleaning routine at 06:00; first Kal around 06:15 always with Na drift limit injury. Data extract extract drawn for evaluation by RDEU RMED | Start of the calibration cleaning routine at 06:00; first Kal around 06:15 always with Na drift limit injury. Data extract extract drawn for evaluation by RDEU RMED | Unknown | Unknown | Unknown | Unknown
ABL825 FLEX | checked data logs for errors, checked hemolyzer and lamp unit, performed complete oxi test | removed Oxi module and replaced lamp unit, adjusted lamp unit and performed complete oxi test, ran tHb cal | D2234ED9AC0DAB5D | No | checked data logs for errors, checked hemolyzer and lamp unit, performed complete oxi test | removed Oxi module and replaced lamp unit, adjusted lamp unit and performed complete oxi test, ran tHb cal | Defect | HW | Oxi module | Lamp unit
ABL827 FLEX | Ventile geprüft und gereinigt --> Neustart--> Kal i.O | Ventile geprüft und gereinigt --> Neustart--> Kal i.O | SWO-169877 | No | Valves tested and cleaned --> Restart--> Kal i.O | Valves tested and cleaned --> Restart--> Kal i.O | Cleaned | HW | El/Met Module | Valve, air
ABL835 FLEX | Found that only scheduled Cal 2's questioning. Cal 1's ok. Suspected Gas 2 valve problem. | Replaced Gas 2 valve and air valve on PH BG Module. Monitored performance over the next 24 hours. Problems went away. | D222C686AC22BEFB | No | Found that only scheduled Cal 2's questioning. Cal 1's ok. Suspected Gas 2 valve problem. | Replaced Gas 2 valve and air valve on PH BG Module. Monitored performance over the next 24 hours. Problems went away. | Defect | HW | Basic Unit | Gas module
ABL90 FLEX | Lösungspack zur Überprüfung abgezogen, Reklamationen erstellt. Mit den Anwendern über Lagerung und Frostindikator gesprochen. Lösungspack in der GS Werkstatt am ABL90+ geprüft, Fehler reproduzierbar. Die QC Lvl.C hat tHb und Met. bei Null, pH, pCO2 und E-Lyte. entsprechen den Kal.2 Lsg. Konzentrationen, Kal.2 Lsg. macht Empf. Fehler. Kal.2 Lsg. QC Lvl.C Lsg. müssen vertauscht sein. | Lösungspack zur Überprüfung abgezogen, Reklamationen erstellt. Mit den Anwendern über Lagerung und Frostindikator gesprochen. Lösungspack in der GS Werkstatt am ABL90+ geprüft, Fehler reproduzierbar. Die QC Lvl.C hat tHb und Met. bei Null, pH, pCO2 und E-Lyte. entsprechen den Kal.2 Lsg. Konzentrationen, Kal.2 Lsg. macht Empf. Fehler. Kal.2 Lsg. QC Lvl.C Lsg. müssen vertauscht sein. | SWO-170195 | No | Solution pack subtracted for review, complaints created. Talked to users about storage and frost indicator. Solution pack tested in the GS workshop on the ABL90+, errors reproducible. The QC Lvl.C has tHb and Met. at zero, pH, pCO2 and E-Lyte. correspond to the Kal.2 Lsg. concentrations, Kal.2 Lsg. makes Empf. Error. Kal.2 Lsg. QC Lvl.C Lsg. must be reversed. | Solution pack subtracted for review, complaints created. Talked to users about storage and frost indicator. Solution pack tested in the GS workshop on the ABL90+, errors reproducible. The QC Lvl.C has tHb and Met. at zero, pH, pCO2 and E-Lyte. correspond to the Kal.2 Lsg. concentrations, Kal.2 Lsg. makes Empf. Error. Kal.2 Lsg. QC Lvl.C Lsg. must be reversed. | Assisted Replacement | SP | SP | Pouch Mixup | XV-3
ABL90 FLEX PLUS |  | B | D2241DBA2214A13D | Yes |  |  | Assisted Replacement | SC | SC | Reference
ABL90 FLEX PLUS | インレットの位置調整 | B | SWO-170764 | No | Positioning inlets |  | Adjustment | HW | Auto Inlet module | Inlet Position Cal.
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-170121 | No | The replacement test card is OK. | The replacement test card is OK.
TCM5 FLEX | Monitor removed for further investigation. | Monitor removed for further investigation. | SWO-170625 | No | Monitor removed for further investigation. | Monitor removed for further investigation.
AQT90 FLEX | 更换636-539一个正常。 by andrew | 更换636-539一个正常。 by andrew | SWO-169773 | No | Replace 636-539 one normal. by andrew | Replace 636-539 one normal. by andrew | Defect | HW | Incubator module | Sensor, Cup | 1997
AQT90 FLEX | 更换905-908一个正常。by andrew | 更换905-908一个正常。by andrew | SWO-170346 | No | Replace 905-908 one normal. by andrew | Replace 905-908 one normal. by andrew | Defect | HW | Process Module | Sensor, BNP | 1235
AQT90 FLEX |  | B | SWO-170117 | No |  |  | Defect | SW | Analyzer Control Module | Off/On
ABL825 FLEX | Found unit in hold status with inlet calibration error. tHB calibration failing with a 314 ( Oxi wavelength deviation too high) error. Found lamp intensity too high. | Exchanged inlet gasket for fault isolation. Readjusted lamp intensity. Ran tHB calibration with pass results. Added power down shutdown message to both new units. Assisted customer with QC setup and scheduling on all analyzers. Restart calibrations. Cl failing level 3 QC high out of range. Replaced membrane. Restart calibrations. | D224541522237EF1 | No | Found unit in hold status with inlet calibration error. tHB calibration failing with a 314 ( Oxi wavelength deviation too high) error. Found lamp intensity too high. | Exchanged inlet gasket for fault isolation. Readjusted lamp intensity. Ran tHB calibration with pass results. Added power down shutdown message to both new units. Assisted customer with QC setup and scheduling on all analyzers. Restart calibrations. Cl failing level 3 QC high out of range. Replaced membrane. Restart calibrations. | Defect | HW | Inlet module | Inlet Gasket
ABL815 FLEX | Pumpenschläuche aus Kundenbestand erneuert. | Pumpenschläuche aus Kundenbestand erneuert. | SWO-170374 | No | Pump hoses from customer base. | Pump hoses from customer base. | Defect | HW | Fluidic manifold | Pump tubes | Why double
ABL815 FLEX | Pumpenschlauch erneuert. | Pumpenschlauch erneuert. | SWO-170375 | No | Pump hose renewed. | Pump hose renewed. | Defect | HW | Fluidic manifold | Pump tubes | Why double
ABL835 FLEX | Found the calibration ok on analyzer. QC levels out. Ran QC and all levels in and OK. Ran a precision run for tHb calibration in service and found precise for both new analyzers. Ran blood sample x7 on both analzyers for a comparision of MetHb and all values commuicated to specialist Keith Clark and all looks good. | No fault found on analylzer. | D22423E6221D60A9 | Yes | Found the calibration ok on analyzer. QC levels out. Ran QC and all levels in and OK. Ran a precision run for tHb calibration in service and found precise for both new analyzers. Ran blood sample x7 on both analzyers for a comparision of MetHb and all values commuicated to specialist Keith Clark and all looks good. | No fault found on analylzer. | Re classified RMED | courtesy call
AQT90 FLEX |  | B | D22438262246CCC2 | No |  |  | Defect | SW | Analyzer Control Module | Off/On
AQT90 FLEX |  | B | SWO-170280 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX | Corrected bends in inlet probe to make it perfectly straight. | Corrected bends in inlet probe to make it perfectly straight. | SWO-170626 | No | Corrected bends in inlet probe to make it perfectly straight. | Corrected bends in inlet probe to make it perfectly straight. | Defect | HW | Inlet module | Probe bend
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-170357 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | pO2
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-170813 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Reference | Leaking
ABL835 FLEX | Tras envío de Electrodo de Lactato el pasado 22 Feb y tras chequear equipo y ver que se ha solucioinado, se vuelve a producir el mismo problema. Chequeo de electrodo, sustitución de membrana, limpieza a fondo de preamplficador, Limpieza general de equipo y resto de preamplificadores | Sustitución electrodo de lactato, limpieza de preamplificadores, sustitución de membranas lact. | D225AD80983E4B69 | No | After sending Lactato Electrode on 22 Feb and after checking equipment and seeing that it has been solucioined, the same problem occurs again. Electrode check, membrane replacement, in-depth cleaning of the preamprifer, General cleaning of equipment and other preamps | Lactate electrode replacement, preamp cleaning, lact membrane replacement. | Assisted Replacement | Electrode | Membrane | Lac
ABL835 FLEX | Tras recibir electroodo de lactato sustituyo éste, remembranado, limpieza de camara de medida, limpieza de preamplificador con restos de oxido, al estar una de las lenguetas suelta se pega y fija, limpieza de preamplificadores, limpieza de camara de medida con bastantes restos secos de liquido en pocillo ref, preamplificadoor de Ref con muelle roto, limpieza de nipples de entrada, Ajuste de lámpara halógena tras fallo de oxi, Impresora rota pegado de pieza negra de plastico partida. instalación de Licencia Net-op sin éxito | Sustitución el lactato, limpieza preamplificador y ajuste, limpieza camaras de medida, nipples de entrada, ajuste de lámpra de módulo oxi. | D225AFDB98394B6D | No | After receiving electroode of lactate I replace this, remembrane, cleaning chamber of measurement, cleaning of preamp with oxide residues, being one of the loose tab sticks and fixed, cleaning of preamps, cleaning of measuring camera with quite a few Dry remnants of liquid in poe ref, Ref preamplificadoor with broken spring, cleaning of inlet nipples, Adjustment of halogen lamp after failure of oxi, Broken printer glued black part part part part part part part split. unsuccessful Net-op License installation | Replacing lactate, preamp cleaning and adjustment, cleaning measuring cameras, input nipples, oxy module lampra adjustment. | Assisted Replacement | Electrode | Membrane | Lac
ABL827 FLEX |  | B | D22539089835EB5E | Yes |  |  | Re classified RMED | Phone support
ABL800 BASIC | ｲﾝﾚｯﾄﾌﾟﾛｰﾌﾞ交換 | B | SWO-170441 | No | Inlet probe replacement |  | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Beim Eintreffen einwandfreie Funktion aller drei Sampler Pos. LEDs. LED manuell bewegt, Schlitten bewegt, keinerlei Auffälligkeiten oder Aussetzer mehr feststellbar => neue PCB vorerst NICHT eingebaut. | Beim Eintreffen einwandfreie Funktion aller drei Sampler Pos. LEDs. LED manuell bewegt, Schlitten bewegt, keinerlei Auffälligkeiten oder Aussetzer mehr feststellbar => neue PCB vorerst NICHT eingebaut. | SWO-170505 | No | On arrival all three sampler pos. Leds. LED manually moved, slide moved, no more abnormalities or dropouts detectable => new PCB NOT installed for the time being. | On arrival all three sampler pos. Leds. LED manually moved, slide moved, no more abnormalities or dropouts detectable => new PCB NOT installed for the time being. | Defect | HW | FlexQ Module | PCB, FlexQ | via short
AQT90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-171704 | No | The replacement test card is OK. | The replacement test card is OK. | Defect | Consumables | Test Kit | Cartridge | Broken, wings
AQT90 FLEX |  | B | SWO-170665 | Yes |  |  | Defect | HW | Process Module | Optical Unit | 129
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-171698 | Yes | Replace the test card | Replace the test card | Re classified RMED | OOB
AQT90 FLEX |  | B | SWO-170704 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL800 BASIC | replace the pressure transducer with unit chasis,check the leakage test for gas 2 and found the flow regulator valve for gas2 was not working,need to replace the the valve. | B | SWO-171314 | No | replace the pressure transducer with unit chasis,check the leakage test for gas 2 and found the flow regulator valve for gas2 was not working,need to replace the the valve. |  | Defect | HW | Gas module | Gas module
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-170824 | Yes | The replacement test card is OK. | The replacement test card is OK.
ABL90 |  | B | D22A9A7E705F69D2 | No |  |  | Assisted Replacement | SC | SC | Glu/Lac
ABL90 |  | B | D2292EF8FA61DCD1 | Yes
ABL90 FLEX PLUS |  | B | D2292661FA416792 | Yes |  |  | Re classified RMED | PER
ABL90 FLEX PLUS |  | B | D22918C7FA25678F | Yes |  |  | Re classified RMED | PER
ABL90 FLEX PLUS | Allgemeine Kontrolle ok. Sensorkassette und Lösungspack entfernt. System durchgespült ok. System mit Sensorkassette durchgepült ok. Alle Tests im Servicememü durchgeführt ok. Neustart ok. Warten 55 min für Initialisierung | Allgemeine Kontrolle ok. Sensorkassette und Lösungspack entfernt. System durchgespült ok. System mit Sensorkassette durchgepült ok. Alle Tests im Servicememü durchgeführt ok. Neustart ok. Warten 55 min für Initialisierung | SWO-170957 | No | General control ok. Sensor cartridge and solution pack. System flushed ok. System with sensor cassette put through ok. All tests performed in the service memü ok. Restart ok. Wait 55 min for initialization | General control ok. Sensor cartridge and solution pack. System flushed ok. System with sensor cassette put through ok. All tests performed in the service memü ok. Restart ok. Wait 55 min for initialization | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | Lösungspack erneuert und installiert. | Lösungspack erneuert und installiert. | SWO-171092 | Yes | Solution pack renewed and installed. | Solution pack renewed and installed. | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Fehlenden Lösungspack geliefert und installiert. SC aus Kundenbestand erneuert. | Fehlenden Lösungspack geliefert und installiert. SC aus Kundenbestand erneuert. | SWO-171093 | Yes | Missing solution pack delivered and installed. SC from customer base. | Missing solution pack delivered and installed. SC from customer base. | Re classified RMED | Logistical upSWO
AQT90 FLEX | Analise do registo de dados e mensagens do sistema. Calibração do Cup Dryer , impossivel calibrar. Substituição da cabeça das bombas, novo tentativa de calibração do Cup Dryer sem sucesso, Abertura do equipamento, limpeza e verificação da placa da unidade de secagem dos copos, limpeza da unidade de lavagem. inicialização. Realização de 2 verificações de sistema com resultados OK. Realização de dois controlos de quailidade OK. | B | D22951A1FA36E77A | No | Analyze data logging and system messages. Calibration of Cup Dryer, impossible to calibrate. Replacement of the pump head, new attempt to calibrate the Cup Dryer without success, opening the equipment, cleaning and checking the plate of the drying unit of the cups, cleaning the washing unit. Startup. Realization of 2 system checks with results OK. Realization of two FINE quaility checks. |  | Defect | HW | Process Module | Compressor | 168
TCM5 BASIC |  | B | D2291052FA32B1CF | No
ABL800 FLEX | Check the Inlet gasket. Check the tubing .Check the upper liquid sensor & lower inlet section. Check the rinse & air valve both valve faulty. Replace the air & rinse valve. calibrate system multiple time. Calibration & qc pass. | Check the Inlet gasket. Check the tubing .Check the upper liquid sensor & lower inlet section. Check the rinse & air valve both valve faulty. Replace the air & rinse valve. calibrate system multiple time. Calibration & qc pass. | SWO-172375 | No | Check the Inlet gasket. Check the tubing . Check the upper liquid sensor & lower inlet section. Check the rinse & air valve both valve faulty. Replace the air & rinse valve. calibrate system multiple time. Calibration & qc pass. | Check the Inlet gasket. Check the tubing . Check the upper liquid sensor & lower inlet section. Check the rinse & air valve both valve faulty. Replace the air & rinse valve. calibrate system multiple time. Calibration & qc pass. | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX |  | B | D22912C9FA51FB94 | No |  |  | Defect | HW | Inlet module | Inlet module
ABL837 FLEX | ＣＰＵボード不良の為交換。 | ＣＰＵボード不良の為交換。 | SWO-171697 | No | Replacement due to cpu board failure. | Replacement due to cpu board failure. | Defect | HW | DMS_PC | CPU
ABL820 FLEX | Inlet area leaking on cals, QC and blood gases. David worked on analyzer Friday. | Flushed from inlet to Oxi area. Replaced tubing block. Replace inlet gasket and probe. | D22965B4FA36913D | No | Inlet area leaking on cals, QC and blood gases. David worked on analyzer Friday. | Flushed from inlet to Oxi area. Replaced tubing block. Replace inlet gasket and probe. | Defect | HW | Inlet module | Inlet Gasket
ABL827 FLEX | Checked data logs for errors. Zero current errors on pO2. | Perform zero current calibration in service. Restart calibrations. | D2292B10FA090F14 | No | Checked data logs for errors. Zero current errors on pO2. | Perform zero current calibration in service. Restart calibrations. | Adjustment | HW | pH/BG Module | Zero current
ABL825 FLEX | Checked data logs for errors. Problem is intermittent occurring on QC most of the time. . Checked fluid path. | Replaced various tubing on pH BG module and seal ring on pH electrode. Flushed inlet probe. Restart calibrations. | D2294587FA600F19 | No | Checked data logs for errors. Problem is intermittent occurring on QC most of the time. . Checked fluid path. | Replaced various tubing on pH BG module and seal ring on pH electrode. Flushed inlet probe. Restart calibrations. | Defect | Electrode | Electrode | pH | Sealing ring
AQT90 FLEX | Analysator geöffnet - 3 Probenehmer entfernt - Funktionsüberprüfung etc. ohne Beanstandung | Analysator geöffnet - 3 Probenehmer entfernt - Funktionsüberprüfung etc. ohne Beanstandung | SWO-170931 | No | Analyzer open - 3 samplers removed - function check etc. without complaint | Analyzer open - 3 samplers removed - function check etc. without complaint | Defect | HW | Aspiration Module | Tube holder
ABL80 FLEX Basic version | Recargado software. Comprobado buen funcionamiento del equipo y pasadas muestras | B | D22A9F3C7012091D | No | Reloaded software. Proven smooth equipment performance and past samples
ABL80 FLEX BASIC | Waited observed the few pointes, it was temperature issue or power failure, explained the same to Biomedical eng, suggested to provide the UPS point and room temperature need to maintain 18 to 22 C, it will be under observation next weel. | Waited observed the few pointes, it was temperature issue or power failure, explained the same to Biomedical eng, suggested to provide the UPS point and room temperature need to maintain 18 to 22 C, it will be under observation next weel. | SWO-172389 | No | Waited observed the few pointes, it was temperature issue or power failure, explained the same to Biomedical eng, suggested to provide the UPS point and room temperature need to maintain 18 to 22 C, it will be under observation next weel. | Waited observed the few pointes, it was temperature issue or power failure, explained the same to Biomedical eng, suggested to provide the UPS point and room temperature need to maintain 18 to 22 C, it will be under observation next weel.
ABL80 | Checked the data logs and verified the problem. | Ordered all new boards and senxsor cassette. I will return on march 1st to finish the repair. | D22A67CC70143CB1 | No | Checked the data logs and verified the problem. | Ordered all new boards and senxsor cassette. I will return on march 1st to finish the repair.
ABL90 FLEX | SC csere után a kalibráció stabilizálódik | B | SWO-171212 | No | After SC replacement, the calibration stabilates |  | Assisted Replacement | SC | SC | Reference
ABL90 FLEX | Gerät überprüft, Fließwege durchgespült, Lösungspack ersetzt, Sensor erneuert, Kal und QC o.k. Eingabe Fixe Zielwerte korregiert | Gerät überprüft, Level A und C grenzwertig, Lösungspack ersetzt | SWO-171401 | No | Device checked, flow paths flushed through, solution pack replaced, sensor renewed, Kal and QC ok. Input Fixed target values corrected | Device checked, level A and C borderline, solution pack replaced | Assisted Replacement | SC | SC | pCO2 & Ca
ABL90 | Performed voltage checks, no issue with voltages, suspect CPU failed. | Replaced CPU, swapped CF card from original CPU, no loss of data. Due to extended period of non-operation sensor cassette no longer meets specification, pCO2 error 378, Lac error 375, QC errors 210, replaced SC, replacement SC failed due to error 378 for pCO2 Lac, error 210 for pCO2 Lac, replaced SC again, printed out credit forms (2) for customer to submit for replacement. Performed start up and standard tests, performed tHB calibration as a courtesy on both ABL90s, overdue, no further issues observed. | D22AD7D17029F57A | No | Performed voltage checks, no issue with voltages, suspect CPU failed. | Replaced CPU, swapped CF card from original CPU, no loss of data. Due to extended period of non-operation sensor cassette no longer meets specification, pCO2 error 378, Lac error 375, QC errors 210, replaced SC, replacement SC failed due to error 378 for pCO2 Lac, error 210 for pCO2 Lac, replaced SC again, printed out credit forms (2) for customer to submit for replacement. Performed start up and standard tests, performed tHB calibration as a courtesy on both ABL90s, overdue, no further issues observed. | Defect | HW | DMS_PC | CPU
ABL90 FLEX | tHb Kal durchgeführt, alle drei QCs durchgeführt | tHb Kal durchgeführt, alle drei QCs durchgeführt | SWO-171387 | No | tHb Kal, all three QCs performed | tHb Kal, all three QCs performed | Adjustment | HW | Oxi module | tHb Cal
ABL90 FLEX PLUS |  | B | D22AA66C70130B04 | Yes |  |  | Re classified RMED | FSE error
ABL90 FLEX PLUS | Fehlersuche, div. Schläuche erneuert, Einlaßdichtung neu Einlaßnadel neu SC erneuert, da nach wie vorFehlermeldung und Fließweg nicht frei Überprüfung Fluid Transport im ServiceProgramm Hemolyser verstopft, läßt sich nicht freispülen Einlaufverfahren fehlerhaft | Hemolyser verstopft Hemolyser Unit erneuert tHb Kal Einlaufverfahren | SWO-171281 | No | Troubleshooting div. Hoses renewed, inlet seal new Inlet needle new SC is renewed, as the error message and flow path are still not free Review Fluid Transport in the ServiceProgram Hemolyser clogged, cannot be rinsed free Inlet method faulty | Hemolyser clogged Hemolyser Unit renewed tHb Kal Entry procedure | Revisit required | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-171090 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Reference | 1361
ABL825 FLEX | 更换834-222修复 by jesson | 更换834-222修复 by jesson | SWO-171076 | No | Replace 834-222 Repair by jesson | Replace 834-222 Repair by jesson | Defect | HW | pH/BG Module | Valve, air | Membrane
ABL800 BASIC | Replaced The Chloride Membrane, Machine Working Ok | Replaced The Chloride Membrane, Machine Working Ok | SWO-172357 | No | Replaced The Chloride Membrane, Machine Working Ok | Replaced The Chloride Membrane, Machine Working Ok | Assisted Replacement | Electrode | Membrane | Cl | Sensitivity
ABL825 FLEX | Analysator heruntergefahren - Überprüfung etc. Messventile - Erneuerung Luftmembran Ventil pH BG Modul - Einstellung Hallogenlampe Oximetrie -Funktionsüberprüfung etc. | Analysator heruntergefahren - Überprüfung etc. Messventile - Erneuerung Luftmembran Ventil pH BG Modul - Einstellung Hallogenlampe Oximetrie -Funktionsüberprüfung etc. | SWO-171241 | No | Analyzer shut down - check etc. Measuring valves - Renewal air membrane valve pH BG module - setting hallogen lamp oximetry -function check etc. | Analyzer shut down - check etc. Measuring valves - Renewal air membrane valve pH BG module - setting hallogen lamp oximetry -function check etc. | Defect | HW | pH/BG Module | Valve, air
ABL815 FLEX | PumpenKal im Service, probehalber Modul- und Lösungspumpenschläuche ersetzt, Einlassdichtung und Einlaßsonde ersetzt, PumpenKal, Pumpenschläuche zurück getauscht, PumpenKal ok, QK ok | PumpenKal im Service, probehalber Modul- und Lösungspumpenschläuche ersetzt, Einlassdichtung und Einlaßsonde ersetzt, PumpenKal, Pumpenschläuche zurück getauscht, PumpenKal ok, QK ok | SWO-171559 | No | PumpsKal in service, test-tested module and solution pump hoses replaced, inlet seal and inlet probe replaced, pumpsKal, pump hoses replaced, pumpsKal ok, QK ok | PumpsKal in service, test-tested module and solution pump hoses replaced, inlet seal and inlet probe replaced, pumpsKal, pump hoses replaced, pumpsKal ok, QK ok | Defect | HW | Fluidic manifold | Pump tubes
ABL835 FLEX |  | B | D22A56E1702A4CD0 | No |  |  | Defect | HW | pH/BG Module | Measuring chamber
AQT90 FLEX |  | B | D22AAF7B70448858 | Yes |  |  | Re classified RMED | HW | Aspiration Module | Needle assembly | This is revisit w/o complaint
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-171699 | No | Replace the test card | Replace the test card
ABL80 | Checked data logs and verified problem. | Replaced the signal data board, sensor board due to corroson, and the sensor cassette due to a leak. | D22B913FE64B2DA4 | No | Checked data logs and verified problem. | Replaced the signal data board, sensor board due to corroson, and the sensor cassette due to a leak.
ABL9 | SD an Kopenhagen weitergeleitet | SD an Kopenhagen weitergeleitet | SWO-171756 | No | SD forwarded to Copenhagen | SD forwarded to Copenhagen
ABL90 FLEX | Gerät überprüft, Level A und C grenzwertig, Lösungspack ersetzt | Gerät überprüft, Level A und C grenzwertig, Lösungspack ersetzt | SWO-171521 | No | Device checked, level A and C borderline, solution pack replaced | Device checked, level A and C borderline, solution pack replaced | Assisted Replacement | SP | SP | pCO2 | RDEU GEMBA
ABL90 | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | Connected analyzer to hospitals time servier and synced up. | D22BF0B1E6563B8D | Yes | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | Connected analyzer to hospitals time servier and synced up.
ABL90 | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | Connected analyzer to hospitals time servier and synced up. | D22BF43EE6243B90 | Yes | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | Connected analyzer to hospitals time servier and synced up.
ABL90 | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | Connected analyzer to hospitals time servier and synced up. | D22BF4FEE6153B92 | Yes | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | Connected analyzer to hospitals time servier and synced up.
ABL90 | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | Connected analyzer to hospitals time servier and synced up. | D22BF5B8E6043B95 | Yes | Upon arrival, verified customers stated complaint and discovered analyzer time was a minute faster and failing to sync with LIS. | Connected analyzer to hospitals time servier and synced up.
ABL90 | Examined data logs. Examined oxi module. | Replaced sensor cassette for troubleshooting purposes. Flushed Oxi module. Performed complete Oxi test. Performed tHb calibration. Verified operation of ABL90. | D22B9E3EE65330AF | Yes | Examined data logs. Examined Oxi module. | Replaced sensor cassette for troubleshooting purposes. Flushed Oxi module. Performed complete Oxi test. Performed tHb calibration. Verified operation of ABL90.
ABL90 | Jay called to advise they are having longstanding issues with Glu Lac constantly falling out of calibration and failing QC | OSPD found intermittent sensitivity falloffs on calibration results. Sensor test shows pass, however issues is very intermittent. Replaced Sensor Assembly PCB, test passed, multiple calibrations passed. | D22BCC9BE604CDC5 | No | Jay called to advise they are having longstanding issues with Glu Lac constantly falling out of calibration and failing QC | OSPD found intermittent sensitivity falloffs on calibration results. Sensor test shows pass, however issues is very intermittent. Replaced Sensor Assembly PCB, test passed, multiple calibrations passed. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX PLUS |  | B | D22D14815C613E66 | Yes
ABL90 FLEX PLUS |  | B | D22B7B14E654DA08 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | Tried to locate solution pack from out laying hospitals but nobody had to borrow. Used the non working analyzers solution pack to get the machine functional. | Tried to locate solution pack from out laying hospitals but nobody had to borrow. Used the non working analyzers solution pack to get the machine functional. | SWO-171452 | Yes | Tried to locate solution pack from out laying hospitals but nobody had to borrow. Used the non working analyzers solution pack to get the machine functional. | Tried to locate solution pack from out laying hospitals but nobody had to borrow. Used the non working analyzers solution pack to get the machine functional. | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Hemolyser verstopft Hemolyser Unit erneuert tHb Kal Einlaufverfahren | Hemolyser verstopft Hemolyser Unit erneuert tHb Kal Einlaufverfahren | SWO-171470 | No | Hemolyser clogged Hemolyser Unit renewed tHb Kal Entry procedure | Hemolyser clogged Hemolyser Unit renewed tHb Kal Entry procedure | Defect | HW | Oxi module | Hemolyzer | Clot
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-177949 | Yes | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Glu/Lac
TCM5 FLEX | センサーシールパーツ交換 | センサーシールパーツ交換 | SWO-171344 | No | Sensor seal parts replacement | Sensor seal parts replacement
AQT90 FLEX | 更换903-364修复 by jesson | 更换903-364修复 by jesson | SWO-171137 | No | Replace 903-364 Repair by jesson | Replace 903-364 Repair by jesson | Defect | HW | Process Module | Optical Unit | 124
ABL825 FLEX |  | B | D22BD5B9E62CF611 | Yes |  |  | Re classified RMED | DEMO Analyzer
ABL835 FLEX | Geräteneustart durchgeführt Touchscreen kalibriert Gerät kalibriert und überprüft | Geräteneustart durchgeführt Touchscreen kalibriert Gerät kalibriert und überprüft | SWO-171568 | No | Device restart performed Touchscreen calibrated Device calibrated and checked | Device restart performed Touchscreen calibrated Device calibrated and checked | Defect | HW | DMS_PC | Touch screen
ABL815 FLEX | Computereinheit demontiert, Druckkopf ausgetauscht, Auto Druckfunktion deaktiviert, Funktionskontrolle ok, Med Technik H. Wahrlich über die Berechnung informiert | B | SWO-171590 | No | Computer unit dismantled, print head replaced, car print function deactivated, function control ok, Med Technique H. Truly informed about the calculation |  | Defect | HW | Printer Module | Printer head
ABL815 FLEX | Der am 26.01.2017 im Rahmen der FAN915-318 eingesetzte Dichtring 834-682 834-618 ist bereits UNDICHT und es tropft aus dem Überlauf der Messkammer, beim Entnehmen der Elektrode verblieb der Dichtring in der Messkammer. Wurde von Frau Klubert neu positioniert. | Der am 26.01.2017 im Rahmen der FAN915-318 eingesetzte Dichtring 834-682 834-618 ist bereits UNDICHT und es tropft aus dem Überlauf der Messkammer; beim Entnehmen der Elektrode verblieb der Dichtring in der Messkammer. Wurde von Frau Klubert neu positioniert. | SWO-171552 | No | The sealing ring 834-682 834-618 used on 26.01.2017 as part of the FAN915-318 is already UNDICHT and it drips out of the overflow of the measuring chamber; when the electrode was removed, the sealing ring remained in the measuring chamber. Was repositioned by Ms. Klubert. | The sealing ring 834-682 834-618 used on 26.01.2017 as part of the FAN915-318 is already UNDICHT and it drips out of the overflow of the measuring chamber; when the electrode was removed, the sealing ring remained in the measuring chamber. Was repositioned by Ms. Klubert. | User error | Electrode | Electrode | pH | Sealing ring
ABL800 BASIC | required new printer head. | required new printer head. | SWO-171520 | No | required new printer head. | required new printer head. | Defect | HW | Printer Module | Printer head
ABL825 FLEX |  | B | D22CD33E5C563D7B | Yes |  |  | Re classified RMED | FAN
ABL820 FLEX | Bei eintreffen war der Analysator messbereit. Frau Neiswirth hat mir mitgeteilt dass pH bei der Messung sporadisch instabil ist. PH Elektrode gereinigt und unteren pH Dichtungsring getauscht. Mehrere Kalibrierungen durchgeführt. QCs für alle Messungen durchgeführt. ABL wieder messbereit. Mehrere Patientenmessungen durchgeführt. Keine instabile Messung mehr feststellbar. | Bei eintreffen war der Analysator messbereit. Frau Neiswirth hat mir mitgeteilt dass pH bei der Messung sporadisch instabil ist. PH Elektrode gereinigt und unteren pH Dichtungsring getauscht. Mehrere Kalibrierungen durchgeführt. QCs für alle Messungen durchgeführt. ABL wieder messbereit. Mehrere Patientenmessungen durchgeführt. Keine instabile Messung mehr feststellbar. | SWO-171658 | No | When the analyzer arrived, it was ready to measure. Mrs Neiswirth told me that pH is sporadically unstable in the measurement. PH electrode cleaned and lower pH sealing ring replaced. Multiple calibrations performed. QCs for all measurements. ABL ready to be measured again. Several patient measurements have been carried out. No more unstable measurement can be detected. | When the analyzer arrived, it was ready to measure. Mrs Neiswirth told me that pH is sporadically unstable in the measurement. PH electrode cleaned and lower pH sealing ring replaced. Multiple calibrations performed. QCs for all measurements. ABL ready to be measured again. Several patient measurements have been carried out. No more unstable measurement can be detected. | No fault found | Electrode | Electrode | pH | Sealing ring
ABL825 FLEX | Checked and cleaned all measuring chamber wells in both modules. Found leakage in Na well. Checked all tubing between measuring chambers. | Replaced Na and Ca membranes. Restart calibrations. Perform zero current calibration in diagnostics. pH still unstable. Replaced pH electrode. Restart testing with blood samples. | D22BDFCAE6012BD9 | No | Checked and cleaned all measuring chamber wells in both modules. Found leakage in Na well. Checked all tubing between measuring chambers. | Replaced Na and Ca membranes. Restart calibrations. Perform zero current calibration in diagnostics. pH still unstable. Replaced pH electrode. Restart testing with blood samples. | Assisted Replacement | Electrode | Electrode | pH
ABL80 FLEX BASIC |  | B | SWO-171692 | No
ABL80 FLEX CO-OX | Gerät an einen den Standort ZNA AQT90Flex ( 838R0096N010) transportiert und dort angebunden. Es konnte keine Verbindung zum Radiance mit dem ABL80COOX aufbauen. Verbindungstest zwischen Notebook und ABL80COOX mit ASTM.Net durchgeführt. Verbindung war OK. Qualitätskontrollen und Patientenbefunde konnten vom Gerät zum Laptop übertragen werden. Gerät wurde gemäß Kundenangaben wieder konfiguriert (IP: 10.96.11.54, Subnet Mask: 255.255.255.0, GT: 10.96.0.200). Verbindung zum Radiance konnte weiterhin nicht aufgebaut werden. | Gerät an einen den Standort ZNA AQT90Flex ( 838R0096N010) transportiert und dort angebunden. Es konnte keine Verbindung zum Radiance mit dem ABL80COOX aufbauen. Verbindungstest zwischen Notebook und ABL80COOX mit ASTM.Net durchgeführt. Verbindung war OK. Qualitätskontrollen und Patientenbefunde konnten vom Gerät zum Laptop übertragen werden. Gerät wurde gemäß Kundenangaben wieder konfiguriert (IP: 10.96.11.54, Subnet Mask: 255.255.255.0, GT: 10.96.0.200). Verbindung zum Radiance konnte weiterhin nicht aufgebaut werden. | SWO-172293 | No | device to a location ZNA AQT90Flex ( 838R0096N010) and connected there. It was unable to connect to the Radiance with the ABL80COOX. Connection test between notebook and ABL80COOX with ASTM.Net performed. Connection was OK. Quality controls and patient findings could be transferred from the device to the laptop. Device has been reconfigured according to customer specifications (IP: 10.96.11.54, Subnet Mask: 255.255.255.0, GT: 10.96.0.200). The connection to the radiance could still not be established. | device to a location ZNA AQT90Flex ( 838R0096N010) and connected there. It was unable to connect to the Radiance with the ABL80COOX. Connection test between notebook and ABL80COOX with ASTM.Net performed. Connection was OK. Quality controls and patient findings could be transferred from the device to the laptop. Device has been reconfigured according to customer specifications (IP: 10.96.11.54, Subnet Mask: 255.255.255.0, GT: 10.96.0.200). The connection to the radiance could still not be established.
ABL80 FLEX | SC80 Installation Failure, leakage found from SC80, Sensor Board Get Defective By Leakage from SC,Placed Order for Connector Board and PER Sent for replacement | SC80 Installation Failure, leakage found from SC80, Sensor Board Get Defective By Leakage from SC,Placed Order for Connector Board and PER Sent for replacement | SWO-172371 | No | SC80 Installation Failure, leakage found from SC80, Sensor Board Get Defective By Leakage from SC,Placed Order for Connector Board and PER Sent for replacement | SC80 Installation Failure, leakage found from SC80, Sensor Board Get Defective By Leakage from SC,Placed Order for Connector Board and PER Sent for replacement
ABL80 CO-OX | Analyzer had physical damage to the display. | Replaced the Display Cover, Touch Screen, and LCD Display. | D22D24635C5EFE1B | No | Analyzer had physical damage to the display. | Replaced the Display Cover, Touch Screen, and LCD Display.
ABL80 FLEX Basic version |  | B | D22D295F5C0557A1 | No
ABL90 FLEX | 更换903-197修复 by jesson | 更换903-197修复 by jesson | SWO-171172 | No | Replace 903-197 Fix by Jesson | Replace 903-197 Fix by Jesson | Defect | HW | Oxi module | Spectrometer | 134
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-172027 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Glu/Lac | 1378
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-177966 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | FT | 1268
AQT90 FLEX |  | B | SWO-171700 | No |  |  | Defect | HW | Process Module | Process Unit sensor
AQT90 FLEX | 更换903-364修复 by jesson | 更换903-364修复 by jesson | SWO-171694 | No | Replace 903-364 Repair by jesson | Replace 903-364 Repair by jesson | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-171781 | No |  |  | Defect | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX |  | B | SWO-171769 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL825 FLEX | Einlassdichtung und Nadel überprüft i.o Gerinnsel in pH entfernt i.o O-Ring pH ersetzt i.o Messkammer pH BG-Modul ausgebaut, ersetzt da Risse im pO2 Zylinder und Messkammer(Nicht die Fehlerursach) i.o Ventile pH BG-Modul ausgebaut und gereinigt i.o Nachfüllung, LS justage und Pumpen Kal. i.o Drei Testmessungen mit Blut i.o Danach 2. Kal. i.o Testmessung mit FlexQ i.o Erneut 2. Kal. i.o Testmessung mit FlexQ i.o QC Level 3 i.o Danach erneut 2. Kal. Na instabil Membranenwartung Na durchgeführt i.o Kal. und QC i.o Testmessung i.o Messung Kunde i.o | B | SWO-171970 | No | Inlet seal and needle checked i.o Clots removed in pH i.o O-ring pH replaces i.o measuring chamber pH BG module removed, replaces cracks in the pO2 cylinder and measuring chamber (Not the error error) i.o Valves pH BG module removed and cleaned i.o Refilling, LS adjustment and pumps Kal. Ok Three test measurements with blood i.o After that 2nd Kal. Ok Test measurement with FlexQ i.o Again 2. Ok Test measurement with FlexQ i.o QC Level 3 i.o Then again 2nd Kal. Na unstable Membrane maintenance Na carried out i.o Kal. and QC i.o Test measurement i.o Measurement customer i.o |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL800 FLEX | Replace the Crea reagent & calibrate system. | Replace the Crea reagent & calibrate system. | SWO-172340 | No | Replace the Crea reagent & calibrate system. | Replace the Crea reagent & calibrate system. | Assisted Replacement | Consumables | Solution bottle | Crea Solution
ABL800 BASIC | 更换902-458修复 by jesson | 更换902-458修复 by jesson | SWO-171430 | No | Replace 902-458 Repair by Jesson | Replace 902-458 Repair by Jesson | Defect | HW | Oxi module | Lamp unit
ABL825 FLEX | PO2 Elektrode geprüft und ausgetauscht, Funktionskontrolle ok | B | SWO-172076 | No | PO2 electrode tested and replaced, function control ok |  | Assisted Replacement | Electrode | Electrode | pO2
ABL825 FLEX | Examined inlet module and QC logs. Found bent inlet probe. | Replaced inlet probe. Replaced inlet heater. Verified operation of ABL825. | D22CFD3D5C28C493 | No | Examined inlet module and QC logs. Found bent inlet probe. | Replaced inlet probe. Replaced inlet heater. Verified operation of ABL825. | Defect | HW | Inlet module | Probe bend
ABL835 FLEX | Neon lamp udskiftet, slutafprøvet. | Neon lamp udskiftet, slutafprøvet. | SWO-171963 | No | Neon lamp replaced, final tested. | Neon lamp replaced, final tested. | Defect | HW | Oxi module | Neon Lamp
ABL825 FLEX | Found error code 314 (Oxi wavelength deviation high) in QC log. Failed tHB calibration in service with status 23 (Wavelength error). Not failing complete oximetry tests. | Replaced spectrometer unit. Adjusted lamp intensity. Ran all tests with pass results. | D22CE4FA5C0CED7F | No | Found error code 314 (Oxi wavelength deviation high) in QC log. Failed tHB calibration in service with status 23 (Wavelength error). Not failing complete oximetry tests. | Replaced spectrometer unit. Adjusted lamp intensity. Ran all tests with pass results. | Defect | HW | Oxi module | Spectrometer
ABL800 BASIC | インレットモジュール入れ替え | B | SWO-171681 | No | Replacing inlet modules |  | Defect | HW | Inlet module | Inlet module
ABL835 FLEX | Arrived onsite and could not get oxy mearsurements with consistant resluts even though sample entering hemolyzer. Error 581. | Replaced the hemolyzer. Adjusted lamp to match other analyzer at 550Pa ro 3200mv. Grey tube next to pinch tubing had pin hole in it. Replaced tubing. Ran qc and error 0374 Inhmog sample for PH BG intermittant. Found pin hole in grey tubing between measuring chambers. Replaced tubing in the flow path of sample. Resolved errors and calibrated without errors. | D22D036C5C032815 | No | Arrived onsite and could not get oxy mearsurements with consistant resluts even though sample entering hemolyzer. Error 581. | Replaced the hemolyzer. Adjusted lamp to match other analyzer at 550Pa ro 3200mv. Grey tube next to pinch tubing had pin hole in it. Replaced tubing. Ran qc and error 0374 Inhmog sample for PH BG intermittant. Found pin hole in grey tubing between measuring chambers. Replaced tubing in the flow path of sample. Resolved errors and calibrated without errors. | Defect | HW | Oxi module | Hemolyzer
ABL835 FLEX | Arrived onsite and the PC02 and the P02 are not calibrating. | Found pin hole in grey tubing to the left of the pintch tubing. Replaced the tubing and all the tubing for the PH BG module. Analyzer calibrating without errors. | D22D14725C2D281C | No | Arrived onsite and the PC02 and the P02 are not calibrating. | Found pin hole in grey tubing to the left of the pintch tubing. Replaced the tubing and all the tubing for the PH BG module. Analyzer calibrating without errors. | Defect | HW | pH/BG Module | Tubing | grey
ABL820 FLEX | Verified time is set correctly in Windows and on the analyzer setup. Also pum calibration error. | Replaced PG pump tubing. Changed BIOS time and tested QC. All tests are successful. | D22D1A155C48E02C | No | Verified time is set correctly in Windows and on the analyzer setup. Also pum calibration error. | Replaced PG pump tubing. Changed BIOS time and tested QC. All tests are successful. | Setup | Time setting | Time setting | Time setting
AQT90 FLEX | Fixed the Problem, Perform Set safe passed Completely, Cleaned analyzer Internally and Externally, Problem Resolve. | Fixed the Problem, Perform Set safe passed Completely, Cleaned analyzer Internally and Externally, Problem Resolve. | SWO-172368 | No | Fixed the Problem, Perform Set safe passed Completely, Cleaned analyzer Internally and Externally, Problem Resolve. | Fixed the Problem, Perform Set safe passed Completely, Cleaned analyzer Internally and Externally, Problem Resolve. | Cleaned | HW | Storage Module | Storage wheel
AQT90 FLEX | Fehlersuche: Dateien überprüft, unter Aktivitäten diverse Meldungen über mechanische Probleme des Storagewheels sowie des Processwheels und weitere ähnliche Fehlermeldungen. Beim Öffnen des Analysators fiel ein Testcup unbekannter Herkunft aus dem Gerät, im Servicemenü safe positions angefahren, OK. Sämtliche in Frage kommende Komponenten sichtkontrolliert und per Serviceprogramm aktiviert, OK. Weder übermäßige Staubbelastung noch mechanische Probleme erkennbar. Im Flows Programm insgesamt 16 Testcups in Folge transportiert, alle einwandfrei absolviert. Montage, Neustart, Endkontrolle: LQC Messung OK | Fehlersuche: Dateien überprüft, unter Aktivitäten diverse Meldungen über mechanische Probleme des Storagewheels sowie des Processwheels und weitere ähnliche Fehlermeldungen. Beim Öffnen des Analysators fiel ein Testcup unbekannter Herkunft aus dem Gerät; im Servicemenü safe positions angefahren, OK. Sämtliche in Frage kommende Komponenten sichtkontrolliert und per Serviceprogramm aktiviert, OK. Weder übermäßige Staubbelastung noch mechanische Probleme erkennbar. Im Flows Programm insgesamt 16 Testcups in Folge transportiert, alle einwandfrei absolviert. Montage, Neustart, Endkontrolle: LQC Messung OK | SWO-172002 | No | Troubleshooting: Checks files, under Activities various messages about mechanical problems of the storage wheel as well as the process wheel and other similar error messages. When the analyzer was opened, a test cup of unknown origin fell from the device; in the safe positions service menu, OK. All eligible components are visually controlled and activated by a service program, OK. No excessive dust exposure or mechanical problems can be detected. Transported in the Flows program a total of 16 test cups in a row, all completed flawlessly. Assembly, restart, final control: LQC measurement OK | Troubleshooting: Checks files, under Activities various messages about mechanical problems of the storage wheel as well as the process wheel and other similar error messages. When the analyzer was opened, a test cup of unknown origin fell from the device; in the safe positions service menu, OK. All eligible components are visually controlled and activated by a service program, OK. No excessive dust exposure or mechanical problems can be detected. Transported in the Flows program a total of 16 test cups in a row, all completed flawlessly. Assembly, restart, final control: LQC measurement OK | Defect | SW | Analyzer Control Module | Off/On
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-172113 | Yes | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX | Cleaned machine Internally and Externally, Done Zeroing, Reinstalled SC80 and Replaced SP with New one, problem Resolve. | Cleaned machine Internally and Externally, Done Zeroing, Reinstalled SC80 and Replaced SP with New one, problem Resolve. | SWO-172372 | No | Cleaned machine Internally and Externally, Done Zeroing, Reinstalled SC80 and Replaced SP with New one, problem Resolve. | Cleaned machine Internally and Externally, Done Zeroing, Reinstalled SC80 and Replaced SP with New one, problem Resolve.
ABL90 |  | B | D22E5AD2D21183DE | No |  |  | Defect | HW | Printer Module | Printer Module
AQT90 FLEX |  | B | SWO-172051 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
AQT90 FLEX |  | B | SWO-172025 | No |  |  | Defect | HW | Analyzer Control Module | Barcode reader, display
ABL835 FLEX | Ref電極交換。 PHｼｰﾘﾝｸﾞﾘﾝｸﾞ張り付きあり。 | Ref電極交換。 PHｼｰﾘﾝｸﾞﾘﾝｸﾞ張り付きあり。 | SWO-172530 | No | Ref electrode replacement. There is with PH sealing ring. | Ref electrode replacement. There is with PH sealing ring. | Assisted Replacement | Electrode | Membrane | Reference
ABL810 FLEX | Messkammer und Elektrode gereinigt, erneut membraniert - Membran defekt- wiederholt, mehrere Kal, QC | Messkammer und Elektrode gereinigt, erneut membraniert - Membran defekt- wiederholt, mehrere Kal, QC | SWO-172249 | No | Measuring chamber and electrode cleaned, re-membraneed - membrane defective- repeated, several kal, QC | Measuring chamber and electrode cleaned, re-membraneed - membrane defective- repeated, several kal, QC | Assisted Replacement | Electrode | Membrane | pCO2
ABL815 FLEX | Sensor entfernt und Anschluss gereinigt, Sensor gereinigt, LS Servicejustage, Pumpenkal, Werte normal, pH instabil, mehrere Kal, QC | Sensor entfernt und Anschluss gereinigt, Sensor gereinigt, LS Servicejustage, Pumpenkal, Werte normal, pH instabil, mehrere Kal, QC | SWO-172250 | No | Sensor removed and connection cleaned, sensor cleaned, LS service adjustment, pump scale, values normal, pH unstable, multiple kal, QC | Sensor removed and connection cleaned, sensor cleaned, LS service adjustment, pump scale, values normal, pH unstable, multiple kal, QC | Defect | HW | Oxi module | LS
ABL825 FLEX | Upon arrival, several parameters were red, reference membrane needed to be changed, and the waste container was full. | Browsed through Control Panel and could not locate the Citron icon that I typically use. Reached out to a few other techs and tech support and everyone seems to think there is not a calibration in Windows 7. The touchscreen is working fine. It is assumed that the issue that the customer experienced has rectified itself. Cleaned touch screen. Changed reference membrane and waste container using customer stock. | D23B1DAF6E089936 | No | Upon arrival, several parameters were red, reference membrane needed to be changed, and the waste container was full. | Browsed through Control Panel and could not locate the Citron icon that I typically use. Reached out to a few other techs and tech support and everyone seems to think there is not a calibration in Windows 7. The touchscreen is working fine. It is assumed that the issue that the customer experienced has rectified itself. Cleaned touch screen. Changed reference membrane and waste container using customer stock. | Adjustment | HW | DMS_PC | Touch screen | Calibration
ABL800 BASIC | Replaced Lactate Membrane , Machine Working Ok | Replaced Lactate Membrane , Machine Working Ok | SWO-172361 | No | Replaced Lactate Membrane , Machine Working Ok | Replaced Lactate Membrane , Machine Working Ok | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL827 FLEX | To gather log files for the investigation of QC not cross over correctly to LIS. | Run service dump program and docklight utility tool to capture data logs. Submitted files to CQC. | D22E719ED20C8447 | Yes | To gather log files for the investigation of QC not cross over correctly to LIS. | Run service dump program and docklight utility tool to capture data logs. Submitted files to CQC. | Re classified RMED | Collect data
ABL825 FLEX | Ca++ & K+ are out of status limits. Cal. 1solution, Ca+ electrode, K+ electrode and reference had been changed by customer. All PCB's check good. Replacement reference electrode used by customer was bad too. | Replaced reference electrode and instrument calibrated normally. All QC is in control. Moved instrument to OR lab. Calibrated and ran QC. Instrument operates normally | D22DEB6FD21AA422 | No | Ca++ & K+ are out of status limits. Cal. 1solution, Ca+ electrode, K+ electrode and reference had been changed by customer. All PCB's check good. Replacement reference electrode used by customer was bad too. | Replaced reference electrode and instrument calibrated normally. All QC is in control. Moved instrument to OR lab. Calibrated and ran QC. Instrument operates normally | Assisted Replacement | Electrode | Electrode | Reference
ABL837 FLEX |  | B | D231CB5F344394C7 | No |  |  | Assisted Replacement | Electrode | Electrode | pCO2
TCM/Tosca | Touch screen will not calibrate. | Replaced the touch screen. | D22E475BD22EDB10 | No | Touch screen will not calibrate. | Replaced the touch screen. | Defect | HW | Touch Screen | Touch Screen
ABL80 FLEX BASIC | 测量仓有血迹，冲洗并做了零点校准后修复 by jesson | 测量仓有血迹，冲洗并做了零点校准后修复 by jesson | SWO-172030 | No | The measuring compartment has blood, rinsed and made a zero-point calibration to repair by jesson | The measuring compartment has blood, rinsed and made a zero-point calibration to repair by jesson
ABL800 BASIC | Cleaned The Inlet By Using Hypo chloride , Machine Working Ok | Cleaned The Inlet By Using Hypo chloride , Machine Working Ok | SWO-172360 | No | Cleaned The Inlet By Using Hypo chloride , Machine Working Ok | Cleaned The Inlet By Using Hypo chloride , Machine Working Ok | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL80 FLEX CO-OX | Kunde hatte kein LP vor Ort und keine Sensor Kassette. LP und Sensor Kassette in der GS geholt. Kontakte an der SC gereinigt, Fliesswege durchgespült. SC initialisiert, Kal und QC sind OK. Neue SC und LP beim Kunden gelassen. | Kunde hatte kein LP vor Ort und keine Sensor Kassette. LP und Sensor Kassette in der GS geholt. Kontakte an der SC gereinigt, Fliesswege durchgespült. SC initialisiert, Kal und QC sind OK. Neue SC und LP beim Kunden gelassen. | SWO-172410 | No | Customer had no LP on site and no sensor cassette. LP and sensor cassette in the GS. Contacts cleaned at the SC, flow paths rinsed through. SC initialized, Kal and QC are OK. New SC and LP left at the customer' | Customer had no LP on site and no sensor cassette. LP and sensor cassette in the GS. Contacts cleaned at the SC, flow paths rinsed through. SC initialized, Kal and QC are OK. New SC and LP left at the customer'
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-172975 | Yes | Repair after replacing test card | Repair after replacing test card
ABL80 FLEX BASIC | Sensor cassette clot clean. Liquid flow line clean. Calibration failed. Sensor cassette replace. Calibration completed. | Sensor cassette clot clean. Liquid flow line clean. Calibration failed. Sensor cassette replace. Calibration completed. | SWO-172479 | No | Sensor cassette clot clean. Liquid flow line clean. Calibration failed. Sensor cassette replace. Calibration completed. | Sensor cassette clot clean. Liquid flow line clean. Calibration failed. Sensor cassette replace. Calibration completed.
ABL835 FLEX | Fehlersuche, alle vom Anwender durchgeführten, möglichen Maßnahmen brachten keinen Erfolg, Kal.- und Aktivitätendateien überprüft, auf Serviceprogrammebene Fluidicmodulventile angesteuert ok, bei Messmodulen vor allem bei der Pumpenkal. Luftblasenbildung im Ventilblock bei pH BG Modul, Ventilmembran (Airventil im pH BG Modul) defekt->erneuert, Neustart, Pumpenkal. und LS Justage ok, Kalibration, QC und Testmessung ok | Fehlersuche, alle vom Anwender durchgeführten, möglichen Maßnahmen brachten keinen Erfolg, Kal.- und Aktivitätendateien überprüft, auf Serviceprogrammebene Fluidicmodulventile angesteuert ok, bei Messmodulen vor allem bei der Pumpenkal. Luftblasenbildung im Ventilblock bei pH BG Modul, Ventilmembran (Airventil im pH BG Modul) defekt->erneuert, Neustart, Pumpenkal. und LS Justage ok, Kalibration, QC und Testmessung ok | SWO-172411 | No | Troubleshooting, all possible measures carried out by the user did not succeed, calendar and activity files checked, controlled at the service program level fluidic module valves ok, with measuring modules especially with the pump scale. Air bubble formation in the valve block at pH BG module, valve membrane (air valve in pH BG module) defective->renewed, restart, pump scale. and LS Adjustment ok, calibration, QC and test measurement ok | Troubleshooting, all possible measures carried out by the user did not succeed, calendar and activity files checked, controlled at the service program level fluidic module valves ok, with measuring modules especially with the pump scale. Air bubble formation in the valve block at pH BG module, valve membrane (air valve in pH BG module) defective->renewed, restart, pump scale. and LS Adjustment ok, calibration, QC and test measurement ok | Defect | HW | pH/BG Module | Valve, air
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-172974 | Yes | Repair after replacing test card | Repair after replacing test card
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-172930 | Yes | Repair after replacing test card | Repair after replacing test card
ABL80 OSM |  | B | D231FAFD3401DCEB | No
ABL80 FLEX | Two Point Calibration Failure, SC not Initialising, tried to reinstall SC, Replaced SC with New one, Problem Resolve. | Two Point Calibration Failure, SC not Initialising, tried to reinstall SC, Replaced SC with New one, Problem Resolve. | SWO-175135 | No | Two Point Calibration Failure, SC not Initialising, tried to reinstall SC, Replaced SC with New one, Problem Resolve. | Two Point Calibration Failure, SC not Initialising, tried to reinstall SC, Replaced SC with New one, Problem Resolve.
ABL80 FLEX Basic version |  | B | D231FC303448DCEE | Yes |  |  | Re classified RMED | Customer training
ABL80 CO-OX | verified all quick keys load to the main screen except external QC | deleted QC and rescanned QC tried to load quick keys external QC still wont load; called tech support and they tried to load this on their 3.14 and it worked; sent pictures to Mark in tech support he is sending them to SENDX.782 | D23215DC345CA3AE | No | verified all quick keys load to the main screen except external QC | deleted QC and rescanned QC tried to load quick keys external QC still wont load; called tech support and they tried to load this on their 3.14 and it worked; sent pictures to Mark in tech support he is sending them to SENDX.782
ABL90 FLEX | Allgemeine Überprüfung - Analysator gespült - Erneuerung Einlassdichtung - Serviceprogramme durchgeführt - Funktionsüberprüfung etc. ohne Beanstandung | Mit dem Kunden noch einmal die Einlassdichtung gereinigt. Nach 10Minuten zeigt das Gerät noch eine Startphase von 50Minuten an. Da das Gerät sporadisch funktioniert ist der Kunde verunsichert und bittet um einen Service und Rückruf: 062181024320 | SWO-172569 | No | General inspection - Analyzer rinsed - Renewal inlet seal - Service programs carried out - Functional check etc. without complaint | Cleaned the inlet seal once again with the customer. After 10 minutes, the device will display a start-up period of 50 minutes. Since the device works sporadically, the customer is unsettled and asks for a service and recall: 062181024320 | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Referenzlösung in der Kassette war ausgelaufen 946-009 Lot R0579. Neue Kassette aus Schwalmstadt geholt und eingesetzt. Gerät überprüft. Vier Lösungspacks abgezogen. | Referenzlösung in der Kassette war ausgelaufen 946-009 Lot R0579. Neue Kassette aus Schwalmstadt geholt und eingesetzt. Gerät überprüft. Vier Lösungspacks abgezogen. | SWO-174116 | No | Reference solution in the cassette had expired 946-009 Lot R0579. New cassette from Schwalmstadt fetched and installed. device checked. Four solution packs subtracted. | Reference solution in the cassette had expired 946-009 Lot R0579. New cassette from Schwalmstadt fetched and installed. device checked. Four solution packs subtracted. | Assisted Replacement | SC | SC | Reference | Leaking
ABL90 FLEX | Lösungspack richtig aktiviert, Einlassdichtung erneuert, Gerät neu gestartet, Kaltstartphase abgewartet, tHb Kal, Kal und QCs durchgeführt | Lösungspack richtig aktiviert, Einlassdichtung erneuert, Gerät neu gestartet, Kaltstartphase abgewartet, tHb Kal, Kal und QCs durchgeführt | SWO-172905 | No | Solution pack correctly activated, inlet seal renewed, device restarted, cold start phase awaited, tHb Kal, Kal and QCs performed | Solution pack correctly activated, inlet seal renewed, device restarted, cold start phase awaited, tHb Kal, Kal and QCs performed | User error | SP | Activation | Activation
ABL90 FLEX | Einstellungen in Roma.ini nochmals eingegeben. Nach Neustart Scanner nochmals ab und angesteckt Scanner funktioniert jetzt. Kalibration und QC's vorgenommen. | Einstellungen in Roma.ini nochmals eingegeben. Nach Neustart Scanner nochmals ab und angesteckt Scanner funktioniert jetzt. Kalibration und QC's vorgenommen. | SWO-172609 | No | settings in Roma.ini. After rebooting scanner again and plugged in Scanner now works. calibration and QC's. | settings in Roma.ini. After rebooting scanner again and plugged in Scanner now works. calibration and QC's. | Setup | SW | Roma.ini | Ext. Barcode reader
ABL90 FLEX PLUS | Gerät geprüft, Einlassnadel (aus Kundenbestand) ersetzt, Funktionskontrolle ok | Gerät geprüft, Einlassnadel (aus Kundenbestand) ersetzt, Funktionskontrolle ok | SWO-172862 | No | Device checked, inlet pin (from customer stock) replaced, function control ok | Device checked, inlet pin (from customer stock) replaced, function control ok | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX PLUS | Preventief mixerunit vervangen. | Preventief mixerunit vervangen. | SWO-172619 | Yes | Replace the preventive mixer unit. | Replace the preventive mixer unit. | Re classified RMED | HW | Basic Unit | Sample Mixer | Agreed WR 217-3-3
ABL90 FLEX PLUS | Preventief mixer vervangen. | Preventief mixer vervangen. | SWO-172628 | Yes | Replace the preventive mixer. | Replace the preventive mixer. | Re classified RMED | HW | Basic Unit | Sample Mixer | Agreed WR 217-3-3
ABL90 FLEX PLUS | Preventief mixer vervangen. | Preventief mixer vervangen. | SWO-172632 | No | Replace the preventive mixer. | Replace the preventive mixer. | Defect | HW | Basic Unit | Sample Mixer | Preventive
ABL90 FLEX PLUS | Preventief mixer vervangen. | Preventief mixer vervangen. | SWO-172642 | Yes | Replace the preventive mixer. | Replace the preventive mixer. | Re classified RMED | HW | Basic Unit | Sample Mixer | Agreed WR 217-3-3
ABL90 FLEX PLUS | Preventief mixer vervangen. | Preventief mixer vervangen. | SWO-172776 | Yes | Replace the preventive mixer. | Replace the preventive mixer. | Re classified RMED | HW | Basic Unit | Sample Mixer | Agreed WR 217-3-3
ABL90 FLEX PLUS | Allgemeine Kontrolle ok. Lösungspack wurde ohne Erfolg ersetzt ok. Hemolyzer ers ok. Schläuche von JW Kit ersetzt ok. Thb Kal. I.o. System i.o Kal. + QC i.o. | B | SWO-172597 | No | General control ok. Solution pack has been replaced ok without success. Hemolyzer ers ok. Hoses replaced by JW Kit ok. Thb Kal. Ok. System i.o Kal. + QC i.o. |  | Defect | HW | Basic Unit | PM kit
ABL90 FLEX | 由于定标不通过导致的更换测试卡频率较高，怀疑机器存在问题，借用903-135检查机器的放大电路和电极接触是否存在问题。目前未发现问题 | 由于定标不通过导致的更换测试卡频率较高，怀疑机器存在问题，借用903-135检查机器的放大电路和电极接触是否存在问题。目前未发现问题 | SWO-172355 | No | Due to the high frequency of replacement test cards due to non-passing the calibration, a problem is suspected of the machine, and 903-135 is used to check the machine's amplification circuit and electrode contact for problems. No problems have been identified at this time | Due to the high frequency of replacement test cards due to non-passing the calibration, a problem is suspected of the machine, and 903-135 is used to check the machine's amplification circuit and electrode contact for problems. No problems have been identified at this time | Assisted Replacement | SC | SC | Unknown
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-172937 | Yes | Repair after replacing test card | Repair after replacing test card
ABL800 BASIC | uninterrupted shutdown causing A D converter fail. shutdown system for 15-20 min & reset the Oxi module. calibrate system calibration pass. guide for take precaution for power supply. | uninterrupted shutdown causing A D converter fail. shutdown system for 15-20 min & reset the Oxi module. calibrate system calibration pass. guide for take precaution for power supply. | SWO-175140 | No | uninterrupted shutdown causing A D converter fail. shutdown system for 15-20 min & reset the Oxi module. calibrate system calibration pass. guide for take precaution for power supply. | uninterrupted shutdown causing A D converter fail. shutdown system for 15-20 min & reset the Oxi module. calibrate system calibration pass. guide for take precaution for power supply. | Power off/on | HW | Oxi module | Off/On
ABL800 BASIC | Clean all the chambers,check Lac membrane status and run cleaning then 2-point calibration which is OK.Take data log for further inspection to resolve the Lac issue. | Clean all the chambers,check Lac membrane status and run cleaning then 2-point calibration which is OK.Take data log for further inspection to resolve the Lac issue. | SWO-172914 | No | Clean all the chambers,check Lac membrane status and run cleaning then 2-point calibration which is OK. Take data log for further inspection to resolve the Lac issue. | Clean all the chambers,check Lac membrane status and run cleaning then 2-point calibration which is OK. Take data log for further inspection to resolve the Lac issue. | Cleaned | HW | El/Met Module | Fluidic pathway
ABL825 FLEX | Elektrode demontiert und geprüft, Messkammern gereinigt, pO2 pCO2 Ca. K Na. Elektroden neu membraniert, Funktionskontrolle ok | Elektrode demontiert und geprüft, Messkammern gereinigt, pO2 pCO2 Ca. K Na. Elektroden neu membraniert, Funktionskontrolle ok | SWO-172976 | No | Electrode dismantled and tested, measuring chambers cleaned, pO2 pCO2 Approx. K Na. Electrodes re-membraneed, function control ok | Electrode dismantled and tested, measuring chambers cleaned, pO2 pCO2 Approx. K Na. Electrodes re-membraneed, function control ok | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX |  | B | D2321259340F3E2D | No |  |  | Cleaned | HW | pH/BG Module | Fluidic pathway
ABL800 BASIC | Cleaned the waste block & checked equipment is functional | Cleaned the waste block & checked equipment is functional | SWO-175107 | No | Cleaned the waste block & checked equipment is functional | Cleaned the waste block & checked equipment is functional | Cleaned | HW | Inlet module | Fluidic pathway
ABL800 FLEX | Option Key installiert | Option Key installiert | SWO-172904 | Yes | Option Key installed | Option Key installed | Re classified RMED | Upgrade option key |  |  | RDEU GEMBA
ABL825 FLEX | Cl. - Elektrode eingebaut und Optionkey installiert. | Cl. - Elektrode eingebaut und Optionkey installiert. | SWO-172902 | Yes | Cl. - Electrode installed and optionkey installed. | Cl. - Electrode installed and optionkey installed. | Re classified RMED | Upgrade option key |  |  | RDEU GEMBA
AQT90 FLEX |  | B | D2320B8D3400D13B | No |  |  | Defect | Consumables | Test Kit | Cartridge, D-Dimer | Broken
AQT90 FLEX | Gone through service program and found inlet wheel and Storage wheel failed. Storage wheel blocking by a broken holder. Re-fix Storage wheel after service and got fixed. Analyser serviced and calibrated properly. Sample done and found values well. Analyser Handed over in afine working condition. | Observation done. and found analyser working well. Handed over in a fine working condition. | SWO-174399 | No | Gone through service program and found inlet wheel and Storage wheel failed. Storage wheel blocking by a broken holder. Re-fix Storage wheel after service and got fixed. Analyser serviced and calibrated properly. Sample done and found values well. Analyser Handed over in afine working condition. | Observation done. and found analyser working well. Handed over in a fine working condition. | Defect | HW | Storage Module | Storage wheel
ABL80 OSM | Had customer run a tHb calibration. Results were good. System cycle and QC's are all within range. Blood results continue to run significantly higher than ABL 80 OSM right next to this analyzer. | Replaced hemolyzer for troubleshooting purposes. Ran tHb calibration. Ran patient blood. Values correlate within acceptable limits. | D2337FC2AA02CCED | No | Had customer run a tHb calibration. Results were good. System cycle and QC's are all within range. Blood results continue to run significantly higher than ABL 80 OSM right next to this analyzer. | Replaced hemolyzer for troubleshooting purposes. Ran tHb calibration. Ran patient blood. Values correlate within acceptable limits.
ABL90 |  | B | D2335A20AA3F658D | No |  |  | Replaced | HW | Inlet module | Probe
ABL90 FLEX | ABL war bei Eintreffen wieder messbereit, ABL im Service überprüft, System- und Funktionskontrolle in Ordnung | ABL war bei Eintreffen wieder messbereit, ABL im Service überprüft, System- und Funktionskontrolle in Ordnung | SWO-173046 | No | ABL was ready to be measured again on arrival, ABL checked in service, system and function control in order | ABL was ready to be measured again on arrival, ABL checked in service, system and function control in order | No fault found | No fault found | No fault found | FT
ABL90 | Found sensor cassette fully clogged. Unable to get solution through it. | Replaced sensor cassette. Flushed lines thoroughly. Cleaned spillage up on analyzer. Performed all necessary calibrations. Used solution pack to verify problem wasn't in solution pack as well. | D2334227AA21BD11 | No | Found sensor cassette fully clogged. Unable to get solution through it. | Replaced sensor cassette. Flushed lines thoroughly. Cleaned spillage up on analyzer. Performed all necessary calibrations. Used solution pack to verify problem wasn't in solution pack as well. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-174226 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Multi
AQT90 FLEX |  | B | SWO-172981 | No |  |  | Defect | HW | Incubator module | Punch up | 5297
ABL825 FLEX | kit de tubos. Susituida cámara de medida del pH. limpieza cámara del medida, circuitos neumáticos, hemolizador, filtro ventilador, limpieza exterior. Comprobación y ajuste de las bombas, válvulas, entrada, sensores de líquidos, oximetro, autocheck, impresora, electrodos, gas, temperatura. Descontaminación. Cal 2 ptos | B | D2331D94AA386EB4 | No | tube kit. Susituted pH measuring chamber. Cleaning measurement camera, pneumatic circuits, hemolizer, fan filter, exterior cleaning. Checking and adjusting pumps, valves, inlet, liquid sensors, oximeter, autocheck, printer, electrodes, gas, temperature. Decontamination. Cal 2 ptos |  | Cleaned | HW | El/Met Module | Measuring chamber
ABL825 FLEX | COHb is running lower on this analyzer on blood samples, it is also running low on QC 5.7 peer 2 is 6.0, Travel from San Rafael | replaced hemolyzer and fiber optic cable; cable replaced old one was crimped; could not adjust lamp after hemolyzer replacement replaced lamp unit replaced the electrode and solution pump tubings | D2336232AA0B02A5 | Yes | COHb is running lower on this analyzer on blood samples; it is also running low on QC 5.7 peer 2 is 6.0, Travel from San Rafael | replaced hemolyzer and fiber optic cable; cable replaced old one was crimped; could not adjust lamp after hemolyzer replacement replaced lamp unit replaced the electrode and solution pump tubings
ABL825 FLEX | Bildschirm geprüft, Leihbildschirm installiert, BG Messkammer gereinigt, Ref. Elektrode neu membraniert, Dichtring an der PH Elektrode ersetzt,Funktionskontrolle ok | Der Kunde wünscht den Tausch des Touchsreens. Der Kunde wünscht das die Jahreswartung vorgezogen wird. Der Kunde bittet um einen Service. | SWO-174686 | No | Screen checked, rental screen installed, BG measuring chamber cleaned, ref. Electrode re-membraneed, sealing ring replaced at THE PH electrode,function control ok | The customer wishes to exchange the touch. The customer wishes that the annual maintenance is preferred. The customer asks for a service. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | issue while entering patient data. | REPLACE DISPLAY SCREEN UNIT. TOUCH SCREEN FUNCTIONS AS EXPECTED. | D23888C3824B8056 | No | issue while entering patient data. | REPLACE DISPLAY SCREEN UNIT. TOUCH SCREEN FUNCTIONS AS EXPECTED. | Defect | HW | DMS_PC | Touch screen
ABL820 FLEX | pCO2 low on level 1 quality control. pCO2 also drifts on calibration intermittently. David and I have tried new pCO2 electrodes also. This is an ABL820 at software version 6.16. | Replaced pCO2 and Ref membranes. Cleaned electrode chamber. | D2331C26AA0B1E81 | No | pCO2 low on level 1 quality control. pCO2 also drifts on calibration intermittently. David and I have tried new pCO2 electrodes also. This is an ABL820 at software version 6.16. | Replaced pCO2 and Ref membranes. Cleaned electrode chamber. | Assisted Replacement | Electrode | Membrane | pCO2
AQT90 FLEX | Observation done. and found analyser working well. Handed over in a fine working condition. | Observation done. and found analyser working well. Handed over in a fine working condition. | SWO-174401 | No | Observation done. and found analyser working well. Handed over in a fine working condition. | Observation done. and found analyser working well. Handed over in a fine working condition. | No fault found | Courtesy | Courtesy | Courtesy
AQT90 FLEX | Fehlende QK-Daten in der Konfiguration angelegt, Schnittstellenkonfiguration geprüft, es ist davon auszugehen, dass ein Anwenderfehler vorliegt -> da fehlende QK mglw. als Patient gemessen wurden Zusätzlich sind sehr viele Verbindungsabbrüche und entsprechende Wiederherstellungen der Verbindung vom AQT zu CobasIT dokumentiert! Hinweis an Kunden über Anlage von QK !! | B | SWO-174188 | No | Missing QK data created in the configuration, interface configuration checked, it is to be assumed that there is a user error -> since missing QK mglw. measured as a patient In addition, a lot of disconnections and corresponding restores of the connection from AQT to CobasIT are documented! Note to customers about qK plant !! |  | Setup | His/lis | His/lis | His/lis
ABL80 FLEX BASIC |  | B | SWO-174218 | No
ABL80 FLEX BASIC | Temp error, during SC installation, found SC was not connected to analyzer, removed the SC reinstalled the SC found Air detection calibration failed, inactivated the Air detection, recalibrated SC found all the parameter was working, found instrument is functional. | Temp error, during SC installation, found SC was not connected to analyzer, removed the SC reinstalled the SC found Air detection calibration failed, inactivated the Air detection, recalibrated SC found all the parameter was working, found instrument is functional. | SWO-181350 | No | Temp error, during SC installation, found SC was not connected to analyzer, removed the SC reinstalled the SC found Air detection calibration failed, inactivated the Air detection, recalibrated SC found all the parameter was working, found instrument is functional. | Temp error, during SC installation, found SC was not connected to analyzer, removed the SC reinstalled the SC found Air detection calibration failed, inactivated the Air detection, recalibrated SC found all the parameter was working, found instrument is functional.
ABL90 FLEX | Abfallschlauch und Ventilschlauch überprüft - ok. Rollenkranz erneuert (aus Neugeräte ausgebaut, da nicht im Bestand). SC und LP aus Kundenbestand erneuert. Kalibrationen und QC's vorgenommen. | Abfallschlauch und Ventilschlauch überprüft - ok. Rollenkranz erneuert (aus Neugeräte ausgebaut, da nicht im Bestand). SC und LP aus Kundenbestand erneuert. Kalibrationen und QC's vorgenommen. | SWO-174422 | No | Waste hose and valve hose checked - ok. Roller wreath renewed (removed from new equipment, as not in stock). SC and LP from customer base. calibrations and QC's. | Waste hose and valve hose checked - ok. Roller wreath renewed (removed from new equipment, as not in stock). SC and LP from customer base. calibrations and QC's. | Unknown | Unknown | Unknown | Unknown
ABL90 |  | B | D234AE952027F2E9 | No |  |  | Replaced | HW | Inlet module | Probe
ABL90 FLEX PLUS |  | B | D234D0F620410E28 | Yes |  |  | Re classified RMED | DEMO Analyzer
ABL90 FLEX |  | B | SWO-174302 | No |  |  | Defect | HW | DMS_PC | DMS_PC
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-174208 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Lac | Response
ABL90 FLEX | 连续更换一个测试卡和一个试剂包修复 | 连续更换一个测试卡和一个试剂包修复 | SWO-176054 | Yes | Replace one test card and one reagent kit repair in a row | Replace one test card and one reagent kit repair in a row
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-177952 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Reference | 1361
AQT90 FLEX |  | B | SWO-174238 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX |  | B | SWO-174230 | No |  |  | Assisted Replacement | Consumables | Test Kit | Cartridge | Myo
AQT90 FLEX |  | B | SWO-174268 | No |  |  | Defect | HW | Analyzer Control Module | PCB, Main Board | 3412
ABL825 FLEX |  | B | D2348A2F2019D135 | Yes
ABL815 FLEX | Bildschirm aus Altgerät eingesetzt. Kalibriert und Töne konfiguriert. | Bildschirm aus Altgerät eingesetzt. Kalibriert und Töne konfiguriert. | SWO-174697 | No | screen made of old equipment. Calibrated and configured totones. | screen made of old equipment. Calibrated and configured totones. | Defect | HW | DMS_PC | Touch screen
ABL815 FLEX | Servicedump gezogen, Lac gereinigt und neu membraniert, akz. Kal abgewartet, Kal und QC i.O. | Servicedump gezogen, Lac gereinigt und neu membraniert, akz. Kal abgewartet, Kal und QC i.O. | SWO-174544 | No | Service dump pulled, Lac cleaned and re-membraneed, akz. Kal waited, Kal and QC i.O. | Service dump pulled, Lac cleaned and re-membraneed, akz. Kal waited, Kal and QC i.O. | Assisted Replacement | Electrode | Membrane | Lac | RDEU GEMBA
ABL825 FLEX | Oxy Modul ersetzt, oxy complete test, Kalibrationen, Validierung der tHb Kalibration, QC Kontrolle | Oxy Modul ersetzt, oxy complete test, Kalibrationen, Validierung der tHb Kalibration, QC Kontrolle | SWO-174516 | No | Oxy module replaced, oxy complete test, calibrations, validation of tHb calibration, QC control | Oxy module replaced, oxy complete test, calibrations, validation of tHb calibration, QC control | Defect | HW | Oxi module | Oxi module | Neon/Spectrometer
ABL827 FLEX | Oxi hardware error, Neon lamp voltage too high. | Swapped out addition oxi parts with no success. Re-installed original parts- Lamp voltage checking OK. Issues return within 3 hours. Probable oxi PCB needed - item on BO. Reseated PCB and module and adjusted lamp. All checking OK again | D2347CD32039DB26 | No | Oxi hardware error, Neon lamp voltage too high. | Swapped out addition oxi parts with no success. Re-installed original parts- Lamp voltage checking OK. Issues return within 3 hours. Probable oxi PCB needed - item on BO. Reseated PCB and module and adjusted lamp. All checking OK again | Defect | HW | Oxi module | Lamp unit | Neon
ABL815 FLEX | Servicdump gezogen, mehrfach kalibriert, Lac neu membraniert, akz. Kal abgewartet, Kal u. QC i.O. | Servicdump gezogen, mehrfach kalibriert, Lac neu membraniert, akz. Kal abgewartet, Kal u. QC i.O. | SWO-174545 | No | Servicdump pulled, multiple calibrated, Lac re-membraneed, akz. Kal waited, Kal & QC i.O. | Servicdump pulled, multiple calibrated, Lac re-membraneed, akz. Kal waited, Kal & QC i.O. | Assisted Replacement | Electrode | Membrane | Lac | RDEU GEMBA
ABL825 FLEX |  | B | D2345D65203C8EE0 | No |  |  | Defect | HW | Gas module | Gas2
AQT90 FLEX |  | B | D23475BE20521762 | No |  |  | Defect | HW | Aspiration Module | P5 | FTC 16
AQT90 FLEX | Probenhalterung getauscht weil gebrochen. Systemüberprüfung durchgeführt. Mehrere Systemreinigungen durchgeführt. QC Level 2 Multicheck durchgeführt --> i.O. AQT wieder messbereit. | Probenhalterung getauscht weil gebrochen. Systemüberprüfung durchgeführt. Mehrere Systemreinigungen durchgeführt. QC Level 2 Multicheck durchgeführt --> i.O. AQT wieder messbereit. | SWO-174355 | No | Sample holder replaced because broken. system verification. Several system cleanings performed. QC Level 2 Multicheck performed --> i.O. AQT ready to measure again. | Sample holder replaced because broken. system verification. Several system cleanings performed. QC Level 2 Multicheck performed --> i.O. AQT ready to measure again. | Defect | HW | Aspiration Module | Tube holder
AQT90 FLEX |  | B | D2347F3D205D53E4 | No |  |  | Assisted Replacement | Consumables | Test Kit | CRP
ABL80 FLEX BASIC | Problem with waste drain, cleaned the waste drain, restarted instrument found instrument is functional. | Problem with waste drain, cleaned the waste drain, restarted instrument found instrument is functional. | SWO-181352 | No | Problem with waste drain, cleaned the waste drain, restarted instrument found instrument is functional. | Problem with waste drain, cleaned the waste drain, restarted instrument found instrument is functional.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-174557 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 | found that they are failing internal and external qc for OXI parameters | back flushed and preformed thb cal on analyzer to correct the issue. | D235CCFF96025412 | No | found that they are failing internal and external qc for OXI parameters | back flushed and preformed thb cal on analyzer to correct the issue. | Adjustment | HW | Oxi module | tHb Cal
ABL90 | Sensors cassettes were failing multiple parameters on calibration. Replacements were ordered through tech support. | Received replacement cassettes and will submit consumable credits. | D23768150C384EFB | Yes | Sensors cassettes were failing multiple parameters on calibration. Replacements were ordered through tech support. | Received replacement cassettes and will submit consumable credits. | Re classified RMED | PER action
ABL90 FLEX PLUS | Strømforsyning udskiftet, slutafprøvet. | Strømforsyning udskiftet, slutafprøvet. | SWO-174673 | No | Power supply replaced, final tested. | Power supply replaced, final tested. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX |  | B | SWO-174584 | No |  |  | Defect | HW | Oxi module | Hemolyzer | 134
ABL825 FLEX | Nyomtató cseréje, Glu, Lac membrán cserék, Mérőkamra tisztítása, Oktatás (QC mérések), Kalibrálás, QC mérés | B | SWO-174665 | No | Printer replacement, Glu, Lac membrane replacements, measuring chamber cleaning, training (QC measurements), calibration, QC measurement |  | Defect | HW | Printer Module | Printer Module
ABL827 FLEX |  | B | D235B78196293E49 | Yes
ABL827 FLEX |  | B | D235CD32962D823A | No |  |  | Setup | Replacement schedule | Replacement schedule | Replacement schedule
AQT90 FLEX | Analyzer Showing Error Sate safe_22, Cartidge not placed Properly, Remove and reinstalled Cartridge, passed setsafe test, Restart the Analyzer, problem resolve. | Analyzer Showing Error Sate safe_22, Cartidge not placed Properly, Remove and reinstalled Cartridge, passed setsafe test, Restart the Analyzer, problem resolve. | SWO-175144 | No | Analyzer Showing Error Sate safe_22, Cartidge not placed Properly, Remove and reinstalled Cartridge, passed setsafe test, Restart the Analyzer, problem resolve. | Analyzer Showing Error Sate safe_22, Cartidge not placed Properly, Remove and reinstalled Cartridge, passed setsafe test, Restart the Analyzer, problem resolve. | User error | Consumables | Test Kit | Cartridge | Set_Safe_22
ABL90 FLEX PLUS | Einlassdichtung (Kundenbestand) und Einlasszwischendichtung (Kundenbestand) erneuert, System und Funktion i.O. | Service Einsatz von Dr. Weingärtner gewünscht | SWO-174964 | No | Inlet seal (customer stock) and inlet intermediate seal (customer stock) renewed, system and function as it is | Service use desired by Dr. Weingärtner | Defect | HW | Auto Inlet module | Connector Gasket
ABL90 FLEX PLUS | LP aus GS besorgt und geliefert, | B | SWO-175113 | Yes | LP from GS, |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Lieferung u. Installation LP, Kal u. QC i.O. | B | SWO-175114 | Yes | Lieferung u. Installation LP, Kal u. QC i.O. |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Sostituito emolizzatore, apparecchio completamente ricalibrato e testato. | Sostituito emolizzatore, apparecchio completamente ricalibrato e testato. | SWO-175257 | No | Replaced molizer, fully recalibrated and tested device. | Replaced molizer, fully recalibrated and tested device. | Defect | HW | Oxi module | Hemolyzer | 134
AQT90 FLEX |  | B | SWO-174860 | No |  |  | Assisted Replacement | Consumables | Test Kit | Cartridge | PCT, NT-proBNP
AQT90 FLEX | 更换备用光学单元903-364（903-364R0231N008）正常。 | B | SWO-175212 | No | Replacing the backup optical unit 903-364 (903-364R0231N008) is normal. |  | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX |  | B | SWO-174888 | No |  |  | Defect | HW | Process Module | Membrane Pump | 168
AQT90 FLEX | 903-138后修复 | 903-138后修复 | SWO-174566 | No | 903-138 Post-Fix | 903-138 Post-Fix | Defect | HW | Incubator module | PCB, Incubator
ABL80 CO-OX |  | B | D237008F0C292611 | No
ABL825 FLEX | Glu. Dichtring und Membran ersetzt. Reinigung, Proteinentfernung und 1.Kal., tHb Kal. durchgeführt, Kal. tHb Test und QCs i.O. 2x 834-618 Dichtring als Ersatz vor Ort gelassen. | Glu. Dichtring und Membran ersetzt. Reinigung, Proteinentfernung und 1.Kal., tHb Kal. durchgeführt, Kal. tHb Test und QCs i.O. 2x 834-618 Dichtring als Ersatz vor Ort gelassen. | SWO-175365 | No | Glu. sealing ring and membrane. Cleansing, protein removal and 1.Cal., tHb Kal. carried out, Kal. tHb test and QCs in the o. 2x 834-618 sealing ring as replacement left on site. | Glu. sealing ring and membrane. Cleansing, protein removal and 1.Cal., tHb Kal. carried out, Kal. tHb test and QCs in the o. 2x 834-618 sealing ring as replacement left on site. | Cleaned | HW | Oxi module | Oxi module | tHb
ABL835 FLEX | Touchscreen erneuert und kalibriert Gerät überprüft | Touchscreen erneuert und kalibriert Gerät überprüft | SWO-175073 | No | Touchscreen renewed and calibrated Device checked | Touchscreen renewed and calibrated Device checked | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | O2elektrode erneuert. Referenz Elektrode aus kundenbestand erneuert. Mehrfach kalibriert. Ca lässt sich nicht einkalibrieren. Ca Elektrode erneuert. Elektrode muss noch einkalibriert werden. Station wird fals erforderlich sich am Montag erneut melden. | O2elektrode erneuert. Referenz Elektrode aus kundenbestand erneuert. Mehrfach kalibriert. Ca lässt sich nicht einkalibrieren. Ca Elektrode erneuert. Elektrode muss noch einkalibriert werden. Station wird fals erforderlich sich am Montag erneut melden. | SWO-175096 | No | O2electrode renewed. Reference electrode renewed from customer stock. Multiple calibrated. Ca cannot be calibrated. Ca electrode renewed. Electrode has yet to be calibrated. Station will report fals required again on Monday. | O2electrode renewed. Reference electrode renewed from customer stock. Multiple calibrated. Ca cannot be calibrated. Ca electrode renewed. Electrode has yet to be calibrated. Station will report fals required again on Monday. | Assisted Replacement | Electrode | Electrode | pO2
ABL837 FLEX |  | B | D23AB5A16E42CB04 | No |  |  | Defect | HW | Auto Check Module | Auto Check Module
ABL815 FLEX | Elektrodenkappe Referenzelektrode gereinigt, Meßkammer und pH Elektrode gesäubert, O-Ring an der pH Elektrodenkappe und Einlassnadel ersetzt. Funktion z.Zt. OK. | Der Kunde bittet um einen Service. | SWO-175255 | No | Electrode cap reference electrode cleaned, measuring chamber and pH electrode cleaned, O-ring replaced at the pH electrode cap and inlet needle. Function currently OK. | The customer asks for a service. | Defect | HW | Inlet module | Probe bend
AQT90 FLEX | Defecte cartridge verwijdert. | B | SWO-175007 | No | Defective cartridge. |  | Defect | Consumables | Test Kit | Cartridge | Broken
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-175221 | Yes | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | Need to replace signal data board | B | SWO-175986 | No | Need to replace signal data board
ABL815 FLEX | Mit Fr. Hintzen die Einlassdichtung erneuert. Nadel überprüft und die beiden Schläuche gespült. Einlass jetzt trocken. | B | SWO-175190 | No | With Fr. Hintzen the inlet seal is renewed. needle and rinsed the two hoses. Entrance now dry. |  | Defect | HW | Inlet module | Inlet Gasket
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-175327 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC |  | B | SWO-175286 | No
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-175705 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX | Rebooted analyser, still same Removed from department to replace display with new Spare analyser put in place New screen fitted and analyser put back in ICU New inlet clip fitted on both analysers | Rebooted analyser, still same Removed from department to replace display with new Spare analyser put in place New screen fitted and analyser put back in ICU New inlet clip fitted on both analysers | SWO-175753 | No | Rebooted analyser, still same Removed from department to replace display with new Spare analyser put in place New screen fitted and analyser put back in ICU New inlet clip fitted on both analysers | Rebooted analyser, still same Removed from department to replace display with new Spare analyser put in place New screen fitted and analyser put back in ICU New inlet clip fitted on both analysers | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX | Bei Ankunft kein Fehler mehr vorhanden. LP und Einlass nochmals überprüft. Keine Auffälligkeiten. Kal. , QC und tHb Kal. durchgeführt. | Bei Ankunft kein Fehler mehr vorhanden. LP und Einlass nochmals überprüft. Keine Auffälligkeiten. Kal. , QC und tHb Kal. durchgeführt. | SWO-175994 | No | There is no more error on arrival. LP and inlet checked again. No abnormalities. Kal. , QC and tHb Kal. Performed. | There is no more error on arrival. LP and inlet checked again. No abnormalities. Kal. , QC and tHb Kal. Performed. | No fault found | No fault found | No fault found | No fault found
TCM5 FLEX |  | B | D23FBCA646501D98 | No
AQT90 FLEX | 更换903-329一个正常。(By Ran) | 更换903-329一个正常。(By Ran) | SWO-175149 | No | Replace the 903-329 one normal. (By Ran) | Replace the 903-329 one normal. (By Ran) | Defect | HW | Aspiration Module | Needle Wash Unit
ABL820 FLEX | Reviewed logs and same as last visit no QC has failed sense 1 16 2017 | Tried a test electrode and values were lower. replaced REF and Pc02 membranes with client stock no change. Put client's electrode back in service, and de-bubbled it - now recovering at 69.6. Will order another electrode and test it once it comes in, if values are still dropping. Tech line case is already open. | D23AC2616E24DDA8 | No | Reviewed logs and same as last visit no QC has failed sense 1 16 2017 | Tried a test electrode and values were lower. replaced REF and Pc02 membranes with client stock no change. Put client's electrode back in service, and de-bubbled it - now recovering at 69.6. Will order another electrode and test it once it comes in, if values are still dropping. Tech line case is already open. | Revisit required | Electrode | Electrode | pCO2 | Technline
ABL825 FLEX | Proteinentfernung durchgeführt - Messkammer gereinigt Ventile gängig gemacht, Ventilmembranen der Luftventile erneuert Gerät kalibriert und überprüft | Proteinentfernung durchgeführt - Messkammer gereinigt Ventile gängig gemacht, Ventilmembranen der Luftventile erneuert Gerät kalibriert und überprüft | SWO-175573 | No | Protein removal performed - measuring chamber cleaned Valves made common, valve membranes of the air valves renewed Device calibrated and checked | Protein removal performed - measuring chamber cleaned Valves made common, valve membranes of the air valves renewed Device calibrated and checked | Defect | HW | El/Met Module | Valve, air | Membrane
ABL825 FLEX | Found the electrode connector missing the spring and could not find the spring. | Replaced electrode connector. Also ran zero current calibration due to a high Pa reading on LAC. Customer was not available when repair was done as he was at main campus, no signature, spoke with Carl Romero upon completion. | D243CEF0A8283E1E | No | Found the electrode connector missing the spring and could not find the spring. | Replaced electrode connector. Also ran zero current calibration due to a high Pa reading on LAC. Customer was not available when repair was done as he was at main campus, no signature, spoke with Carl Romero upon completion. | User error | HW | pH/BG Module | Electrode connector
AQT90 FLEX |  | B | D23AA0C06E5AE5F3 | No |  |  | Defect | HW | Cover and Chassis | Front Cover
AQT90 FLEX | Nadel bestellt. | Nadel erneuert und justiert. | SWO-175754 | No | Needle ordered. | Needle renewed and adjusted. | Revisit required | HW | Aspiration Module | Bend needle | 4794
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-176254 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-175658 | Yes | The replacement test card is OK. | The replacement test card is OK.
ABL90 |  | B | D23C10A3E4215C05 | No |  |  | Defect | HW | Inlet module | Probe bend
ABL90 FLEX PLUS |  | B | D23BF727E45B6F4F | No |  |  | Defect | SW | Windows | Communication | V3.1 MR5
ABL90 FLEX PLUS | Gerät gespült, defekten Ventilschlauch ersetzt, checkout and adjustment, Gerät kalibriert und mit QCs überprüft,Ok | Gerät gespült, defekten Ventilschlauch ersetzt, checkout and adjustment, Gerät kalibriert und mit QCs überprüft,Ok | SWO-175985 | No | Device rinsed, defective valve hose replaced, checkout and adjustment, device calibrated and checked with QCs,Ok | Device rinsed, defective valve hose replaced, checkout and adjustment, device calibrated and checked with QCs,Ok | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX |  | 更换903-197修复 by jesson | SWO-175643 | No |  | Replace 903-197 Fix by Jesson | Defect | HW | Oxi module | Spectrometer | 134
ABL90 | Chequeo de conexiones, cable y de configuración observando que la puerta de enlace ha cambiado y no es la correcta, se configura conecta a Radiance, configuración de vista de registro de pacientes ya que quieren que se vea la IDMuestra. prueba de envío de muestras, pcalibración y pasado qc. | configuración de puerta de enlace, configuración de roma.ini | D23C32B7E461AFE7 | No | Checking connections, cable, and configuration by observing that the gateway has changed and is not correct, it is configured connect to Radiance, patient log view settings since they want the SampleID to be seen. sample submission test, precalibration and past qc. | gateway configuration, roma.ini configuration | Setup | His/lis | Roma.ini | Radiance
ABL90 FLEX | 更换测试卡及试剂包后修复 | 更换测试卡及试剂包后修复 | SWO-175690 | Yes | Repair after replacing test cards and reagent kits | Repair after replacing test cards and reagent kits
ABL827 FLEX |  | B | D23BF5A6E40EABDE | No |  |  | Defect | HW | Oxi module | Oxi module
ABL800 FLEX | Replaced printer head from a working ABL800, printer is working now. | Replaced printer head from a working ABL800, printer is working now. | SWO-175693 | No | Replaced printer head from a working ABL800, printer is working now. | Replaced printer head from a working ABL800, printer is working now. | Defect | HW | Printer Module | Printer head
ABL825 FLEX |  | B | D23C16C3E4488855 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL835 FLEX | Hemolyzer udskiftet, opjusteret, service tHb kalibreret, slutafprøvet. | B | SWO-175772 | No | Hemolyzer replaced, upwards adjusted, service tHb calibrated, final tested. |  | Defect | HW | Oxi module | Hemolyzer
ABL827 FLEX | checked over the analyser and but experienced no issues whilst investigating on site. decided to run a screen calibration within the windows control panel. advised customer to monitor and report back if problem persists. | checked over the analyser and but experienced no issues whilst investigating on site. decided to run a screen calibration within the windows control panel. advised customer to monitor and report back if problem persists. | SWO-175937 | No | checked over the analyser and but experienced no issues whilst investigating on site. decided to run a screen calibration within the windows control panel. advised customer to monitor and report back if problem persists. | checked over the analyser and but experienced no issues whilst investigating on site. decided to run a screen calibration within the windows control panel. advised customer to monitor and report back if problem persists. | Defect | HW | DMS_PC | Touch screen
ABL827 FLEX | Arrived on site and attended analyser. found that the analyser had some qc staining on the auto check cover which indicated it may have been leaking from the inlet gasket after the QC run that morning. opened the analyser and ran a qc then monitored the rinse afterwards. no leak coming from the gasket. checked the tubings for blockages or perforations. found none. flushed the waste lines and soaked with protein remover to break down any possible clots or blockages. checked for debris being flushed out of the waste bottle nipple onto the tissue placed below it. none removed. checked the inlet by removing the gasket and inspecting the seal. this looked in good condition. inspected the inlet probe and the blue inlet probe tubing. these both looked good. calibrated the inlet probe as a precaution. refitted all parts and ran a number of rinses in a row to check the rinse procedure and monitor the inlet for leakage. none found and all looked to be in working order. tried clearing the blue inlet nipple on the fluidic module to ensure this was clear. upon rinsing, there was no fluid drawn in to the nipple. inspected the valve operation. this did not seem correct. checked in service programs for any flagged failures. none shown. powered down the device and removed inspected and cleaned all valves for signs of perforation or blockage. refitted all parts and powered up the analyser. set all settings to default and retested the blue inlet nipple. this now took liquid in and rinsed as it should. suspect small blockage in waste valve. calibrated and qcd analyser. customer will monitor. | Arrived on site and attended analyser. found that the analyser had some qc staining on the auto check cover which indicated it may have been leaking from the inlet gasket after the QC run that morning. opened the analyser and ran a qc then monitored the rinse afterwards. no leak coming from the gasket. checked the tubings for blockages or perforations. found none. flushed the waste lines and soaked with protein remover to break down any possible clots or blockages. checked for debris being flushed out of the waste bottle nipple onto the tissue placed below it. none removed. checked the inlet by removing the gasket and inspecting the seal. this looked in good condition. inspected the inlet probe and the blue inlet probe tubing. these both looked good. calibrated the inlet probe as a precaution. refitted all parts and ran a number of rinses in a row to check the rinse procedure and monitor the inlet for leakage. none found and all looked to be in working order. tried clearing the blue inlet nipple on the fluidic module to ensure this was clear. upon rinsing, there was no fluid drawn in to the nipple. inspected the valve operation. this did not seem correct. checked in service programs for any flagged failures. none shown. powered down the device and removed inspected and cleaned all valves for signs of perforation or blockage. refitted all parts and powered up the analyser. set all settings to default and retested the blue inlet nipple. this now took liquid in and rinsed as it should. suspect small blockage in waste valve. calibrated and qcd analyser. customer will monitor. | SWO-175940 | No | Arrived on site and attended analyser. found that the analyser had some qc staining on the auto check cover which indicated it may have been leaking from the inlet gasket after the QC run that morning. opened the analyser and ran a qc then monitored the rinse afterwards. no leak coming from the gasket. checked the tubings for blockages or perforations. found none. flushed the waste lines and soaked with protein remover to break down any possible clots or blockages. checked for debris being flushed out of the waste bottle nipple onto the tissue placed below it. none removed. checked the inlet by removing the gasket and inspecting the seal. this looked in good condition. inspected the inlet probe and the blue inlet probe tubing. these both looked good. calibrated the inlet probe as a precaution. refitted all parts and ran a number of rinses in a row to check the rinse procedure and monitor the inlet for leakage. none found and all looked to be in working order. tried clearing the blue inlet nipple on the fluidic module to ensure this was clear. upon rinsing, there was no fluid drawn in to the nipple. inspected the valve operation. this did not seem correct. checked in service programs for any flagged failures. none shown. powered down the device and removed inspected and cleaned all valves for signs of perforation or blockage. refitted all parts and powered up the analyser. set all settings to default and retested the blue inlet nipple. this now took liquid in and rinsed as it should. suspect small blockage in waste valve. calibrated and qcd analyser. customer will monitor. | Arrived on site and attended analyser. found that the analyser had some qc staining on the auto check cover which indicated it may have been leaking from the inlet gasket after the QC run that morning. opened the analyser and ran a qc then monitored the rinse afterwards. no leak coming from the gasket. checked the tubings for blockages or perforations. found none. flushed the waste lines and soaked with protein remover to break down any possible clots or blockages. checked for debris being flushed out of the waste bottle nipple onto the tissue placed below it. none removed. checked the inlet by removing the gasket and inspecting the seal. this looked in good condition. inspected the inlet probe and the blue inlet probe tubing. these both looked good. calibrated the inlet probe as a precaution. refitted all parts and ran a number of rinses in a row to check the rinse procedure and monitor the inlet for leakage. none found and all looked to be in working order. tried clearing the blue inlet nipple on the fluidic module to ensure this was clear. upon rinsing, there was no fluid drawn in to the nipple. inspected the valve operation. this did not seem correct. checked in service programs for any flagged failures. none shown. powered down the device and removed inspected and cleaned all valves for signs of perforation or blockage. refitted all parts and powered up the analyser. set all settings to default and retested the blue inlet nipple. this now took liquid in and rinsed as it should. suspect small blockage in waste valve. calibrated and qcd analyser. customer will monitor. | Cleaned | Fluidic pathway | Inlet module | Inlet module
AQT90 FLEX | Nadel erneuert und justiert. | Nadel erneuert und justiert. | SWO-175755 | No | Needle renewed and adjusted. | Needle renewed and adjusted. | Defect | HW | Aspiration Module | Bend needle
ABL80 FLEX BASIC | 连续更换2张测试卡后修复 | 连续更换2张测试卡后修复 | SWO-177455 | No | Repair after 2 consecutive test cards | Repair after 2 consecutive test cards
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-179664 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | FT | 181
ABL835 FLEX |  | B | D23D80A05A185C0F | Yes
ABL815 FLEX | Lac neu membraniert mit D7077, Lot R3473, akz. Kal abgewartet, Kal u. QC i.O. | Lac neu membraniert mit D7077, Lot R3473, akz. Kal abgewartet, Kal u. QC i.O. | SWO-176103 | No | Lac re-membraneed with D7077, lot R3473, akz. Kal waited, Kal & QC i.O. | Lac re-membraneed with D7077, lot R3473, akz. Kal waited, Kal & QC i.O. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL815 FLEX | Lac neu membraniert (Lot R 3473), akz. Kal abgewartet, Kal u. QC i.O. | Lac neu membraniert (Lot R 3473), akz. Kal abgewartet, Kal u. QC i.O. | SWO-176154 | No | Lac re-membraneed (Lot R 3473), akz. Kal waited, Kal & QC i.O. | Lac re-membraneed (Lot R 3473), akz. Kal waited, Kal & QC i.O. | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL835 FLEX | Haken bei Kalibration nach jeder Messung raus genommen. Glu und Lac neu membraniert. Mehrfach Kalibriert. Optionkeys installiert. Fehlersuche an der Netzwerkverbindung durchgeführt, fehler der Kis Verbindung nicht beheben können. | Mit Hilfe von AleSp die Radianceeinstellungen über Remote vorgenommen und dem Leitenden Kardiotechniker Herrn Schreiber telefonisch informiert. Herr Schreiber wird sich wieder melden falls ihm noch etwas auffällt. | SWO-176341 | No | Hook scalding taken out after each measurement. Glu and Lac. Multiple calibrated. Optionkeys installed. Troubleshooting the network connection, the Kis connection fails. | With the help of AleSp, the radiance settings were made remotely and informed mr. Schreiber by telephone. Mr Schreiber will come back if he notices anything else. | Setup | His/lis | His/lis | His/lis
AQT90 FLEX | CPU Unit in der GS eingebaut. Startphase abgewartet, NT-pro BNP kalibriert. | CPU Unit in der GS eingebaut. Startphase abgewartet, NT-pro BNP kalibriert. | SWO-177501 | No | CPU unit built into the GS. start-up phase, NT-pro BNP calibrated. | CPU unit built into the GS. start-up phase, NT-pro BNP calibrated. | Defect | HW | Analyzer Control Module | CPU
ABL80 FLEX BASIC | 更换902-972一个正常。(By Ran) | 更换902-972一个正常。(By Ran) | SWO-175727 | No | Replace 902-972 one normal. (By Ran) | Replace 902-972 one normal. (By Ran)
ABL80 FLEX BASIC | 更换备用机件902-972后修复(By Ran) | 更换备用机件902-972后修复(By Ran) | SWO-177482 | No | Repair after replacement of spare parts 902-972 (By Ran) | Repair after replacement of spare parts 902-972 (By Ran)
ABL80 FLEX BASIC | 更换测试卡及换试剂包后修复 | 更换测试卡及换试剂包后修复 | SWO-177459 | Yes | Repair after replacing test cards and reagent kits | Repair after replacing test cards and reagent kits
ABL80 FLEX | Leakage from SC80 while installation, Sensor connector get Defective, clean Sensor Board and Tried to reinstall, Problem Remains, Need Sensor Board, SC80 PER need to sent. waiting for Spare. | B | SWO-177820 | No | Leakage from SC80 while installation, Sensor connector get Defective, clean Sensor Board and Tried to reinstall, Problem Remains, Need Sensor Board, SC80 PER need to sent. waiting for Spare.
ABL90 FLEX PLUS | Reseated inlet connector gasket and Oximetry pinch valve tube. Cleaned Sensor Cassette bay contact pins. | Reseated inlet connector gasket and Oximetry pinch valve tube. Cleaned Sensor Cassette bay contact pins. | SWO-177697 | No | Reseated inlet connector gasket and Oximetry pinch valve tube. Cleaned Sensor Cassette bay contact pins. | Reseated inlet connector gasket and Oximetry pinch valve tube. Cleaned Sensor Cassette bay contact pins. | Defect | HW | Oxi module | Pinch valve tube/blood
ABL90 FLEX PLUS | Replaced Oximetry pinch valve tube. | Replaced Oximetry pinch valve tube. | SWO-177698 | No | Replaced Oximetry pinch valve tube. | Replaced Oximetry pinch valve tube. | Defect | HW | Oxi module | Pinch valve tube/blood
ABL90 FLEX PLUS | Reseated Oximetry pinch valve tube. | Reseated Oximetry pinch valve tube. | SWO-177696 | No | Reseated Oximetry pinch valve tube. | Reseated Oximetry pinch valve tube. | Defect | HW | Oxi module | Pinch valve tube/blood
ABL90 FLEX PLUS | Gerät aus dem OP abgezogen, Hemolyzer erneuert, checkout and adjustment, Gerät kalibriert, mit QCs überprüft und div. Blutmessungen durchgeführt,OK. defekt: 908-552R1841N019 | Gerät aus dem OP abgezogen, Hemolyzer erneuert, checkout and adjustment, Gerät kalibriert, mit QCs überprüft und div. Blutmessungen durchgeführt,OK. defekt: 908-552R1841N019 | SWO-177427 | No | Device removed from the operating room, Hemolyzer renewed, checkout and adjustment, device calibrated, checked with QCs and various blood measurements performed,OK. defective: 908-552R1841N019 | Device removed from the operating room, Hemolyzer renewed, checkout and adjustment, device calibrated, checked with QCs and various blood measurements performed,OK. defective: 908-552R1841N019 | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | 更换903-183一个正常。(By Ran) | 更换903-183一个正常。(By Ran) | SWO-176019 | No | Replace 903-183 one normal. (By Ran) | Replace 903-183 one normal. (By Ran) | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX | 更换903-197后修复(By Ran) | 更换903-197后修复(By Ran) | SWO-177249 | No | Post-replacement 903-197 repair (By Ran) | Post-replacement 903-197 repair (By Ran) | Defect | HW | Oxi module | Spectrometer | 12
ABL800 FLEX | Need to replace printer head | B | SWO-176284 | No | Need to replace printer head |  | Revisit required | HW | Printer Module | Printer head
ABL800 BASIC | Checked the connection between DMS PC and Analyzer Control PCB. After reconnecting the cable analyser found working fine . But repeating the same error intermittently. Has to check with new 636- 506 :Power Communication Cable, Analyzer Control | B | SWO-177299 | No | Checked the connection between DMS PC and Analyzer Control PCB. After reconnecting the cable analyser found working fine . But repeating the same error intermittently. Has to check with new 636- 506 :Power Communication Cable, Analyzer Control |  | Revisit required | HW | Basic Unit | Analyzer control PCB
ABL810 FLEX | 5 Blood samples taken with low pH readings. No patient injury occurred as stated by blood gas coordinator. Possible cause could have been blood clot. Follwoing the 5 consecutive blood samples,on the next calibration pH drift occurred. Data logs show no further calibration errors since event. | Talked with Shelia about incident. Did a electrode channel check (service mode).. | D23EB60CD0487930 | No | 5 Blood samples taken with low pH readings. No patient injury occurred as stated by blood gas coordinator. Possible cause could have been blood clot. Follwoing the 5 consecutive blood samples,on the next calibration pH drift occurred. Data logs show no further calibration errors since event. | Talked with Shelia about incident. Did a electrode channel check (service mode).. | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL810 FLEX | QC results for all levels have been 2 mmHg lower since instal and trended down begining 12-18-16. There have been many rinse errors from pH BG upper LS sensor. There is evidence of a leak at top of module pump tubing. Gas 2 is also the original since install. | Changed pump tubing, reference membrane, inspected all electrode seals and cleaned measuring chamber. Rinse errors subsided and all QC results increased slightly. Replaced gas 2 and pCO2 electrode sensitivity increased 2%. Instrument is operating normally. | D23EA37FD00FAF60 | No | QC results for all levels have been 2 mmHg lower since instal and trended down begining 12-18-16. There have been many rinse errors from pH BG upper LS sensor. There is evidence of a leak at top of module pump tubing. Gas 2 is also the original since install. | Changed pump tubing, reference membrane, inspected all electrode seals and cleaned measuring chamber. Rinse errors subsided and all QC results increased slightly. Replaced gas 2 and pCO2 electrode sensitivity increased 2%. Instrument is operating normally. | Defect | HW | pH/BG Module | Pump tubes
ABL825 FLEX | Customer had already replaced pump tubing on pH BG module. Checked pump rotor assy. Pump calibration showed calibration error. | Cleaned grease from bottom of pump rotor assy. Replaced pump tubing again for fault isolation. | D23E9841D04368D5 | No | Customer had already replaced pump tubing on pH BG module. Checked pump rotor assy. Pump calibration showed calibration error. | Cleaned grease from bottom of pump rotor assy. Replaced pump tubing again for fault isolation. | Defect | HW | pH/BG Module | Pump tubes
AQT90 FLEX | Servicedump gezogen. Daten zur weiteren Auswertung nach Willich geschickt. | Servicedump gezogen. Daten zur weiteren Auswertung nach Willich geschickt. | SWO-177404 | No | Service dump. data sent to Willich for further evaluation. | Service dump. data sent to Willich for further evaluation. | Revisit required | SW | Aspiration Module | Barcode reader, Inlet
AQT90 FLEX |  | B | D23E50AAD0301FCA | No |  |  | Setup | His/lis | His/lis | His/lis
ABL80 FLEX BASIC | 更换试剂包和测试卡后修复 | 更换试剂包和测试卡后修复 | SWO-177878 | No | Repair after replacement of reagent kits and test cards | Repair after replacement of reagent kits and test cards
ABL80 CO-OX | Found P02 disabled and failing system cycles on P02. Customer has replaced two SC and one SP and P02 is still not calibrating. | Ran zeroing cassette and tested again, analyzer passed with no issues after zero cal. | D2489904800E4BB4 | No | Found P02 disabled and failing system cycles on P02. Customer has replaced two SC and one SP and P02 is still not calibrating. | Ran zeroing cassette and tested again, analyzer passed with no issues after zero cal.
ABL80 FLEX CO-OX | neue SC eingesetzt, Initialisierungsabbruch nach 29%, Zero Check durchgeführt und rekalibriert, SC erneut initialisiert, Kal. und QC i.O. | Service Einsatz | SWO-177633 | No | new SC used, initialization abort after 29%, zero check performed and recalibrated, SC reinitialized, Kal. and QC in the first place | Service use
Hb 201+ Analyzer | Checked with new adapter and batteries | B | SWO-177673 | Yes | Checked with new adapter and batteries
Hb 201+ Analyzer | Cleaned with Hemocue cleaner, but E02 did not clear | B | SWO-177688 | Yes | Cleaned with Hemocue cleaner, but E02 did not clear
ABL90 FLEX | Gerät überprüft und prophylaktisch 10-15 durchgespült. Halterung der Einlassdichtung seitlich gebrochen. Einlassdichtung mit Halterung durch Kundeneigene ersetzt. LP und SC erneuert, Gerät kalibriert, SC erneuert. Absprache mit Kundin getroffen, dass eine telefonische Rücksprache nach Initialisierung der SC erfolgen wird. | Gerät überprüft und prophylaktisch 10-15 durchgespült. Halterung der Einlassdichtung seitlich gebrochen. Einlassdichtung mit Halterung durch Kundeneigene ersetzt. LP und SC erneuert, Gerät kalibriert, SC erneuert. Absprache mit Kundin getroffen, dass eine telefonische Rücksprache nach Initialisierung der SC erfolgen wird. | SWO-177774 | No | device and rinsed prophylactic 10-15. Bracket of the inlet seal broken sideways. Inlet seal with bracket replaced by customer's own. LP and SC renewed, device calibrated, SC renewed. Agreement with customer that a telephone consultation will take place after initialization of the SC. | device and rinsed prophylactic 10-15. Bracket of the inlet seal broken sideways. Inlet seal with bracket replaced by customer's own. LP and SC renewed, device calibrated, SC renewed. Agreement with customer that a telephone consultation will take place after initialization of the SC. | Assisted Replacement | SC | SC | Glu
ABL90 FLEX | ich habe die Sotware noch einmal aufgespielt, | B | SWO-177733 | No | I have replayed the Sotware, |  | Defect | SW | SW | Freeze
ABL90 FLEX PLUS |  | B | D24036594647345F | Yes |  |  |  | PER |  |  | HEAD LINE = "Consumable Failure" = PER
ABL90 FLEX PLUS |  | B | D244B8861E5D4519 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Replaced Oximetry pinch valve tube. Flushed and reseated Sensor Cassette. | Replaced Oximetry pinch valve tube. Flushed and reseated Sensor Cassette. | SWO-177699 | No | Replaced Oximetry pinch valve tube. Flushed and reseated Sensor Cassette. | Replaced Oximetry pinch valve tube. Flushed and reseated Sensor Cassette. | Defect | HW | Oxi module | Pinch valve tube/blood
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-177955 | Yes | Repair after replacing test card | Repair after replacing test card
AQT90 FLEX | 更换903-364后修复 | 更换903-364后修复 | SWO-176334 | No | Repair after replacement 903-364 | Repair after replacement 903-364 | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | 重新定标后修复(By Ran) | 重新定标后修复(By Ran) | SWO-177480 | No | Repair after re-calibration (By Ran) | Repair after re-calibration (By Ran) | Adjustment | Consumables | Test Kit | Cartridge | 5784 (Broken)
AQT90 FLEX | 更换886-638一个正常。by jesson | 更换886-638一个正常。by jesson | SWO-178261 | No | Replace 886-638 one normal. by jesson | Replace 886-638 one normal. by jesson | Defect | HW | Cover and Chassis | Inlet Door
ABL800 FLEX | Replace the Crea reagent & calibrate system. | Replace the Crea reagent & calibrate system. | SWO-177841 | No | Replace the Crea reagent & calibrate system. | Replace the Crea reagent & calibrate system. | Assisted Replacement | Electrode | Membrane | Crea Solution
ABL825 FLEX | Membranventile der Messmodule kontrolliert, Verstopfung am Gaseinlaßventil beseitigt, Spülfehler, Kontrolle Spülzyklus, Lufteinlaßventil am ElMetModul ersetzt, Inbetriebnahme, korrekte Kal abgewartet, QK ok | B | SWO-177861 | No | Diaphragm valves of the measuring modules checked, clogging on the gas inlet valve eliminated, flushing error, control flushing cycle, air inlet valve replaced at ElMetmodule, commissioning, correct calving waited, QK ok |  | Defect | HW | El/Met Module | Valve, air
ABL800 BASIC | 修理完了装置返却 | 修理完了装置返却 | SWO-177622 | No | Return repair completion equipment | Return repair completion equipment | Assisted Replacement | Electrode | Membrane | pCO2
ABL815 FLEX | Anbindung gemäß Vorgaben durchgeführt, Kontrolle der Verbindung, Kal und QC | Kunde bittet um einen Service. | SWO-177663 | No | Connection carried out according to specifications, control of the connection, Kal and QC | Customer asks for a service. | Setup | His/lis | His/lis | His/lis
ABL825 FLEX | Unit getting error code 0316: Neon Number Too Low. Tested oximetry module. Failing Neon position too low. Tried adjusted lamp and running tHB calibration. This did not correct issue. | Replaced Spectrometer Unit. Ran complete oximetry tests and tHB calibration with pass results. Restart calibrations. | D240029E46473A1D | No | Unit getting error code 0316: Neon Number Too Low. Tested oximetry module. Failing Neon position too low. Tried adjusted lamp and running tHB calibration. This did not correct issue. | Replaced Spectrometer Unit. Ran complete oximetry tests and tHB calibration with pass results. Restart calibrations. | Defect | HW | Oxi module | Spectrometer | Lamp unit ?
AQT90 FLEX | Neue Charge von NT-ProBNP kalibriert, Systemüberprüfung gestartet, Systemüberprüfungen alle in Ordnung, ServiceDump gezogen, nach ServiceDump Software Version 8.8.16 erneut installiert, Startphase abgewartet, eine weitere Systemüberprüfung gestartet - alles in Ordnung, neue LQC Charge eingelesen, LQC MCC2 Level 2 - Lot 06162 gemessen, TnT und NT-ProBNP war erfolgreich, System- und Funktionskontrolle in Ordnung | Neue Charge von NT-ProBNP kalibriert, Systemüberprüfung gestartet, Systemüberprüfungen alle in Ordnung, ServiceDump gezogen, nach ServiceDump Software Version 8.8.16 erneut installiert, Startphase abgewartet, eine weitere Systemüberprüfung gestartet - alles in Ordnung, neue LQC Charge eingelesen, LQC MCC2 Level 2 - Lot 06162 gemessen, TnT und NT-ProBNP war erfolgreich, System- und Funktionskontrolle in Ordnung | SWO-177727 | No | New batch of NT-ProBNP calibrated, system check started, system checks all in order, ServiceDump pulled, reinstalled after ServiceDump software version 8.8.16, start phase waited, another system check started - all right, new LQC batch read in, LQC MCC2 Level 2 - Lot 06162 measured, TnT and NT-ProBNP was successful, system and function control in order | New batch of NT-ProBNP calibrated, system check started, system checks all in order, ServiceDump pulled, reinstalled after ServiceDump software version 8.8.16, start phase waited, another system check started - all right, new LQC batch read in, LQC MCC2 Level 2 - Lot 06162 measured, TnT and NT-ProBNP was successful, system and function control in order | Assisted Replacement | Consumables | Test Kit | NT-Pro BNP
AQT90 FLEX |  | B | D23F9CA1460A3E48 | No |  |  | Defect | HW | Incubator module | Incubator Lid
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-177815 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Reference
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-178268 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Lac | Sensitivity
AQT90 FLEX | 更换925-016后修复(By Ran) | 更换925-016后修复(By Ran) | SWO-177817 | No | Repair after replacement 925-016 (By Ran) | Repair after replacement 925-016 (By Ran) | Defect | HW | Cover and Chassis | Inlet Door
ABL825 FLEX | Autocheck aborting QC without user message. Activity log shows error 0742, Autocheck Carrosel failure - couldn't move. | When removing Autocheck module for inspection, cooling fan fell out on ground. Replaced cooling fan. Also replaced ribbon cable. Check Autocheck in service - all okay. Also removed old decommissioned analyzer for local disposal. Been sitting for removal since 1-26. | D2420EAA323F35F3 | No | Autocheck aborting QC without user message. Activity log shows error 0742, Autocheck Carrosel failure - couldn't move. | When removing Autocheck module for inspection, cooling fan fell out on ground. Replaced cooling fan. Also replaced ribbon cable. Check Autocheck in service - all okay. Also removed old decommissioned analyzer for local disposal. Been sitting for removal since 1-26. | Defect | HW | Auto Check Module | FPC Cable
ABL80 FLEX BASIC | 更换备件903-334后修复 by andrew | 更换备件903-334后修复 by andrew | SWO-177834 | No | Repair by Andrew after replacing spare parts 903-334 | Repair by Andrew after replacing spare parts 903-334
ABL80 FLEX BASIC | Need to replace power supply board | B | SWO-179690 | No | Need to replace power supply board
ABL80 FLEX Basic version |  | B | D2440037A843B5F4 | No
ABL80 CO-OX | Paula said the solution pack door would come back a little when she latches it closed. She believes it is not pushing the SP in tight enough. Replaced door hinge and it is tighter. Will also claim the COOX zero cassette to get off my inventory. | Replaced solution pack molded door hinge assembly. | D2435D0FA85FFB31 | No | Paula said the solution pack door would come back a little when she latches it closed. She believes it is not pushing the SP in tight enough. Replaced door hinge and it is tighter. Will also claim the COOX zero cassette to get off my inventory. | Replaced solution pack molded door hinge assembly.
ABL90 FLEX | Touch Kal. im Controll Panel durchgeführt. Funktionen wieder normal. | Touch Kal. im Controll Panel durchgeführt. Funktionen wieder normal. | SWO-178837 | No | Touch Kal. performed in the Controll Panel. functions back to normal. | Touch Kal. performed in the Controll Panel. functions back to normal. | Adjustment | HW | DMS_PC | Touch screen
TCM400 |  | B | SWO-178015 | No
ABL90 FLEX |  | 更换903-199后仍出现间歇性故障，需要重新借903-664及910-285.此单先关单 bu andrew | SWO-177869 | No |  | Intermittent failure after replacement of 903-199 and needs to be re-borrowed 903-664 and 910-285. | Defect | HW | Basic Unit | Cassette retainer | 134
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-177891 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Reference | 1361
AQT90 FLEX | 重新定标正常。 by andrew | 重新定标正常。 by andrew | SWO-177850 | No | Recalibration is normal. by andrew | Recalibration is normal. by andrew | Assisted Replacement | Consumables | Test Kit | Cartridge, Myo | 5297
AQT90 FLEX |  | B | SWO-177910 | No |  |  | Defect | HW | Aspiration Module | Wet Section, Manifold
AQT90 FLEX |  | B | SWO-178037 | No |  |  | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL825 FLEX |  | B | D24D8AF458503B45 | No |  |  | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX |  | B | SWO-177976 | No |  |  | Defect | HW | Inlet module | Nipple kit
ABL800 BASIC | Cleaned The Lactate Chamber & Lactate Electrode Tip, Run The 2 Point Calibration. After Passing 2 Point Calibration Again Run QC Level 1 ,In Result Lactate Values Are In Range . Machine Working Fine. | Cleaned The Lactate Chamber & Lactate Electrode Tip, Run The 2 Point Calibration. After Passing 2 Point Calibration Again Run QC Level 1 ,In Result Lactate Values Are In Range . Machine Working Fine. | SWO-179545 | No | Cleaned The Lactate Chamber & Lactate Electrode Tip, Run The 2 Point Calibration. After Passing 2 Point Calibration Again Run QC Level 1 ,In Result Lactate Values Are In Range . Machine Working Fine. | Cleaned The Lactate Chamber & Lactate Electrode Tip, Run The 2 Point Calibration. After Passing 2 Point Calibration Again Run QC Level 1 ,In Result Lactate Values Are In Range . Machine Working Fine. | Cleaned | Fluidic pathway | El/Met Module | Measuring chamber | Lac
ABL827 FLEX | Messkammer im pH BG Modul gereinigt, Tausch des pH Dichtrings, mehrere Kal. durchgeführt, Proteinentfernung durchgeführt, System- und Funktionskontrolle in Ordnung | Messkammer im pH BG Modul gereinigt, Tausch des pH Dichtrings, mehrere Kal. durchgeführt, Proteinentfernung durchgeführt, System- und Funktionskontrolle in Ordnung | SWO-178079 | No | Measuring chamber cleaned in pH BG module, replacement of pH sealing ring, several cal. performed, protein removal performed, system and function control in order | Measuring chamber cleaned in pH BG module, replacement of pH sealing ring, several cal. performed, protein removal performed, system and function control in order | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL815 FLEX | Lac Elektrode gründlich mit Proteinentferner gereinigt, Membrane erneut getauscht - sie schien undicht zu sein, Messkammer gereinigt, Glasfläche der Elektrode gründlich gereinigt, Einlassnadel und Einlassdichtung aus Kundenbestand getauscht - Einlassnadel war stark verbogen (Knick), Lufteinlass gespült, mehrere Kal. durchgeführt, System- und Funktionskontrolle in Ordnung | Lac Elektrode gründlich mit Proteinentferner gereinigt, Membrane erneut getauscht - sie schien undicht zu sein, Messkammer gereinigt, Glasfläche der Elektrode gründlich gereinigt, Einlassnadel und Einlassdichtung aus Kundenbestand getauscht - Einlassnadel war stark verbogen (Knick), Lufteinlass gespült, mehrere Kal. durchgeführt, System- und Funktionskontrolle in Ordnung | SWO-178154 | No | Lac electrode thoroughly cleaned with protein remover, membrane replaced - it seemed to be leaking, measuring chamber cleaned, glass surface of the electrode thoroughly cleaned, inlet needle and inlet seal replaced from customer stock - inlet needle was strong bent (kink), air inlet rinsed, several cal. performed, system and function control in order | Lac electrode thoroughly cleaned with protein remover, membrane replaced - it seemed to be leaking, measuring chamber cleaned, glass surface of the electrode thoroughly cleaned, inlet needle and inlet seal replaced from customer stock - inlet needle was strong bent (kink), air inlet rinsed, several cal. performed, system and function control in order | Assisted Replacement | Electrode | Membrane | Lac
ABL815 FLEX | Messkammer gereinigt. Ref Elektrode membraniert, Ca+ Elektrode Membraniert. Gerät kalibriert und QC durchgeführt. | B | SWO-178795 | No | measuring chamber. Ref electrode membranes, Ca+ electrode membraneed. calibrated and QC performed. |  | Cleaned | Fluidic pathway | El/Met Module | Measuring chamber
ABL825 FLEX | Customer remembraned electrodes and still the same issue. | Cleaned measuring chambers and electrode contacts. Performed gas1 and gas2 tests in service diagnostics.Replaced inlet probe. | D243BF53A821FF63 | No | Customer remembraned electrodes and still the same issue. | Cleaned measuring chambers and electrode contacts. Performed gas1 and gas2 tests in service diagnostics. Replaced inlet probe. | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL825 FLEX | Replaced inlet nipples and inlet gasket. | All tests are successful. | D243C300A84DFF68 | No | Replaced inlet nipples and inlet gasket. | All tests are successful. | Defect | HW | Inlet module | Inlet Gasket
AQT90 FLEX | after Checking found Analyzer , Showing Error 3050 Corroptxml, Restart The Analyzer, Done Auto Refresh for Incubation Temperature, Analyzer took 20 mins to stabilize incubation temperature, problem resolve, Analyzer working ok. | after Checking found Analyzer , Showing Error 3050 Corroptxml, Restart The Analyzer, Done Auto Refresh for Incubation Temperature, Analyzer took 20 mins to stabilize incubation temperature, problem resolve, Analyzer working ok. | SWO-179534 | No | after Checking found Analyzer , Showing Error 3050 Corroptxml, Restart The Analyzer, Done Auto Refresh for Incubation Temperature, Analyzer took 20 mins to stabilize incubation temperature, problem resolve, Analyzer working ok. | after Checking found Analyzer , Showing Error 3050 Corroptxml, Restart The Analyzer, Done Auto Refresh for Incubation Temperature, Analyzer took 20 mins to stabilize incubation temperature, problem resolve, Analyzer working ok. | Defect | SW | Analyzer Control Module | Off/On | 35
AQT90 FLEX |  | B | D24427E9A84A7BF8 | No |  |  | User error | User error | User error | User error
ABL80 FLEX | After checking found that Touch Panal Get Defective, Replaced Touch Panel with New One, And Reinstalled SC and SP, Problem Resolve. | After checking found that Touch Panal Get Defective, Replaced Touch Panel with New One, And Reinstalled SC and SP, Problem Resolve. | SWO-179536 | No | After checking found that Touch Panal Get Defective, Replaced Touch Panel with New One, And Reinstalled SC and SP, Problem Resolve. | After checking found that Touch Panal Get Defective, Replaced Touch Panel with New One, And Reinstalled SC and SP, Problem Resolve.
ABL80 FLEX Basic version |  | B | D244ADB01E0D1D31 | No
ABL90 FLEX PLUS | Time server IP address added to Internet time server setting and test connection ok. | Time server IP address added to Internet time server setting and test connection ok. | SWO-178501 | No | Time server IP address added to Internet time server setting and test connection ok. | Time server IP address added to Internet time server setting and test connection ok. | Setup | Time setting | Time setting | Time setting
ABL90 FLEX PLUS | Replaced inlet probe and inlet gasket w. Holder. Fluid transport tests ok. Start up tests completed. PCO2 sensitivity out of range.(rising with every calibration-customer to monitor and replace if necessary) | Replaced inlet probe and inlet gasket w. Holder. Fluid transport tests ok. Start up tests completed. PCO2 sensitivity out of range.(rising with every calibration-customer to monitor and replace if necessary) | SWO-180249 | No | Replaced inlet probe and inlet gasket w. Holder. Fluid transport tests ok. Start up tests completed. PCO2 sensitivity out of range. (rising with every calibration-customer to monitor and replace if necessary) | Replaced inlet probe and inlet gasket w. Holder. Fluid transport tests ok. Start up tests completed. PCO2 sensitivity out of range. (rising with every calibration-customer to monitor and replace if necessary) | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX |  | B | SWO-178220 | No |  |  | Defect | HW | DMS_PC | CPU
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-179175 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Glu/Lac | 21
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-178920 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Reference | 1361
AQT90 FLEX | 更换903-364一个正常。 by andrew | 更换903-364一个正常。 by andrew | SWO-177867 | Yes | Replace 903-364 one normal. by andrew | Replace 903-364 one normal. by andrew | Defect | HW | Process Module | Optical Unit | 124
ABL825 FLEX | CONFIRMED ISSUE WHEN ONSITE. EI MET MODULE WAS ALREADY REPLACED BY HJP ON FEB 21 SO ISSUE COULD BE CAUSED BY MINOR RINSE SOLUTION MIXED W CAL SOLUTION IN RIGHT FLUIDIC MODULE. | REPLACED V2, V3 & V8 VALVES FOR TIGHT FLUIDIC SEAL. ALSO REPLACED ALL MEMBRANES ON V1, V4, V5 AND V6 WHILE MODULE WAS OUT. | D244C6B71E580287 | No | CONFIRMED ISSUE WHEN ONSITE. EI MET MODULE WAS ALREADY REPLACED BY HJP ON FEB 21 SO ISSUE COULD BE CAUSED BY MINOR RINSE SOLUTION MIXED W CAL SOLUTION IN RIGHT FLUIDIC MODULE. | REPLACED V2, V3 & V8 VALVES FOR TIGHT FLUIDIC SEAL. ALSO REPLACED ALL MEMBRANES ON V1, V4, V5 AND V6 WHILE MODULE WAS OUT. | Defect | HW | El/Met Module | Valve
ABL825 FLEX | Checked calibration log for errors. Unit failing with a :0251 Halogen Lamp intensity too high. | Adjusted halogen lamp down to 3100. Ran completed oximetry tests in diagnostics with pass results. Restart calibrations. Now getting Neon Number Too Low, error code :0316. Replaced spectrometer unit. Restart testing and completed tHB calibration. | D244C5D01E3E8D37 | No | Checked calibration log for errors. Unit failing with a :0251 Halogen Lamp intensity too high. | Adjusted halogen lamp down to 3100. Ran completed oximetry tests in diagnostics with pass results. Restart calibrations. Now getting Neon Number Too Low, error code :0316. Replaced spectrometer unit. Restart testing and completed tHB calibration. | Adjustment | HW | Oxi module | Lamp unit | Halogen
AQT90 FLEX |  | B | D244BF211E4DCFFA | No |  |  | Defect | HW | Aspiration Module | P4 | 1218
AQT90 FLEX | Vanmorgen binnengekomen nieuwe CRP kit (12521) met succes gecalibreerd. | Vanmorgen binnengekomen nieuwe CRP kit (12521) met succes gecalibreerd. | SWO-178436 | No | This morning, new CRP Kit (12521) was successfully calibrated. | This morning, new CRP Kit (12521) was successfully calibrated. | Defect | Consumables | Test Kit | Cartridge, TnT | Broken
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-178933 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-179180 | No | Replace the test card | Replace the test card
ABL90 FLEX | Akku defekt, Ersatzteil bestellt | Austausch der Batterie Option, alter Akku 911-311P0072N001 | SWO-178664 | No | Battery defective, spare part ordered | Battery replacement option, old battery 911-311P0072N001 | Revisit required | HW | Battery | Battery
ABL90 FLEX | Fluid Transport-Serviceprogramme durchgeführt. Durch Pumpenkalibration Fehler beseitigt. Nach Kaltstart Fehler 1396, 1340 sowie Kalibrierfehler Glu & pCO2. Sensorkassettentausch empfohlen und nach erneuter gescheiterter Kalibration durchgeführt. Einlassdichtung und Anschlussdichtung erneuert und mit dem Kunden Vorgehensweise besprochen. Einlaufphase der Sensorkassette beobachtet. Funktion fehlerfrei. | Fluid Transport-Serviceprogramme durchgeführt. Durch Pumpenkalibration Fehler beseitigt. Nach Kaltstart Fehler 1396, 1340 sowie Kalibrierfehler Glu & pCO2. Sensorkassettentausch empfohlen und nach erneuter gescheiterter Kalibration durchgeführt. Einlassdichtung und Anschlussdichtung erneuert und mit dem Kunden Vorgehensweise besprochen. Einlaufphase der Sensorkassette beobachtet. Funktion fehlerfrei. | SWO-178951 | No | Fluid Transport service programs. Eliminated by pump calibration errors. After cold start error 1396, 1340 as well as calibration error Glu & pCO2. sensor cartridge replacement and performed after another failed calibration. Inlet seal and connection seal renewed and discussed with the customer's procedure. inlet phase of the sensor cartridge. Function error-free. | Fluid Transport service programs. Eliminated by pump calibration errors. After cold start error 1396, 1340 as well as calibration error Glu & pCO2. sensor cartridge replacement and performed after another failed calibration. Inlet seal and connection seal renewed and discussed with the customer's procedure. inlet phase of the sensor cartridge. Function error-free. | Adjustment | HW | Service | Pump Calibration
ABL90 FLEX | Fehlersuche Netzwerkeinstellungen vorgenommen. 3 Geräte im Haus stillgelegt. | Fehlersuche Netzwerkeinstellungen vorgenommen. 3 Geräte im Haus stillgelegt. | SWO-178969 | No | Troubleshooting network settings. 3 appliances in the house shut down. | Troubleshooting network settings. 3 appliances in the house shut down. | Setup | His/lis | His/lis | Radiance
ABL90 FLEX PLUS |  | B | D248DB5B80580DFD | No |  |  | Assisted Replacement | SC | SC | Reference
ABL825 FLEX | 更换GLU电极膜正常。 | 更换GLU电极膜正常。 | SWO-179173 | No | Replace the GLU electrode film normally. | Replace the GLU electrode film normally. | Assisted Replacement | Electrode | Membrane | Glu
ABL825 FLEX | Bei eintreffen war die Nadel verbogen. Deswegen Nadel und Einlassdichtung getauscht. ausKundenbestand entnommen. Ventile und Ventilsitze gereinigt aus pH bG und Ele Met Modul. Pumpenschlscueh getauscht. Mehrere Kalibrierungen durchgeführt. tHb Kal durchgeführt. --> i.O. Fcuv.: 0,9933 Kal Verifizierung durchgeführt--> i.O. Kappiläre Testmessung mit QC Lösung durchgeführt. --> Lösung wird ohne Probleme angesaugt. QCs für alle Level durchgeführt. ABL wieder messbereit. | Bei eintreffen war die Nadel verbogen. Deswegen Nadel und Einlassdichtung getauscht. ausKundenbestand entnommen. Ventile und Ventilsitze gereinigt aus pH bG und Ele Met Modul. Pumpenschlscueh getauscht. Mehrere Kalibrierungen durchgeführt. tHb Kal durchgeführt. --> i.O. Fcuv.: 0,9933 Kal Verifizierung durchgeführt--> i.O. Kappiläre Testmessung mit QC Lösung durchgeführt. --> Lösung wird ohne Probleme angesaugt. QCs für alle Level durchgeführt. ABL wieder messbereit. | SWO-178870 | No | When it arrived, the needle was bent. Therefore, needle and inlet seal are exchanged. taken from customer stock. Valves and valve seats cleaned from pH bG and Ele Met module. Pump scueh. Multiple calibrations performed. tHb Kal. --> i.O. Fcuv.: 0.9933 Kal Verification performed--> i.O. Kappilular test measurement performed with QC solution. --> Solution is sucked in without any problems. QCs for all levels. ABL ready to be measured again. | When it arrived, the needle was bent. Therefore, needle and inlet seal are exchanged. taken from customer stock. Valves and valve seats cleaned from pH bG and Ele Met module. Pump scueh. Multiple calibrations performed. tHb Kal. --> i.O. Fcuv.: 0.9933 Kal Verification performed--> i.O. Kappilular test measurement performed with QC solution. --> Solution is sucked in without any problems. QCs for all levels. ABL ready to be measured again. | Defect | HW | Inlet module | Probe bend
ABL820 FLEX | CONFIRMED ISSUE WHEN ONSITE. | REBOOT ANALYZER AND ATTACHED MOUSE TO ANALYZER USB PORT UNDER SCREEN. HYPERLINK IS NOW LINKING INDIVIDUALLY TOPICS SUCCESSFULLY TO RESPECTIVE SUBJECTS. | D2467113944B029E | No | CONFIRMED ISSUE WHEN ONSITE. | REBOOT ANALYZER AND ATTACHED MOUSE TO ANALYZER USB PORT UNDER SCREEN. HYPERLINK IS NOW LINKING INDIVIDUALLY TOPICS SUCCESSFULLY TO RESPECTIVE SUBJECTS. | Setup | His/lis | His/lis | His/lis
ABL835 FLEX | Pleura fluid not installed. | Installed pleura fluid option and verified that the profile was setup and the syringe mode was added. | D246042294343C69 | Yes | Pleura fluid not installed. | Installed pleura fluid option and verified that the profile was setup and the syringe mode was added. | Re classified RMED | User request
ABL835 FLEX | Pleura key not installed on analyzer. | Installed pleura key on analyzer and verified that profiled was setup and button added to syringe mode. | D2460956942F3C6D | Yes | Pleura key not installed on analyzer. | Installed pleura key on analyzer and verified that profiled was setup and button added to syringe mode. | Re classified RMED | User request
AQT90 FLEX |  | B | D2461B2694494788 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX |  | B | D2460D5E941F2206 | No |  |  | No fault found | No fault found | No fault found | No fault found | 5669
ABL80 FLEX BASIC | 拆备用机上的902-972一个更换后正常。 by andrew | 拆备用机上的902-972一个更换后正常。 by andrew | SWO-178582 | No | Dismantling the 902-972 on the spare machine is normal after replacement. by andrew | Dismantling the 902-972 on the spare machine is normal after replacement. by andrew
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-178932 | Yes | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Einlass-Sonde wurde begradigt und das gerät im option Perform pressurtest mehrmals kalibriert | B | SWO-179145 | No | Inlet probe has been straightened and this is calibrated several times in the option Perform pressurtest |  | Replaced | HW | Inlet module | Probe bend
ABL90 FLEX |  | Der Kunde bittet um einen Service. | SWO-179423 | No |  | The customer asks for a service. | User error | Pre analytical | Pre analytical | Sample mix
ABL90 | Sodium correlation high value value compared to ABBOT architect | Ran sensor simulator test alls channel are OK internal Qcs and external Qcs are in the midlle of the range | D2473B050A59E780 | No | Sodium correlation high value value compared to ABBOT architect | Ran sensor simulator test alls channel are OK internal Qcs and external Qcs are in the midlle of the range | No fault found | Perfomance | Perfomance | Na
ABL90 FLEX | IT Connection Ethernet überprüft. Keine feste IP und Netzwerkdaten gespeichert, Stand DHCP. Eingabe der Netzwerkdaten. Speicherung auf Windowsebene. Verbindung hergestellt. Neustart. Keine Verbindung nach Neustart. Fehlermeldung Ethernetadapter. Ethernet Adapter speichert Daten nicht ab. Überprüfung und ändern der Etherneteinstellungen. Erneute Eingabe der Berbindungsdaten. Speicherung auf Windowsebene. Nicht erfolgreich nach Neustart. Deinstallation der Netzwerkkarte. Neuinstallation nach Neustart. Eingabe der Verbindungsdaten und Speicherung. Netzwerkverbindung nach Neustart nicht hergestellt. Nochmals manuell hergestellt. Verbindung wird nach einem Neustart wieder verloren gehen! CPU inkl. Netzwerkkarte wird bestellt und nach Erhalt eigebaut, keine Speicherung der Anbindungsdaten möglich. | IT Connection Ethernet überprüft. Keine feste IP und Netzwerkdaten gespeichert, Stand DHCP. Eingabe der Netzwerkdaten. Speicherung auf Windowsebene. Verbindung hergestellt. Neustart. Keine Verbindung nach Neustart. Fehlermeldung Ethernetadapter. Ethernet Adapter speichert Daten nicht ab. Überprüfung und ändern der Etherneteinstellungen. Erneute Eingabe der Berbindungsdaten. Speicherung auf Windowsebene. Nicht erfolgreich nach Neustart. Deinstallation der Netzwerkkarte. Neuinstallation nach Neustart. Eingabe der Verbindungsdaten und Speicherung. Netzwerkverbindung nach Neustart nicht hergestellt. Nochmals manuell hergestellt. Verbindung wird nach einem Neustart wieder verloren gehen! CPU inkl. Netzwerkkarte wird bestellt und nach Erhalt eigebaut, keine Speicherung der Anbindungsdaten möglich. | SWO-179058 | No | IT Connection Ethernet. No fixed IP and network data stored, status DHCP. Entering the network data. Windows-level storage. Connection established. Restart. No connection after reboot. Error message Ethernet adapter. Ethernet Adapter does not store data. Check and change Ethernet settings. Re-enter the binding data. Windows-level storage. Not successful after reboot. Uninstall the network card. Reinstall after reboot. Enter the connection data and save it. Network connection not established after restart. Manufactured manually again. Connection will be lost again after a restart! CPU incl. network card is ordered and built after receipt, no storage of the connection data possible. | IT Connection Ethernet. No fixed IP and network data stored, status DHCP. Entering the network data. Windows-level storage. Connection established. Restart. No connection after reboot. Error message Ethernet adapter. Ethernet Adapter does not store data. Check and change Ethernet settings. Re-enter the binding data. Windows-level storage. Not successful after reboot. Uninstall the network card. Reinstall after reboot. Enter the connection data and save it. Network connection not established after restart. Manufactured manually again. Connection will be lost again after a restart! CPU incl. network card is ordered and built after receipt, no storage of the connection data possible. | Revisit required | HW | DMS_PC | CPU | His/Lis
ABL90 FLEX PLUS |  | B | D247954C0A3B5F7C | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX PLUS | First of all the solution pack was removed and replaced but this didn't fix the problem. The inlet module was removed and a blood clot found at the end of the probe, this was flushed out and the inlet gasket flushed but this still didn't cure the problem. A new inlet gasket was fitted and following this the analyser did pass its pressure test. | B | SWO-179301 | No | First of all the solution pack was removed and replaced but this didn't fix the problem. The inlet module was removed and a blood clot found at the end of the probe, this was flushed out and the inlet gasket flushed but this still didn't cure the problem. A new inlet gasket was fitted and following this the analyser did pass its pressure test. |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D2474FC10A185F77 | No |  |  | Defect | HW | Sample mixer | Sample mixer | LEH Mail
AQT90 FLEX | 更换903-364一个正常。by andrew | 更换903-364一个正常。by andrew | SWO-178293 | No | Replace 903-364 one normal. by andrew | Replace 903-364 one normal. by andrew | Defect | HW | Process Module | Optical Unit | 135
ABL800 BASIC | After Sudden Power Failure , Machine Started Again & Gives The Leak Detected Error. Shut Down The Machine By Standard Method, Checked All The Connections, Again Started The Machine, Machine Initialised & Passed 2 Point Calibration. Run The Sample , Machine Working Ok | After Sudden Power Failure , Machine Started Again & Gives The Leak Detected Error. Shut Down The Machine By Standard Method, Checked All The Connections, Again Started The Machine, Machine Initialised & Passed 2 Point Calibration. Run The Sample , Machine Working Ok | SWO-179547 | No | After Sudden Power Failure , Machine Started Again & Gives The Leak Detected Error. Shut Down The Machine By Standard Method, Checked All The Connections, Again Started The Machine, Machine Initialised & Passed 2 Point Calibration. Run The Sample , Machine Working Ok | After Sudden Power Failure , Machine Started Again & Gives The Leak Detected Error. Shut Down The Machine By Standard Method, Checked All The Connections, Again Started The Machine, Machine Initialised & Passed 2 Point Calibration. Run The Sample , Machine Working Ok | User error | Mains Power Failure | Mains Power Failure | Mains Power Failure
ABL825 FLEX |  | B | D247503B0A1FC50E | Yes |  |  | Re classified RMED | Phone support
ABL825 FLEX | Messkammer gereinigt und Ref Kontakte nachgebogen - keine Verbesserung, Ref Elektrode erneuert - keine Verbesserung, Ref Elektrode zurück getauscht, Ref Connector erneuert - Empfindlichkeit und Status aller Elektrolyte und pH OK, mehrfach Kalibriert und alle 3 QC Lvl durchgeführt. | Der Kunde bittet um einen Service und Rückruf. | SWO-179121 | No | Measuring chamber cleaned and ref contacts bent - no improvement, ref electrode renewed - no improvement, ref electrode replaced, ref connector renewed - sensitivity and status of all electrolytes and pH OK, multiple calibrated and all 3 QC Lvl Performed. | The customer asks for a service and callback. | Assisted Replacement | Electrode | Electrode | Reference
ABL800 BASIC | Lactate Membrane Need To Be Replace, Replaced The Lactate Membrane, Machine Passed 2 Point Calibration. Run The Sample, Machine Working Fine | Lactate Membrane Need To Be Replace, Replaced The Lactate Membrane, Machine Passed 2 Point Calibration. Run The Sample, Machine Working Fine | SWO-179551 | No | Lactate Membrane Need To Be Replace, Replaced The Lactate Membrane, Machine Passed 2 Point Calibration. Run The Sample, Machine Working Fine | Lactate Membrane Need To Be Replace, Replaced The Lactate Membrane, Machine Passed 2 Point Calibration. Run The Sample, Machine Working Fine | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Verified customer's complaint. | Performed inlet calibration. Replaced inlet gasket, with customer stock. Verified operation of ABL825. | D24761760A0B6D7E | No | Verified customer's complaint. | Performed inlet calibration. Replaced inlet gasket, with customer stock. Verified operation of ABL825. | Adjustment | HW | Inlet module | Probe cal.
ABL837 FLEX |  | B | D24EE66FCE632D60 | No |  |  | Revisit required | HW | Auto Check Module | Auto Check Module
ABL810 FLEX | UPON ARRIVED DISCOVED THE REPORT FOR BOTH ANALYZERS HAD BEEN RECEIVED AFTER INSTALLATION. MISIDENTIFIED IN EMAIL. CUSTOMER WAS ABLE TO PRINT AND REVIEW THE REPORTS. ALSO STATED THEY NEEDED HELP WITH CONNECTION FOR HIS LIS SETTINGS | WORKED WITH CUSTOMER IT TO RESOLVE CONNECTION SETTINGS FOR ANALYZER IP AND HOST IP. | D247D0EC0A1DBA7C | No | UPON ARRIVED DISCOVED THE REPORT FOR BOTH ANALYZERS HAD BEEN RECEIVED AFTER INSTALLATION. MISIDENTIFIED IN EMAIL. CUSTOMER WAS ABLE TO PRINT AND REVIEW THE REPORTS. ALSO STATED THEY NEEDED HELP WITH CONNECTION FOR HIS LIS SETTINGS | WORKED WITH CUSTOMER IT TO RESOLVE CONNECTION SETTINGS FOR ANALYZER IP AND HOST IP. | Setup | His/lis | His/lis | His/lis
ABL80 FLEX | QC Results out of Range, Cleaned SC80, Cleaned Connector perform Zeroing, Reinstalled SC, Run QC Problem Remains,Need to replace Sensor Board, TO Placed for Spare. | B | SWO-179539 | No | QC Results out of Range, Cleaned SC80, Cleaned Connector perform Zeroing, Reinstalled SC, Run QC Problem Remains,Need to replace Sensor Board, TO Placed for Spare.
ABL80 FLEX BASIC | 更换903-239修复 by andrew | 更换903-239修复 by andrew | SWO-178972 | No | Replace 903-239 Repair by Andrew | Replace 903-239 Repair by Andrew
ABL90 FLEX | Austausch der Batterie Option, alter Akku 911-311P0072N001 | Austausch der Batterie Option, alter Akku 911-311P0072N001 | SWO-179382 | No | Battery replacement option, old battery 911-311P0072N001 | Battery replacement option, old battery 911-311P0072N001 | Defect | HW | Battery | Battery
ABL90 | Four sensor cassettes had to be installed to get one that would work. | Replaced the sensor cassette interface board. | D24888FA803575D2 | No | Four sensor cassettes had to be installed to get one that would work. | Replaced the sensor cassette interface board. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX PLUS | Einlass Positionierung im Servicemenü überprüft, OK. INLET POSITION DETECTOR PCB 903-337 ERSETZT (aus Demogerät 092R0104N005), Neustart, Test der Funktionen im Servicemenü, OK, Endkontrolle: QC Messung OK | Einlass Positionierung im Servicemenü überprüft, OK. INLET POSITION DETECTOR PCB 903-337 ERSETZT (aus Demogerät 092R0104N005), Neustart, Test der Funktionen im Servicemenü, OK, Endkontrolle: QC Messung OK | SWO-179287 | No | Inlet positioning checked in the service menu, OK. INLET POSITION DETECTOR PCB 903-337 REPLACEMENT (from demo device 092R0104N005), restart, test of functions in the service menu, OK, final control: QC measurement OK | Inlet positioning checked in the service menu, OK. INLET POSITION DETECTOR PCB 903-337 REPLACEMENT (from demo device 092R0104N005), restart, test of functions in the service menu, OK, final control: QC measurement OK | Defect | HW | Auto Inlet module | PCB, Inlet Position
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-179227 | No | The replacement kit is OK. | The replacement kit is OK. | Assisted Replacement | SP | SP | FT | 163
ABL90 FLEX | 更换903-139一个正常。 by andrew | 更换903-139一个正常。 by andrew | SWO-178994 | No | Replace the 903-139 one normal. by andrew | Replace the 903-139 one normal. by andrew | Defect | HW | Inlet module | Inlet module
ABL825 FLEX |  | B | D24885DF803C5111 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX |  | B | D2547362A634E69C | No |  |  | Defect | HW | El/Met Module | Pump tubes
ABL90 FLEX | 更换903-185，580-417修复 by andrew | 更换903-185，580-417修复 by andrew | SWO-179242 | No | Replace 903-185,580-417 Fix by Andrew | Replace 903-185,580-417 Fix by Andrew | Defect | HW | Wet section | PCB, Flow Selector | 1177
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-179549 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Reference | 1361
AQT90 FLEX | 更换903-175和903-198后，仍旧未修复，追加借件910-330一个 by andrew | 更换903-175和903-198后，仍旧未修复，追加借件910-330一个 by andrew | SWO-179570 | No | After replacing 903-175 and 903-198, still unrepaired, added borrowing 910-330 a by andrew | After replacing 903-175 and 903-198, still unrepaired, added borrowing 910-330 a by andrew | Revisit required | HW | Analyzer Control Module | PCB, Main Board
AQT90 FLEX | Cup Stuck inside incubator wheel. Cup Removed from incubator wheel. Analyzer serviced and calibrated properly. Handed over in a fine working condition. one test run to check. working fine. | Cup Stuck inside incubator wheel. Cup Removed from incubator wheel. Analyzer serviced and calibrated properly. Handed over in a fine working condition. one test run to check. working fine. | SWO-179569 | No | Cup Stuck inside incubator wheel. Cup Removed from incubator wheel. Analyzer serviced and calibrated properly. Handed over in a fine working condition. one test run to check. working fine. | Cup Stuck inside incubator wheel. Cup Removed from incubator wheel. Analyzer serviced and calibrated properly. Handed over in a fine working condition. one test run to check. working fine. | Defect | Consumables | SP | Cup Stack
ABL80 FLEX BASIC |  | B | SWO-179603 | No
AQT90 FLEX |  | B | SWO-179614 | No |  |  | Defect | HW | Aspiration Module | Barcode reader, Inlet
AQT90 FLEX |  | 更换903-175和903-198后，仍旧未修复，追加借件910-330一个 by andrew | SWO-179601 | No |  | After replacing 903-175 and 903-198, still unrepaired, added borrowing 910-330 a by andrew | Defect | HW | Analyzer Control Module | PCB, Main Board
ABL80 FLEX BASIC |  | B | SWO-179743 | No
ABL80 CO-OX | The customer said the SP was leaking | I checked out the latch and the SP. It looks like there was a leak but it is not at this time. I ran a system cycle to see if it would leak and it was dry. the lot number of SP is 235553. | D24C8325E2098B9E | No | The customer said the SP was leaking | I checked out the latch and the SP. It looks like there was a leak but it is not at this time. I ran a system cycle to see if it would leak and it was dry. the lot number of SP is 235553.
ABL80 FLEX | O2 Values out of range, QC Failure, Cleaned Connector and Interior Dust and Reinstalled. Problem Remains, Replaced Sensor Cassette with New one. | O2 Values out of range, QC Failure, Cleaned Connector and Interior Dust and Reinstalled. Problem Remains, Replaced Sensor Cassette with New one. | SWO-181286 | No | O2 Values out of range, QC Failure, Cleaned Connector and Interior Dust and Reinstalled. Problem Remains, Replaced Sensor Cassette with New one. | O2 Values out of range, QC Failure, Cleaned Connector and Interior Dust and Reinstalled. Problem Remains, Replaced Sensor Cassette with New one.
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-180000 | No | Repair after replacing test card | Repair after replacing test card
ABL80 CO-OX | the OXI wauld not cal and was turned off | I pulled out the OXI and checked all the tubings and reseated them. It is working now. | D24C58DEE201AA21 | No | the OXI wauld not cal and was turned off | I pulled out the OXI and checked all the tubings and reseated them. It is working now.
ABL80 CO-OX | CONFIRMED ISSUE WHEN ONSITE. CUSTOMER HAS ALREADY TRIED DIFFERENT SENSOR CASSETTE BUT LEAK CONTINUED. | TUBING CONNECTING WASTE PORT AND ACRYLIC VALVE KINKED, CAUSING LEAK. BROUGHT FORWARD PM DUE IN JULY TO REPLACE ALL TUBINGS AND PUMPS IN PM KIT. | D24C63D6E2141C50 | No | CONFIRMED ISSUE WHEN ONSITE. CUSTOMER HAS ALREADY TRIED DIFFERENT SENSOR CASSETTE BUT LEAK CONTINUED. | TUBING CONNECTING WASTE PORT AND ACRYLIC VALVE KINKED, CAUSING LEAK. BROUGHT FORWARD PM DUE IN JULY TO REPLACE ALL TUBINGS AND PUMPS IN PM KIT.
ABL90 FLEX |  | Site visit by engineer required. | SWO-180028 | Yes |  | Site visit by engineer required. | Re classified RMED | Phone call, FSE dispatch
ABL90 FLEX | Gerät überprüft. Lös Pack scheint ok. Hemolysator und Tubings überprüft, ok. Einlassdichtung ersetzt. Ohne Erfolg ,erneute Einlassdichtung, dann Sol Pack Check mehrfach ok. Pressuretest jetzt auch ok. Initialisierung abgewartet. Nach mehrfacher Kal und Kalibration und QC wieder ok. Blutmessungen durchgeführt. | Der Kunde (die Med-Technik) bittet um einen Service. | SWO-179941 | No | device checked. Lös Pack seems ok. Hemolysator and tubings checked, ok. replaced inlet seal. Without success, new inlet seal, then Sol Pack Check several times ok. Pressuretest now ok. initialization awaited. After multiple cal and calibration and QC ok again. blood measurements. | The customer (the Med-Technik) asks for a service. | Defect | HW | Inlet module | Inlet Gasket
ABL90 | Checked printer connections. | Replaced the printer. | D24C92D1E22F81BA | No | Checked printer connections. | Replaced the printer. | Defect | HW | Printer Module | Printer Module
ABL90 FLEX PLUS | Firstly replaced solution pack but only corrected glu error, Replaced sensor cassette and all ok. | Firstly replaced solution pack but only corrected glu error, Replaced sensor cassette and all ok. | SWO-179988 | No | Firstly replaced solution pack but only corrected glu error, Replaced sensor cassette and all ok. | Firstly replaced solution pack but only corrected glu error, Replaced sensor cassette and all ok. | Assisted Replacement | SC | SC | Glu
ABL90 FLEX PLUS | Preventief voedings-unit vervangen. | Preventief voedings-unit vervangen. | SWO-179815 | No | Preventive power unit replacement. | Preventive power unit replacement. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX PLUS | Voeding vervangen en een nieuwe accu geplaatst. | Voeding vervangen en een nieuwe accu geplaatst. | SWO-179864 | Yes | Replace the power supply and put a new battery. | Replace the power supply and put a new battery. | Re classified RMED | HW | Power Supply | Power Supply | Agreed WR 217-3-3
ABL90 FLEX PLUS | Preventief powersupply vervangen. Na weken door kunnnn spuiten met demiwater. | Preventief powersupply vervangen. Na weken door kunnnn spuiten met demiwater. | SWO-179831 | Yes | Preventive powersupply replacement. After weeks by kunnnn spraying with demi water. | Preventive powersupply replacement. After weeks by kunnnn spraying with demi water. | Re classified RMED | HW | Power Supply | Power Supply | Agreed WR 217-3-3
ABL90 | Sustitución de SP cogido de otro servicio Sustitución de Sonda, gasket y junta de conexión Inicio completo | Sustitución de SP cogido de otro servicio Sustitución de Sonda, gasket y junta de conexión Inicio completo | D24CC909E24A25A3 | No | Replacement of SP caught from other service Probe replacement, gasket and connection joint Full start | Replacement of SP caught from other service Probe replacement, gasket and connection joint Full start | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX |  | B | SWO-179720 | No |  |  | Defect | HW | Oxi module | Hemolyzer
AQT90 FLEX |  | B | SWO-179747 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
ABL815 FLEX | 自行更换备用件910-274后修复 by jesson | 自行更换备用件910-274后修复 by jesson | SWO-179773 | No | Repair by jesson after replacing spare parts 910-274 | Repair by jesson after replacing spare parts 910-274 | Defect | HW | Power Supply | Power Supply
ABL800 BASIC | Cleaned the Lactate Electrode Tip & Measuring Chamber , Machine Passed 2 Point Calibration & QC Level 1. If Problem Persist Again Then Lactate Membrane Need To Be Change | Cleaned the Lactate Electrode Tip & Measuring Chamber , Machine Passed 2 Point Calibration & QC Level 1. If Problem Persist Again Then Lactate Membrane Need To Be Change | SWO-181279 | No | Cleaned the Lactate Electrode Tip & Measuring Chamber , Machine Passed 2 Point Calibration & QC Level 1. If Problem Persist Again Then Lactate Membrane Need To Be Change | Cleaned the Lactate Electrode Tip & Measuring Chamber , Machine Passed 2 Point Calibration & QC Level 1. If Problem Persist Again Then Lactate Membrane Need To Be Change | Cleaned | Fluidic pathway | El/Met Module | Measuring chamber | Lac
ABL825 FLEX | S7940 and S7950 levels of the verification material VK-R7 lot #459 is low out of limits. | Performed complete oxi alignment. QC check on mean value. | D24C8F4BE22230E3 | No | S7940 and S7950 levels of the verification material VK-R7 lot #459 is low out of limits. | Performed complete oxi alignment. QC check on mean value. | Adjustment | HW | Oxi module | Oxi Complete
ABL825 FLEX | tHb is borderline low on S7940 and S7950 levels. | Performed complete oxi alignment. | D24C90DFE26130E4 | No | tHb is borderline low on S7940 and S7950 levels. | Performed complete oxi alignment. | Adjustment | HW | Oxi module | Oxi Complete
ABL815 FLEX | Fehlersuche, Layout+Messmodi und Referentwerte Standart wiederhergestellt, Layout und Refwerte neu eingegeben, Messmodi angepasst, mehrere Test durchgeführt, Werte verglichen - i.O., Konfiguration gespeichert | B | SWO-180037 | No | Troubleshooting, Layout+Measurement Modes and Referent values Standset restored, layout and ref values re-entered, measurement modes adjusted, multiple tests performed, values compared - i.O., configuration saved |  | Setup | SW | Patient layout | Patient layout
ABL825 FLEX | Had a zero current calibration error. Customer remembraned and cleaned chambers | Performed zero calibration procedure in service. Replaced blue o-ring on pH electrode | D24C60AEE2481163 | No | Had a zero current calibration error. Customer remembraned and cleaned chambers | Performed zero calibration procedure in service. Replaced blue o-ring on pH electrode | Assisted Replacement | Electrode | Electrode | Sealing ring
ABL827 FLEX | Unit is connected serially via a lantronix box. Not connected to Radiance. Unit operating normally otherwise. | Installed blind modules and removed Oxi modules. Setuo appears OK and customer can ping lantronix boxes. | D24C5C9CE20A30E0 | No | Unit is connected serially via a lantronix box. Not connected to Radiance. Unit operating normally otherwise. | Installed blind modules and removed Oxi modules. Setuo appears OK and customer can ping lantronix boxes. | Setup | His/lis | His/lis | His/lis
ABL825 FLEX | Durchsicht de Kalibrationen. Kalibrationen seit Membranierung der Referenz fehlerhaft. Messkammer gereinigt und Ref neu membraniert. Kalibrationen mehrfach durchgeführt. QC in allen Bereichen nachgeholt, ohne Beanstandung | Frau Hessberger wünscht einen Vorortservice. | SWO-179851 | No | Review of the calibrations. calibrations since membraneization of the reference faulty. measuring chamber cleaned and ref re-membraneed. calibrations performed several times. QC caught up in all areas, without complaint | Ms. Hessberger wants a suburban service. | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX | Investigation found error 316 (neon number low) 643 (DMS not connected) and 238 OXI hardware errors. | Replaced the OXI PCB and adjusted the lamp to manufacture standards. Complete OXI test passed. Service THB, pump cal, LS cal and user THB. Calibrations and QC to test. All passed with no errors. | D24C6D3FE20F7BEE | No | Investigation found error 316 (neon number low) 643 (DMS not connected) and 238 OXI hardware errors. | Replaced the OXI PCB and adjusted the lamp to manufacture standards. Complete OXI test passed. Service THB, pump cal, LS cal and user THB. Calibrations and QC to test. All passed with no errors. | Defect | HW | Oxi module | PCB, Oxi
AQT90 FLEX |  | B | D24C71F8E2155DA5 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL80 FLEX BASIC |  | 更换933-303后修复 by jesson | SWO-180093 | No |  | Repair by jesson after replacement 933-303
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-180161 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX | Flushed analyser vigourously and carried out several clot removed procedures, eventually got large clot of blood to release through inlet gasket. | Flushed analyser vigourously and carried out several clot removed procedures, eventually got large clot of blood to release through inlet gasket. | SWO-180673 | No | Flushed analyser vigourously and carried out several clot removed procedures, eventually got large clot of blood to release through inlet gasket. | Flushed analyser vigourously and carried out several clot removed procedures, eventually got large clot of blood to release through inlet gasket. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | Big CLOT :)
ABL90 FLEX | Replaced inlet probe and gasket, still same Replaced tubing for valve Pressure test ok | Replaced inlet probe and gasket, still same Replaced tubing for valve Pressure test ok | SWO-180380 | No | Replaced inlet probe and gasket, still same Replaced tubing for valve Pressure test ok | Replaced inlet probe and gasket, still same Replaced tubing for valve Pressure test ok | Defect | HW | Oxi module | Pinch valve tube
ABL90 | Powered the analyzer on and installed a new SC and SP and was able to re-create the troubleshooting loop issue. | Replaced inlet connecter with new style inlet connector. | D24D9FA55855B84B | No | Powered the analyzer on and installed a new SC and SP and was able to re-create the troubleshooting loop issue. | Replaced inlet connect with new style inlet connector. | Defect | HW | Wet section | Connector Gasket
ABL90 FLEX | Datensicherung durchgeführt. Analysator geöffnet. Computerunit ausgetauscht, Installation der Kundenspezifischen Speicherkart. Neustart und Initialisierungsphase. Eintragen der Netzwerkverbindungsdaten. Speicherung auf Windowsebene. Fehlermeldung des Ethernetadapters mit Austausch beseitigt. Neustart und Initialisierungsphase. Anbindungsdaten wurden nicht gespeichert. Erneutes herstellen der Verbindung und Speicherung. Anbindung wird nach Neustart wieder verloren gehen. (Notstromtest am 01.04.2017) Austausch der CF Karte notwendig. Dr. Pietsch wünscht keine weitere Blockierung des Analysators! Blanko CF Karte aus neuer CPU wird vorkonfiguriert und nach Terminabsprache mit Dr.Pietsch eingebaut. | CF Karte in neuen ABL90 eingesetzt. Windows und Applikationssoftware installiert. Aufspielen der Kundenkonfiguration. Speicherung der Netzwerkanbindung. Speicherung CF. Neustart. Netzwerkeinstellungen werden übernommen. Ausbau und testweise Einbau in anderen ABL90. Initialisierung. Netzwerkdaten weiterhin vorhanden. Vorinstallierte CF wird nach Terminvergabe in Kundengerät eingebaut. | SWO-180178 | No | data backup. Analyzer open. Computer unit, installation of customer-specific memory cart. Restart and initialization phase. Enter the network connection data. Windows-level storage. Eliminates the error message of the Ethernet adapter with replacement. Restart and initialization phase. Connection data was not saved. Reconnect and save. Connection will be lost again after restart. (Emergency power test on 01.04.2017) Replacement of the CF card is necessary. Dr. Pietsch does not want any further blocking of the analyzer! Blank CF card made of new CPU is preconfigured and installed upon appointment with Dr.Pietsch. | CF card inserted in new ABL90. Windows and application software. Playing on the customer configuration. Storage of the network connection. Storage CF. Restart. Network settings are applied. Removal and test installation in other ABL90. Initialization. Network data still exists. Pre-installed CF is installed in customer equipment after appointment. | Defect | HW | DMS_PC | CPU | CF Card ?
ABL90 | Customer came in this morning and analyzer was in intervention screen requesting a flush. Error code 1350. Seems to start after midnight after clot detect. Happens every day. | Disassemble analyzer. Flush out all ports and pathways. Put all parts from PM kit in analyzer. Analyzer started up fine. Ran startup Cal - all green. All internal QC fine. Ran external QC. All fine. Will check back with Pam in AM on status. If same issue occurs, I will create Techline case. Retrieved a service dump from analyzer for this. | D24DC579583389FF | No | Customer came in this morning and analyzer was in intervention screen requesting a flush. Error code 1350. Seems to start after midnight after clot detect. Happens every day. | Disassemble analyzer. Flush out all ports and pathways. Put all parts from PM kit in analyzer. Analyzer started up fine. Ran startup Cal - all green. All internal QC fine. Ran external QC. All fine. Will check back with Pam in AM on status. If same issue occurs, I will create Techline case. Retrieved a service dump from analyzer for this. | Defect | HW | Basic Unit | PM kit
ABL90 FLEX PLUS |  | B | D24E017058317DE2 | Yes |  |  |  | PER
ABL90 FLEX PLUS |  | B | D24DC9B158167EDF | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX |  | B | SWO-180132 | No |  |  | Defect | HW | Basic Unit | PCB, SP ID
AQT90 FLEX |  | B | SWO-180050 | No |  |  | Defect | HW | Cover and Chassis | Inlet Door | 4548
ABL800 BASIC | Cleaned and aligned laser position, it was not possible to adjust the value 191-193 and adjusted at 196. Only two QC ampoules were scanned. | Cleaned and aligned laser position, it was not possible to adjust the value 191-193 and adjusted at 196. Only two QC ampoules were scanned. | SWO-180690 | No | Cleaned and aligned laser position, it was not possible to adjust the value 191-193 and adjusted at 196. Only two QC ampoules were scanned. | Cleaned and aligned laser position, it was not possible to adjust the value 191-193 and adjusted at 196. Only two QC ampoules were scanned. | Cleaned | HW | Auto Check Module | Barcode reader, Autocheck Module
ABL815 FLEX |  | B | D24DC5EE58534690 | No |  |  | Assisted Replacement | Electrode | Electrode | pO2
ABL825 FLEX | pCO2 Elektrode (E788 XK-01) geprüft und (aus Kundenbestand) ausgetauscht, Funktionskontrolle ok | Kunde bittet um Service | SWO-180199 | No | pCO2 electrode (E788 XK-01) tested and replaced (from customer stock), function control ok | Customer asks for service | Assisted Replacement | Electrode | Electrode | pCO2 | Broken
ABL810 FLEX | pCO2 QC results trended low compared to peer since 12-18 reference membrane change. The results plot shows increases over time with different consumable repalcement but results have still been lower than original history. | Replaced all inlet tubings, probe & inlet nipple. Cleaned PH BG air valve & replaced module heater tubing. Reseated pH tip seal. Replaced pCO2 electrode and membrane. pCO2 QC results came up 1mmHg. Let instrument run overnight and pCO2 results are above or same as history from 12-18-2016. Instrument is operating normally. | D24DD3C75801F428 | No | pCO2 QC results trended low compared to peer since 12-18 reference membrane change. The results plot shows increases over time with different consumable repalcement but results have still been lower than original history. | Replaced all inlet tubings, probe & inlet nipple. Cleaned PH BG air valve & replaced module heater tubing. Reseated pH tip seal. Replaced pCO2 electrode and membrane. pCO2 QC results came up 1mmHg. Let instrument run overnight and pCO2 results are above or same as history from 12-18-2016. Instrument is operating normally. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL817 FLEX | Anpassung des Ausdrucks | B | SWO-180155 | No | Adaptation of the expression |  | Setup | SW | Printer Layout | Printer Layout
AQT90 FLEX |  | B | D24DCA0558003EE8 | Yes |  |  | Re classified RMED | SWU
ABL80 FLEX CO-OX OSM | Gerät zerlegt, Hemolyzer mehrfach durchgespült. tHB Kal durchgeführt. Mehrfach blutige Messungen durchgeführt. | Hr. Kruse meldet Kalkulationsfehler ist wieder aufgetreten, er sollte sich melden damit vom KDB E Teil bestellt wird, Einsatz erforderlich | SWO-180756 | No | device disassembled, Hemolyzer rinsed several times. tHB Kal. Several bloody measurements performed. | Mr. Kruse reports calculation error has occurred again, he should report so that the KDB E part is ordered, use required
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-180410 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 |  | B | D24EE02ACE2CB203 | Yes
ABL90 FLEX | Speicherkarte (mitgeliefert in CPU) in GS vorinstalliert (Windows, Applikationssoftware und Kundenkonfig vom 27.03.17) Öffnen des Analysators und einsetzen der CF Karte. Neustart und Initialisierungsphase. Neu einlesen der Betriebsmittel. Initialisierung 50Min. Kal und QC ohne Beanstandung Dauerhafte Speicherung der Radianceverbindung erfolgreich. | Speicherkarte (mitgeliefert in CPU) in GS vorinstalliert (Windows, Applikationssoftware und Kundenkonfig vom 27.03.17) Öffnen des Analysators und einsetzen der CF Karte. Neustart und Initialisierungsphase. Neu einlesen der Betriebsmittel. Initialisierung 50Min. Kal und QC ohne Beanstandung Dauerhafte Speicherung der Radianceverbindung erfolgreich. | SWO-180403 | No | Memory card (supplied in CPU) pre-installed in GS (Windows, application software and customer confectionery from 27.03.17) Open the analyzer and insert the CF card. Restart and initialization phase. Re-reading the equipment. Initialization 50min. Kal and QC without complaint Permanent storage of the radiance connection successfully. | Memory card (supplied in CPU) pre-installed in GS (Windows, application software and customer confectionery from 27.03.17) Open the analyzer and insert the CF card. Restart and initialization phase. Re-reading the equipment. Initialization 50min. Kal and QC without complaint Permanent storage of the radiance connection successfully. | Defect | HW | DMS_PC | CF Card
AQT90 FLEX |  | B | SWO-180296 | No |  |  | Defect | HW | Aspiration Module | Bend needle | 4794
AQT90 FLEX |  | B | SWO-180319 | No |  |  | Defect | HW | Aspiration Module | Tube holder
ABL825 FLEX | S7950 is out of limits low. S7940 is out or borderline low. | Repllaced and aligned hemolyzer. No help. reinstalled original part. | D24EEE86CE602038 | No | S7950 is out of limits low. S7940 is out or borderline low. | Repllaced and aligned hemolyzer. No help. reinstalled original part. | Defect | HW | Oxi module | Hemolyzer
ABL805 FLEX | TS-Eingabe überprüft, bei meinem Eintreffen waren alle Funktionen wieder einwandfrei, aktuell keine Fehler feststellbar! Kal. und QC Messung im Rahmen der FAN915-318 durchgeführt. | TS-Eingabe überprüft, bei meinem Eintreffen waren alle Funktionen wieder einwandfrei, aktuell keine Fehler feststellbar! Kal. und QC Messung im Rahmen der FAN915-318 durchgeführt. | SWO-180488 | No | TS input checked, when I arrived all functions were flawless again, currently no errors can be detected! Kal. and QC measurement within the scope of the FAN915-318. | TS input checked, when I arrived all functions were flawless again, currently no errors can be detected! Kal. and QC measurement within the scope of the FAN915-318. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | found that the lac key did not install right at installation of analyzers. | reinstalled option key for lac | D24EE3B5CE4BE4E0 | No | found that the lac key did not install right at installation of analyzers. | reinstalled option key for lac | Setup | Demo key exp | Demo Key Exp | Demo key expired
ABL820 FLEX | Ventile überprüft und gereinigt defekte Airventilmembran im pH BG Modul ersetzt Kal- und QC | Ventile überprüft und gereinigt defekte Airventilmembran im pH BG Modul ersetzt Kal- und QC | SWO-180465 | No | Valves checked and cleaned defective airvalve membrane replaced in pH BG module Kal- and QC | Valves checked and cleaned defective airvalve membrane replaced in pH BG module Kal- and QC | Defect | HW | pH/BG Module | Valve, air
ABL825 FLEX | Air line needs flushed. Also noticed halogen lamp voltage too high on calibrations intermittently. | Flushed air line through V7. Adjusted halogen lamp voltage down to proper specification. | D24F6B15CE2D06CF | No | Air line needs flushed. Also noticed halogen lamp voltage too high on calibrations intermittently. | Flushed air line through V7. Adjusted halogen lamp voltage down to proper specification. | Cleaned | Fluidic pathway | Oxi module | Oxi module
ABL817 FLEX | AutoCheck Modul im Service überprüft, Ampullen werden eingelesen, sie werden konditioniert, die Ampullen werden aufgebrochen, das Karussell dreht sich, der Deckel (Lid) öffnet sich, der Becher worin das AutoCheck-Karussell gelagert ist hebt sich bei Ansteuerung nicht an, zum Test das 14-polige Flachbandkabel getauscht - brachte keine Besserung, erforderliche Ersatzteile werden bestellt... AutoCheck Ampullen werden manuell gemessen | Einsatz zur näheren Überprüfung erforderlich + Rückruf bei Fr. schaefer 09131 8543855 | SWO-180458 | No | AutoCheck module checked in the service, ampoules are read in, they are conditioned, the ampoules are broken open, the carousel turns, the lid opens, the cup in which the AutoCheck carousel is stored does not lift at the control, to the Test the 14-pin ribbon cable replaced - brought no improvement, required spare parts are ordered... AutoCheck ampoules are measured manually | Use required for further verification + recall at Fr. schaefer 09131 8543855 | Revisit required | HW | Autocheck module | Autocheck module
ABL825 FLEX |  | B | D2506D904402D23E | No |  |  | Defect | HW | Oxi module | Spectrometer | LEH , why
ABL825 FLEX |  | B | D2506FAD4449D242 | No |  |  | Defect | HW | Oxi module | Spectrometer | LEH , why
ABL80 FLEX BASIC | 更换903-381后修复by jesson | 更换903-381后修复by jesson | SWO-180434 | No | Repair by jesson after replacement 903-381 | Repair by jesson after replacement 903-381
ABL80 OSM | Checked data logs. Ran comparison bloods obtained from another customer site. COHGB consistently ran 1.0 higher on this analyzer compared to the OSM right next to it. | Replaced hemolyzer for troubleshooting purposes. Ran comparison blood samples again between the two analyzers. Now this analyzer runs about 1.0 less than the other OSM. Unable to run tHB calibration on the other analyzer at this time due to customer running a study. Will return in the morning to run new blood samples and to run tHB calibration on both analyzers with the same lot of tHB cal solution. Also will run a VKR5 kit to validate the analyzer after troubleshooting is done. | D250808144329889 | No | Checked data logs. Ran comparison bloods obtained from another customer site. COHGB consistently ran 1.0 higher on this analyzer compared to the OSM right next to it. | Replaced hemolyzer for troubleshooting purposes. Ran comparison blood samples again between the two analyzers. Now this analyzer runs about 1.0 less than the other OSM. Unable to run tHB calibration on the other analyzer at this time due to customer running a study. Will return in the morning to run new blood samples and to run tHB calibration on both analyzers with the same lot of tHB cal solution. Also will run a VKR5 kit to validate the analyzer after troubleshooting is done.
ABL80 CO-OX | Found the molded door hinge latch broken. It was not keeping the solution pack tight against the waste luer. When I got it back up it failed the Glu low on level 1 external QC. Put in a new sensor cassette and it gave a sensor circuit error. Replaced it with another SC and all came in. | Replaced Solution pack molded door hinge assembly. Replaced SC with customer stock. | D2502CA844590CCF | No | Found the molded door hinge latch broken. It was not keeping the solution pack tight against the waste luer. When I got it back up it failed the Glu low on level 1 external QC. Put in a new sensor cassette and it gave a sensor circuit error. Replaced it with another SC and all came in. | Replaced Solution pack molded door hinge assembly. Replaced SC with customer stock.
ABL90 FLEX | Gerät überprüft, Standort kontrolliert, Sensor erneuert, tHb Kal durchgeführt, Lösungspack ausgetauscht | B | SWO-180786 | No | Device checked, location checked, sensor renewed, tHb Kal performed, solution pack replaced |  | Assisted Replacement | SC | SC | Reference
ABL90 FLEX | Probenweg gespült Verschleißteile und Schläuche sowie Quetschventil erneuert | Probenweg gespült Verschleißteile und Schläuche sowie Quetschventil erneuert | SWO-180835 | No | Sample path rinsed Wear parts and hoses as well as crushing valve renewed | Sample path rinsed Wear parts and hoses as well as crushing valve renewed | Defect | HW | Basic Unit | PM kit
ABL90 FLEX PLUS | Changed settings in RMED Dst managed service from manual to automatic as advised by RMED. All changes saved to CM card ok. Test time sync with time server ok. All 11 analysers have been done. | now all anlaysers that aer netwoerked have reverted back and lost and hour. escalte to FSE to visit and set DST service to auto.. | SWO-180875 | No | Changed settings in RMED Dst managed service from manual to automatic as advised by RMED. All changes saved to CM card ok. Test time sync with time server ok. All 11 analysers have been done. | now all anlaysers that aer netwoerked have reverted back and lost and hour. escalte to FSE to visit and set DST service to auto.. | Defect | SW | Time setting | Time setting
ABL90 FLEX PLUS |  | B | SWO-180866 | No |  |  | Assisted Replacement | SP | SP | FT
ABL90 FLEX PLUS | 2 LPs auf die E2 gebracht als Reserve für das Haus | 2 LPs auf die E2 gebracht als Reserve für das Haus | SWO-180717 | Yes | 2 LPs brought to the E2 as a reserve for the house | 2 LPs brought to the E2 as a reserve for the house | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS |  | B | D2505CED445902DE | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB
AQT90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-180944 | No | Repair after replacing test card | Repair after replacing test card | Adjustment | Consumables | Test Kit | Cartridge | 5784 (Broken)
ABL800 BASIC | Replaced Lactate membrane With New One ,Machine Passed 2 Point Calibration & QC . Machine Working Fine | Replaced Lactate membrane With New One ,Machine Passed 2 Point Calibration & QC . Machine Working Fine | SWO-181280 | No | Replaced Lactate membrane With New One ,Machine Passed 2 Point Calibration & QC . Machine Working Fine | Replaced Lactate membrane With New One ,Machine Passed 2 Point Calibration & QC . Machine Working Fine | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL825 FLEX | 3 different lot #'s of verification run low out of limits on s7950 solution. Unit is operating normally otherwise with all QC near mean values. | Replaced pump tubes. Checked measuring chambers again. Recalibrated tHb cal. tHb check values good. | D2501DEF441B344B | No | 3 different lot #'s of verification run low out of limits on s7950 solution. Unit is operating normally otherwise with all QC near mean values. | Replaced pump tubes. Checked measuring chambers again. Recalibrated tHb cal. tHb check values good. | Defect | HW | Fluidic manifold | Pump tubes
ABL825 FLEX | # lots of material tested. An occasional result came in on the low limit. Unit operating normally otherwise with no cal or QC problems. | Replaced pump tubes. Calibrated pumps. Recalibrated tHb. tHb test results good. | D2503D44443C122B | No | # lots of material tested. An occasional result came in on the low limit. Unit operating normally otherwise with no cal or QC problems. | Replaced pump tubes. Calibrated pumps. Recalibrated tHb. tHb test results good. | Defect | HW | Fluidic manifold | Pump tubes
ABL815 FLEX | ich habe die Ref.elektrode noch einmal neu bezogen und das Gerät kalibriert, Probemessungen durchgeführt und mit dem Gerät auf der Intensivst. verglichen, | ich habe die Ref.elektrode noch einmal neu bezogen und das Gerät kalibriert, Probemessungen durchgeführt und mit dem Gerät auf der Intensivst. verglichen, | SWO-180824 | No | I have re-received the ref.electrode and calibrated the device, carried out test measurements and used the device on the intensive test. Compared | I have re-received the ref.electrode and calibrated the device, carried out test measurements and used the device on the intensive test. Compared | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX |  | B | D2500E2C441B7748 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | Calibration results showing errors 0258: Neon Lamp Intensity Too Low. Ran complete oximetry test with Neon Intensity failing at 24%. | Replaced lamp unit. Adjusted halogen lamp intensity. Ran complete oximetry test with pass results. Ran tHB calibration in service mode. | D25038434415A4AE | No | Calibration results showing errors 0258: Neon Lamp Intensity Too Low. Ran complete oximetry test with Neon Intensity failing at 24%. | Replaced lamp unit. Adjusted halogen lamp intensity. Ran complete oximetry test with pass results. Ran tHB calibration in service mode. | Defect | HW | Oxi module | Lamp unit
ABL825 FLEX | Gerät mehrfach gespült und Proteinentfernungen durchgeführt. | Gerät mehrfach gespült und Proteinentfernungen durchgeführt. | SWO-181037 | No | device rinsed several times and carried out protein removals. | device rinsed several times and carried out protein removals. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL817 FLEX | Tausch des AutoCheck Moduls, AutoCheck Modul im Service überprüft, System- und Funktionskontrolle in Ordnung | Einsatz zur näheren Überprüfung erforderlich + Rückruf bei Fr. schaefer 09131 8543855 | SWO-180693 | No | Replacement of the AutoCheck module, AutoCheck module checked in service, system and function control in order | Use required for further verification + recall at Fr. schaefer 09131 8543855 | Defect | HW | Auto Check Module | Auto Check Module
ABL80 FLEX BASIC |  | B | SWO-181063 | No
ABL80 FLEX BASIC | 更换903-228一个正常。by jesson | 更换903-228一个正常。by jesson | SWO-180684 | No | Replace 903-228 one normal. by jesson | Replace 903-228 one normal. by jesson
ABL90 FLEX | tHb lässt sich nich Kalibrieren, da die Oxiwerte abgewählt sind. Nur pH, pCO2, pO2 , K, Lac. Seriennummer stimmt mit CRM und Liste von Guido Bohn. In den nächsten Tagen wird mit Oberärztin Fr. Manal geklärt ob die Konfiguration so gewollt ist. | B | SWO-181370 | Yes | tHb cannot be calibrated because the oxi values have been deselected. Only pH, pCO2, pO2 , K, Lac. Serial number matches CRM and list of Guido Bohn. In the next few days, Chief Physician Fr. Manal will clarify whether the configuration is so intentional. |  | Re classified RMED | OXI de-selected, could not do tHb Cal
ABL90 FLEX | Nach Haken bei Daylight saving rein raus stimmt Uhrzeit mit Update Now nicht mehr, Haken gestetzt, Timer Sever Test Uhrzeit okay | Nach Haken bei Daylight saving rein raus stimmt Uhrzeit mit Update Now nicht mehr, Haken gestetzt, Timer Sever Test Uhrzeit okay | SWO-181086 | No | After hooking at Daylight saving in out time is no longer correct with Update Now, hook stung, Timer Sever test time okay | After hooking at Daylight saving in out time is no longer correct with Update Now, hook stung, Timer Sever test time okay | Setup | Time setting | Time setting | Time setting
ABL90 FLEX PLUS | Wasteconnector gedemonteerd om lipje weer op zijn plaats te schuiven. | Wasteconnector gedemonteerd om lipje weer op zijn plaats te schuiven. | SWO-181119 | No | The Wasteconnector disassembled to slide the tab back into place. | The Wasteconnector disassembled to slide the tab back into place. |  | LEH: Is this Connector gasket???
ABL90 FLEX PLUS | Accu gereset. Accu spanning blijft 0%, laad niet op. Ook na een half uur nog steeds 0%. Zal een batterij bestellen. | Accu gereset. Accu spanning blijft 0%, laad niet op. Ook na een half uur nog steeds 0%. Zal een batterij bestellen. | SWO-181078 | No | Battery Reset. Battery voltage remains 0%, do not charge. Even after half an hour still 0%. Will order a battery. | Battery Reset. Battery voltage remains 0%, do not charge. Even after half an hour still 0%. Will order a battery. | Revisit required | HW | Battery | Battery
ABL90 FLEX PLUS | 回路洗浄 | B | SWO-181381 | No | Circuit cleaning |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1316
ABL90 FLEX PLUS |  | B | D25151E5BA18F54A | No |  |  | Defect | HW | Oxi module | Spectrometer
ABL90 FLEX PLUS |  | B | D2517126BA03E525 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | SC Clot
TCM CombiM | 更换910-246后修复 by jesson | 更换910-246后修复 by jesson | SWO-180682 | No | Repair by jesson after replacement 910-246 | Repair by jesson after replacement 910-246
ABL90 FLEX |  | B | SWO-180924 | No |  |  | Defect | HW | Oxi module | Hemolyzer | 134
AQT90 FLEX | 更换903-030后修复 by jesson | 更换903-030后修复 by jesson | SWO-180603 | No | Repair by jesson after replacement 903-030 | Repair by jesson after replacement 903-030 | Defect | HW | Storage Module | PCB, Cartridge Lock | 199.22.-1
ABL800 FLEX | Replace the crea regent. Calibrate system multiple time & calibration pass. | Replace the crea regent. Calibrate system multiple time & calibration pass. | SWO-181209 | No | Replace the crea regent. Calibrate system multiple time & calibration pass. | Replace the crea regent. Calibrate system multiple time & calibration pass. | Assisted Replacement | Consumables | Solution bottle | Crea expired
ABL825 FLEX | Remplacement de l'imprimante et essai : OK | Remplacement de l'imprimante et essai : OK | SWO-181197 | No | Printer replacement and test: OK | Printer replacement and test: OK | Defect | HW | Printer Module | Printer head
ABL825 FLEX |  | B | D2512E7DBA11E520 | No |  |  | Cleaned | Fluidic pathway | Oxi module | Fluidic pathway
ABL825 FLEX |  | B | D2512F7DBA07E523 | No |  |  | Cleaned | Fluidic pathway | Oxi module | Fluidic pathway
ABL800 BASIC | pCO2電極・電極コネクター交換 Clメンブラン交換 LS調整実施 CAL・QC測定データ確認 | pCO2電極・電極コネクター交換 Clメンブラン交換 LS調整実施 CAL・QC測定データ確認 | SWO-181341 | No | Exchange of pCO2 electrodes and electrode connectors Cl-Menbran Replacement Ls adjustment implemented Cal/QC measurement data confirmation | Exchange of pCO2 electrodes and electrode connectors Cl-Menbran Replacement Ls adjustment implemented Cal/QC measurement data confirmation | Assisted Replacement | Electrode | Membrane | pCO2
ABL800 BASIC | pCO2電極・電極コネクター交換 Clメンブラン交換 LS調整実施 CAL・QC測定データ確認 | B | SWO-181342 | No | Exchange of pCO2 electrodes and electrode connectors Cl-Menbran Replacement Ls adjustment implemented Cal/QC measurement data confirmation |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Reference has been remembraned X3 by customer with lot #2022. pH sens. 111.9, K 105.4, Na 87.9 and Ca 122. | Cleaned contacts, then tried a new connector. No help. A new reference electrode made no difference at first but after several calibrations and changing solutions results improved. | D251A9CEBA4E1D36 | No | Reference has been remembraned X3 by customer with lot #2022. pH sens. 111.9, K 105.4, Na 87.9 and Ca 122. | Cleaned contacts, then tried a new connector. No help. A new reference electrode made no difference at first but after several calibrations and changing solutions results improved. | Assisted Replacement | Electrode | Electrode | Reference
ABL815 FLEX | Fehlersuche, Glu und Lac Membran neu - Elektroden mit Cleaner behandelt, pCO2 wurde eben erst membraniert, Messkammern gereinigt, mehrfach kalibriert, fehlerhafte QC wiederholt | B | SWO-181396 | No | Troubleshooting, Glu and Lac membrane new - electrodes treated with cleaner, pCO2 has just been membraneed, measuring chambers cleaned, multiple calibrated, faulty QC repeated |  | Assisted Replacement | Electrode | Membrane | Glu
ABL825 FLEX | Examined data logs.This is a repetitive random issue. Examined inlet and pH BG module. | Replaced nipple connector on inlet module. Rerouted solution tubing. Verified operation of ABL825. | D25182BFBA349E5F | No | Examined data logs. This is a repetitive random issue. Examined inlet and pH BG module. | Replaced nipple connector on inlet module. Rerouted solution tubing. Verified operation of ABL825. | Defect | HW | Inlet module | Nipple kit
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-181282 | No | The replacement test card is OK. | The replacement test card is OK.
AQT90 FLEX | 2 Probenröhrchen aus Gerät entfernt. Tube Holder erneuert. Systemüberprüfung ok. 1x CRP Testmessungen zu Funktionsüberprüfung durchgeführt - ok. Messung wurde ohne Fehler durchgeführt. LQC konnte nicht durchgeführt werden, da Labor geschlossen war. | 2 Probenröhrchen aus Gerät entfernt. Tube Holder erneuert. Systemüberprüfung ok. 1x CRP Testmessungen zu Funktionsüberprüfung durchgeführt - ok. Messung wurde ohne Fehler durchgeführt. LQC konnte nicht durchgeführt werden, da Labor geschlossen war. | SWO-181389 | No | Removed 2 sample tubes from the device. Tube Holder renewed. System check ok. 1x CRP test measurements for function verification performed - ok. Measurement was carried out without errors. LQC could not be performed because Labor was closed. | Removed 2 sample tubes from the device. Tube Holder renewed. System check ok. 1x CRP test measurements for function verification performed - ok. Measurement was carried out without errors. LQC could not be performed because Labor was closed. | Defect | HW | Aspiration Module | Tube holder
ABL90 FLEX | Kunden darüber aufgeklärt das der Einsatz berechnet wird. LP ausgeliefert. | Kunden darüber aufgeklärt das der Einsatz berechnet wird. LP ausgeliefert. | SWO-181354 | Yes | informed customers that the use will be charged. LP delivered. | informed customers that the use will be charged. LP delivered. | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS |  | B | D253C6CFA63E7C11 | No |  |  | Setup | SW | Time setting | Time setting | Summer time
ABL825 FLEX |  | B | D253B3AFA60FEA15 | No |  |  | Setup | SW | Time setting | Time setting
ABL837 FLEX |  | B | D253BC46A61A7C0D | No |  |  | Setup | SW | Time setting | Time setting
ABL80 FLEX CO-OX | Fehler im Zuge der Jahreswartung beseitigt. | Fehler im Zuge der Jahreswartung beseitigt. | SWO-181465 | No | Fixed errors during annual maintenance. | Fixed errors during annual maintenance.
ABL80 OSM | Checked data logs. Ran service test. Failed cavitation absorption. | Replaced hemolyzer. Performed tHb calibration. | D25542AA1C11331B | Yes | Checked data logs. Ran service test. Failed cavitation absorption. | Replaced hemolyzer. Performed tHb calibration.
ABL80 FLEX Basic version |  | B | D25855117E366305 | No
ABL80 CO-OX | Tried a reboot and tHb calibration-No change | Checked for leaks and manually exercised pinch valve Complete oxi tests and alignment in diagnostics | D25523431C51FEE8 | No | Tried a reboot and tHb calibration-No change | Checked for leaks and manually exercised pinch valve Complete oxi tests and alignment in diagnostics
ABL90 FLEX | Fitted new probe and gasket, fitted new inlet clip which was missing Carried out all fluid transport check in service program ok | Fitted new probe and gasket, fitted new inlet clip which was missing Carried out all fluid transport check in service program ok | SWO-181623 | No | Fitted new probe and gasket, fitted new inlet clip which was missing Carried out all fluid transport check in service program ok | Fitted new probe and gasket, fitted new inlet clip which was missing Carried out all fluid transport check in service program ok | Replaced | HW | Inlet module | Probe bend
ABL90 FLEX | Touchscreen Kalibration konnte nicht durchgeführt werden. Bereiche der linken Displayhälfte werden nicht erkannt. Analysator neu gestartet. Weiterhin fehlerhaft. Touchscreen wird bestellt und nach Erhalt eingebaut | Analysator geöffnet und CPU ausgebaut. Austausch des Displays, einsetzen des Barcodescanners. Initialisierungsphase Touchscreen kalibriert und gespeichert. Kal und QC ohne Beanstandung | SWO-181626 | No | Touchscreen calibration could not be performed. Areas of the left half of the display are not recognized. Analyzer restarted. Still faulty. Touchscreen is ordered and installed upon receipt | Analyzer opened and CPU expanded. Replacement of the display, inserting the barcode scanner. Initialization phase touchscreen calibrated and stored. Kal and QC without complaint | Revisit required | HW | DMS_PC | Touch screen
ABL90 | Marge informed me that the analyzer fell off the table and the white screen appeared after start up. | Removed the back cover to check all connections to main boards and monitor. Reseated all connections and the problem remained after start up. Removed the monitor from the base unit and disassembled for troubleshooting. Reseated all connections (found one loose) and returned the monitor to the base unit. Monitor is working after power up. Replacing both consumables due to expire dates. Lac and Glu sensitivity low and will require time to pass cals and QC. | D255271B1C2EF8AF | No | Marge informed me that the analyzer fell off the table and the white screen appeared after start up. | Removed the back cover to check all connections to main boards and monitor. Reseated all connections and the problem remained after start up. Removed the monitor from the base unit and disassembled for troubleshooting. Reseated all connections (found one loose) and returned the monitor to the base unit. Monitor is working after power up. Replacing both consumables due to expire dates. Lac and Glu sensitivity low and will require time to pass cals and QC. | User error | HW | Analyzer dropped | Analyzer dropped
ABL90 FLEX PLUS | Entered service program and done all tests but showed no errors, replaced solution pack as analyzer requested but soon after replacing solution pack required replacement of sensor cassette as well. | Entered service program and done all tests but showed no errors, replaced solution pack as analyzer requested but soon after replacing solution pack required replacement of sensor cassette as well. | SWO-181821 | No | Entered service program and done all tests but showed no errors, replaced solution pack as analyzer requested but soon after replacing solution pack required replacement of sensor cassette as well. | Entered service program and done all tests but showed no errors, replaced solution pack as analyzer requested but soon after replacing solution pack required replacement of sensor cassette as well. | Assisted Replacement | SC | SC | pO2
ABL90 FLEX PLUS | Neue SC90 installiert, Credit Claim Form erstellt, Startphase abgewartet, Kal u. QC i.O. Defekte SC90: Lot R0595 - Verfallsdatum erreicht (Verfallsdatum 2017-08-07, vorher nicht benutzt!), Lot: 1009-191 - pCO2 Kalibrier- u. QC-Fehler | Neue SC90 installiert, Credit Claim Form erstellt, Startphase abgewartet, Kal u. QC i.O. Defekte SC90: Lot R0595 - Verfallsdatum erreicht (Verfallsdatum 2017-08-07, vorher nicht benutzt!), Lot: 1009-191 - pCO2 Kalibrier- u. QC-Fehler | SWO-181741 | No | New SC90 installed, Credit Claim Form created, start-up phase awaited, Kal & QC i.O. Defective SC90: Lot R0595 - Expiration date reached (expiry date 2017-08-07, not used before!), Lot: 1009-191 - pCO2 Calibration and QC error | New SC90 installed, Credit Claim Form created, start-up phase awaited, Kal & QC i.O. Defective SC90: Lot R0595 - Expiration date reached (expiry date 2017-08-07, not used before!), Lot: 1009-191 - pCO2 Calibration and QC error | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS |  | B | D2555C421C2A4D24 | Yes |  |  | Re classified RMED
ABL90 FLEX PLUS | Einlasssystem zerlegt, gereinigt, montiert und geprüft, verbogene Einlassnadel ausgerichtet. Neustart, abwarten der Kalibration und QC Kontrolle oxymetrie ausser Bereich. Hemolyzer mit proteinentferner und Wasser gespült, clot detect remove und weitere Programme im Service durchgeführt, mehrere Kalibration erforderlich wg thb drift, QC- Kontrolle | Einlasssystem zerlegt, gereinigt, montiert und geprüft, verbogene Einlassnadel ausgerichtet. Neustart, abwarten der Kalibration und QC Kontrolle oxymetrie ausser Bereich. Hemolyzer mit proteinentferner und Wasser gespült, clot detect remove und weitere Programme im Service durchgeführt, mehrere Kalibration erforderlich wg thb drift, QC- Kontrolle | SWO-181673 | No | Inlet system disassembled, cleaned, assembled and tested, bent inlet needle aligned. Restart, wait for calibration and QC control oxymetry off range. Hemolyzer with protein remover and water rinsed, clot detect remove and other programs performed in the service, multiple calibration required wg thb drift, QC control | Inlet system disassembled, cleaned, assembled and tested, bent inlet needle aligned. Restart, wait for calibration and QC control oxymetry off range. Hemolyzer with protein remover and water rinsed, clot detect remove and other programs performed in the service, multiple calibration required wg thb drift, QC control | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX PLUS | Anwender in Fehlersuche eingewiesen, neues defektes SP ersetzt, SC mit Glucosefehler auch erneuert, Gerät gespült, Kal und QCs | Anwender in Fehlersuche eingewiesen, neues defektes SP ersetzt, SC mit Glucosefehler auch erneuert, Gerät gespült, Kal und QCs | SWO-182033 | No | User instructed in troubleshooting, new defective SP replaced, SC with glucose error also renewed, device flushed, Kal and QCs | User instructed in troubleshooting, new defective SP replaced, SC with glucose error also renewed, device flushed, Kal and QCs | Assisted Replacement | SP | SP | FT
TCM400-2 |  | B | D255A1C31C0CE1D2 | No
ABL825 FLEX | Baja pantalla | B | D25514811C523413 | Yes | Low screen |  | Re classified RMED | Run 15
ABL817 FLEX |  | B | D256B327923B46B2 | No |  |  | Cleaned | HW | El/Met Module | Valve, air
ABL825 FLEX | Examined data logs. Found Oxi Hardware failure codes. | Performed halogen lamp calibration. Performed tHb and 2-point calibration. Rebooted unit. Verified operation of ABL825. | D255490C1C539F18 | No | Examined data logs. Found Oxi Hardware failure codes. | Performed halogen lamp calibration. Performed tHb and 2-point calibration. Rebooted unit. Verified operation of ABL825. | Adjustment | HW | Oxi module | Lamp unit
ABL837 FLEX |  | B | D25574A81C31E1D0 | Yes |  |  | Re classified RMED | Phone support
ABL835 FLEX | Inlet modul skiftet | Inlet modul skiftet | SWO-181512 | No | Inlet Module Shift | Inlet Module Shift | Defect | HW | Inlet module | Inlet module
ABL800 BASIC | 液センサー交換 | 液センサー交換 | SWO-181719 | No | Liquid sensor replacement | Liquid sensor replacement | Defect | HW | pH/BG Module | LS Sensor
ABL815 FLEX |  | B | D256B1C8922546AA | Yes |  |  | Re classified RMED | UPS Faulty
AQT90 FLEX |  | B | D25549D01C2370CC | No |  |  | Defect | HW | Process Module | Pump 3
ABL80 FLEX CO-OX OSM | Systemüberprüfung durchgeführt - keine Fehlermeldungen. Fehler tritt nur sporadisch auf. Oximetrie Ventil ausgebaut und gereinigt, Schläuche erneuert. QA Zyklus durchgeführt. Blutige Messungen durchgeführt. Keine Fehler aufgetreten. | Systemüberprüfung durchgeführt - keine Fehlermeldungen. Fehler tritt nur sporadisch auf. Oximetrie Ventil ausgebaut und gereinigt, Schläuche erneuert. QA Zyklus durchgeführt. Blutige Messungen durchgeführt. Keine Fehler aufgetreten. | SWO-182103 | No | System check performed - no error messages. Error occurs only sporadically. Oximetry valve removed and cleaned, hoses renewed. QA cycle. Bloody measurements carried out. No errors occurred. | System check performed - no error messages. Error occurs only sporadically. Oximetry valve removed and cleaned, hoses renewed. QA cycle. Bloody measurements carried out. No errors occurred.
ABL90 FLEX | SABİTLEYİCİ SEKMAN TAKILDI. | SABİTLEYİCİ SEKMAN TAKILDI. | SWO-182788 | Yes | THE STABLE TAB HAS BEEN INSTALLED. | THE STABLE TAB HAS BEEN INSTALLED.
ABL90 FLEX | tHb-Kal durchgeführt, Kal, QC | tHb-Kal durchgeführt, Kal, QC | SWO-181900 | No | tHb-Kal durchgeführt, Kal, QC | tHb-Kal durchgeführt, Kal, QC | Adjustment | HW | Oxi module | tHb Cal
TCM5 FLEX |  | B | D2565937921CB360 | No
ABL825 FLEX | Modul im Servicemenü überprüft, Funktion OK. Modul ausgebaut und Mechanik getestet, keine Fehler feststellbar. Montage, ausgiebiger Test aller Funktionen im Servicemenü, keine Fehler feststellbar. Vorbeugend Datenextrakt gezogen. Neustart, Kal und QC Messung OK | Modul im Servicemenü überprüft, Funktion OK. Modul ausgebaut und Mechanik getestet, keine Fehler feststellbar. Montage, ausgiebiger Test aller Funktionen im Servicemenü, keine Fehler feststellbar. Vorbeugend Datenextrakt gezogen. Neustart, Kal und QC Messung OK | SWO-181964 | No | Module checked in the service menu, function OK. Module removed and mechanics tested, no faults detectable. Installation, extensive testing of all functions in the service menu, no errors detectable. Preventive data extract drawn. Restart, Kal and QC measurement OK | Module checked in the service menu, function OK. Module removed and mechanics tested, no faults detectable. Installation, extensive testing of all functions in the service menu, no errors detectable. Preventive data extract drawn. Restart, Kal and QC measurement OK | Adjustment | HW | FlexQ Module | FlexQ Module
ABL825 FLEX | checked printer and printer voltage | removed DMS and replaced print head | D256E05D92414D62 | No | checked printer and printer voltage | removed DMS and replaced print head | Defect | HW | Printer Module | Printer head
ABL827 FLEX | Crea-Elektroden membraniert, pCO2 gereinigt, Lac konnte nicht membraniert werden, da keine Membranen vorhanden | Crea-Elektroden membraniert; pCO2 gereinigt; Lac konnte nicht membraniert werden, da keine Membranen vorhanden | SWO-182022 | No | Crea electrodes membraneed, pCO2 cleaned, Lac could not be membraneed as there are no membranes | Crea electrodes membraneed; pCO2 cleaned; Lac could not be membraned because there are no membranes | User error | Electrode | Membrane | Crea expired
ABL80 CO-OX | Able to reproduce error. | Service test, Thb cal, Cleaning and all 3 Qc levels ran. | D257735E08068B3D | No | Able to reproduce error. | Service test, Thb cal, Cleaning and all 3 Qc levels ran.
ABL90 FLEX | Analysator geöffnet und CPU ausgebaut. Austausch des Displays, einsetzen des Barcodescanners. Initialisierungsphase Touchscreen kalibriert und gespeichert. Kal und QC ohne Beanstandung | Analysator geöffnet und CPU ausgebaut. Austausch des Displays, einsetzen des Barcodescanners. Initialisierungsphase Touchscreen kalibriert und gespeichert. Kal und QC ohne Beanstandung | SWO-182245 | No | Analyzer opened and CPU expanded. Replacement of the display, inserting the barcode scanner. Initialization phase touchscreen calibrated and stored. Kal and QC without complaint | Analyzer opened and CPU expanded. Replacement of the display, inserting the barcode scanner. Initialization phase touchscreen calibrated and stored. Kal and QC without complaint | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX | Vorhergehende Konfiguration des letzte Woche ersetzten ABL90 zurückgesichert, Probelauf ok, Berichtbaren Messbereich für Ca nach unten auf 0,1mmol L korrigiert, Testmessung ok | Vorhergehende Konfiguration des letzte Woche ersetzten ABL90 zurückgesichert, Probelauf ok, Berichtbaren Messbereich für Ca nach unten auf 0,1mmol L korrigiert, Testmessung ok | SWO-182371 | No | Previous configuration of last week's replaced ABL90 backsecured, test run ok, reportable measuring range for Ca corrected down to 0.1mmol L, test measurement ok | Previous configuration of last week's replaced ABL90 backsecured, test run ok, reportable measuring range for Ca corrected down to 0.1mmol L, test measurement ok | Setup | SW | Print layout | Print layout
ABL800 FLEX | Analyser showing Sample problem while reporting and give all with ? mark. Inlet gasket and sample probe cleaned. Lac and Glu membrane changed. Analyser serviced and calibrated properly. Handed over in a fine working condition. | Analyser showing Sample problem while reporting and give all with ? mark. Inlet gasket and sample probe cleaned. Lac and Glu membrane changed. Analyser serviced and calibrated properly. Handed over in a fine working condition. | SWO-182411 | No | Analyser showing Sample problem while reporting and give all with ? mark. Inlet gasket and sample probe cleaned. Lac and Glu membrane changed. Analyser serviced and calibrated properly. Handed over in a fine working condition. | Analyser showing Sample problem while reporting and give all with ? mark. Inlet gasket and sample probe cleaned. Lac and Glu membrane changed. Analyser serviced and calibrated properly. Handed over in a fine working condition. | Assisted Replacement | Electrode | Membrane | Glu
ABL800 BASIC | インレットホルダー交換 インレットモジュール調整 タンパク除去、LS調整・ポンプCAL確認 CAL・QC測定データ確認 | 配管チューブ交換(リンスライン) タンパク除去、LS調整・ポンプCAL確認 エアライン清掃 CAL GAS2交換 CAL・QC測定データ確認 | SWO-184176 | No | Inlet holder replacement Inlet module adjustment Protein removal, LS adjustment and pump CAL confirmation Cal/QC measurement data confirmation | Pipe tube replacement (rinse line) Protein removal, LS adjustment and pump CAL confirmation Airline Cleaning CAL GAS2 replacement Cal/QC measurement data confirmation | Cleaned | Fluidic pathway | pH/BG Module | pCO2
ABL835 FLEX |  | B | D25A6424F434010F | No |  |  | Defect | HW | Oxi module | Hemolyzer | phase err
ABL800 BASIC | cleaned the pH electrode and did calibration found all results with in the range. kept instrument observation. | B | SWO-183229 | No | cleaned the pH electrode and did calibration found all results with in the range. kept instrument observation. |  | Cleaned | Electrode | Electrode | pH
ABL835 FLEX | Reviewed data logs ,and on 2Pt cals the PO2zero falls below 0 | Adjusted zero on PO2 zero. Performed several 2 Pt Cals, reads 1.0 mmHg on zero | D257BFF50835BEBA | No | Reviewed data logs ,and on 2Pt cals the PO2zero falls below 0 | Adjusted zero on PO2 zero. Performed several 2 Pt Cals, reads 1.0 mmHg on zero | Adjustment | SW | pH/BG Module | Zero current
ABL825 FLEX | checked for possible areas of printer burning and printer conections | removed DMS and replaced printer head | D25821730862945B | No | checked for possible areas of printer burning and printer conections | removed DMS and replaced printer head | Defect | HW | Printer Module | Printer head
ABL800 BASIC | プローブホルダー交換作業一式 | プローブホルダー交換作業一式 | SWO-187017 | No | Probe holder replacement work set | Probe holder replacement work set | Defect | HW | Inlet module | Probe holder
ABL800 BASIC | プローブホルダー交換作業一式 | プローブホルダー交換作業一式 | SWO-187023 | No | Probe holder replacement work set | Probe holder replacement work set | Defect | HW | Inlet module | Probe holder
ABL825 FLEX | Testweise neuen kompl. Drucker eingesetzt. Drucker druckt nicht , Papiervorschub o.k. Streckverbinder bis zur CPU kontrolliert. Keine Auffälligkeiten.! Danach einmalige Funktionalität des Druckers , dann wieder keine Reaktion auf Druckauftrag. Eventuell Flachbandkabel und oder CPU- Board defekt. Teile müssen bestellt werden. | Kabelverbindungen getestet, kein Erfolg. PC Modul getestet. Bei abgesteckten Anschlüssen 1 und 3 Funktioniert der Drucker, werden diese wieder angesteckt ( blauer und schwarzer Stecker ) keine Reaktion des Druckers. CPU Unit complete DMS910 erneuert. Danach wieder vollständige Funktion gegeben. | SWO-182518 | No | Test-wise new compl. printer. Printer does not print , paper feed ok. extend connectors up to the CPU. No abnormalities.! After that, one-time functionality of the printer , then again no reaction to print job. Possibly ribbon cable and or CPU board defective. Parts must be ordered. | Cable connections tested, no success. PC module tested. With plugged-in connectors 1 and 3 If the printer works, these are plugged back on (blue and black plug) no reaction of the printer. CPU Unit complete DMS910 renewed. Then full function is given again. | Revisit required | HW | DMS_PC | CPU
ABL90 |  | B | D25909A17E1DE7A8 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | Corrupt files in windows. | Replaced the CF card and reloaded software. | D25917C77E1B20C8 | No | Corrupt files in windows. | Replaced the CF card and reloaded software. | Defect | HW | DMS_PC | CF Card
ABL90 | Investigation found 1184 (leak detected) and 1063 (Pressure test vacuum error) | Replaced the inlet conn. gasket w holder Performed fluid transport program in service. Pressure test X2 with passing results. | D258D02C7E55C9E0 | No | Investigation found 1184 (leak detected) and 1063 (Pressure test vacuum error) | Replaced the inlet conn. gasket w holder Performed fluid transport program in service. Pressure test X2 with passing results. | Defect | HW | Inlet module | Inlet Gasket
ABL90 | Multiple SC's and SP's replaced, instrument has been in and out, will replace sensor interconnect PCB to test for defective Preamp. Tried a new SC lot 1005-198 and same error ( note same lot as the one on analyzer. it too had Glu and Lac errors | Replaced sensor connector PCB to test electronics but no failure was found. Replaced SC lot 1005-198 and was able to duplicate errors in another unit and visa versa. Returned 4 7 2017 with another lot number SC and replaced with lot 1002 and the analyzer calibrated normally. Ran 10 waste bloods and external QC, all test passed. | D25969AB7E1AF22A | No | Multiple SC's and SP's replaced, instrument has been in and out, will replace sensor interconnect PCB to test for defective Preamp. Tried a new SC lot 1005-198 and same error ( note same lot as the one on analyzer. it too had Glu and Lac errors | Replaced sensor connector PCB to test electronics but no failure was found. Replaced SC lot 1005-198 and was able to duplicate errors in another unit and visa versa. Returned 4 7 2017 with another lot number SC and replaced with lot 1002 and the analyzer calibrated normally. Ran 10 waste bloods and external QC, all test passed. | Defect | HW | Basic Unit | Sensor Interface PCB | Glu/Lac
ABL90 FLEX PLUS | Gerät überprüft, Einlaß im Sevice kontrolliert, Anschlussdichtung getauscht, vermeintlich defekte Dichtung eingebaut, Funktion o.k. | Gerät überprüft, JW durchgeführt | SWO-182793 | No | Device checked, inlet checked in the Sevice, connection seal exchanged, supposedly defective seal installed, function ok. | Device checked, JW performed | Revisit required | HW | Auto Inlet module | Auto Inlet module | Gasket?
ABL90 FLEX PLUS | Zeiteinstellung durchgeführt | Zeiteinstellung durchgeführt | SWO-182792 | No | Time setting performed | Time setting performed | Setup | SW | Time setting | Time setting | Summer time
ABL90 FLEX PLUS | Replaced faulty printer unit. | Replaced faulty printer unit. | SWO-182589 | No | Replaced faulty printer unit. | Replaced faulty printer unit. | Defect | HW | DMS_PC | Printer Module
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-182762 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Reference
ABL90 FLEX | 备件没有使用，清理了一下后修复(by tim) | 备件没有使用，清理了一下后修复(by tim) | SWO-182146 | No | Spare parts are not in use, cleanup after repair (by tim) | Spare parts are not in use, cleanup after repair (by tim) | Cleaned | HW | Basic Unit | SC Contacts
ABL825 FLEX | OXI PCB csere, SW frissítés, halogén lámpa intenzítás beállítása, teljes OXI teszt, tHb kalibráció, 2 pontos kalibráció, QC mérés | B | SWO-182695 | No | OXY PCB replacement, SW update, halogen lamp intensity adjustment, full OXY test, tHb calibration, 2 point calibration, QC measurement |  | Defect | HW | Oxi module | PCB, Oxi
ABL825 FLEX | 更换942-058后修复(by tim) | 更换942-058后修复(by tim) | SWO-182053 | No | Repair after replacement 942-058 (by tim) | Repair after replacement 942-058 (by tim) | Assisted Replacement | Electrode | Membrane | Reference
ABL835 FLEX |  | B | D25A6646F4330114 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Troubleshot with customer. Could not fix. | Sent cutomer ph and ref electrodes. We ended up replacing both. I believe ref was damaged during troubleshooting. | D258F47C7E061A3F | Yes | Troubleshot with customer. Could not fix. | Sent cutomer ph and ref electrodes. We ended up replacing both. I believe ref was damaged during troubleshooting. | Re classified RMED | Phone support
ABL800 BASIC | replaced the lactate membrane, did the calibration and samples all results found with in the range. | B | SWO-183232 | No | replaced the lactate membrane, did the calibration and samples all results found with in the range. |  | Assisted Replacement | Electrode | Membrane | Lac
ABL800 BASIC | Cleaned The Inlet Gasket , Calibrate Inlet , Pump & Liquid Sensor , Perform Protein Removal , Machine Passed 2 Point Calibration & Run The QC Level 3 Successfully. | Cleaned The Inlet Gasket , Calibrate Inlet , Pump & Liquid Sensor , Perform Protein Removal , Machine Passed 2 Point Calibration & Run The QC Level 3 Successfully. | SWO-183100 | No | Cleaned The Inlet Gasket , Calibrate Inlet , Pump & Liquid Sensor , Perform Protein Removal , Machine Passed 2 Point Calibration & Run The QC Level 3 Successfully. | Cleaned The Inlet Gasket , Calibrate Inlet , Pump & Liquid Sensor , Perform Protein Removal , Machine Passed 2 Point Calibration & Run The QC Level 3 Successfully. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | Nadel ersetzt während Rep, | Nadel ersetzt während Rep, | SWO-182743 | Yes | Needle replaced during Rep, | Needle replaced during Rep, | Re classified RMED | Needle replacement Warning
AQT90 FLEX | Fleißstrecke mit Hypochlorid gespült, Sensoren Storagewheel, Cartrige Bewegung, Proccesswheel , Inkubatorwheel, Punch down und Nadel gereinigt, Motoren gefettet, Systemüberprüfung Systemreinigung durchgeführt, während Überprüfung QC Meldung 4549 Openclose_3 - Stecker Sensor Punch down Sitz korrigiert, Sensoren und Antrieb Inlet Door gereinigt, QC durchgeführt Blindmessungen durchgeführt i.O , Anwender beobachtet und meldet sich gegebenenfalls wieder | Fleißstrecke mit Hypochlorid gespült, Sensoren Storagewheel, Cartrige Bewegung, Proccesswheel , Inkubatorwheel, Punch down und Nadel gereinigt, Motoren gefettet, Systemüberprüfung Systemreinigung durchgeführt, während Überprüfung QC Meldung 4549 Openclose_3 - Stecker Sensor Punch down Sitz korrigiert, Sensoren und Antrieb Inlet Door gereinigt, QC durchgeführt Blindmessungen durchgeführt i.O , Anwender beobachtet und meldet sich gegebenenfalls wieder | SWO-182742 | No | Hard track with hypochloride rinsed, sensors storagewheel, cartrige movement, proccesswheel, incubator wheel, punch down and needle cleaned, motors greased, system verification system cleaning performed while checking QC message 4549 Openclose_3 - plug Sensor punch down seat corrected, sensors and drive Inlet Door cleaned, QC performed blind measurements carried out i.O , user observed and reports again if necessary | Hard track with hypochloride rinsed, sensors storagewheel, cartrige movement, proccesswheel, incubator wheel, punch down and needle cleaned, motors greased, system verification system cleaning performed while checking QC message 4549 Openclose_3 - plug Sensor punch down seat corrected, sensors and drive Inlet Door cleaned, QC performed blind measurements carried out i.O , user observed and reports again if necessary | Cleaned | HW | Aspiration Module | Needle Wash Unit
AQT90 FLEX | Herrn Gieße über die augenblickliche Situation zur Installation des Etikettendruckers von Zebra unterrichtet. Über den Sachverhalt aufgeklärt, dass derzeit ( laut offizieller Information vom letzten SEKO Meeting ) Radiometer keine Unterstützung und Verantwortung für einen externen Drucker und dessen Installation mittels Druckertreiber übernimmt. Die Situation wird derzeit mit TS und Koppenhagen erörtert und dann entschieden wie weiterhin in dieser Frage gehandelt werden darf im Hinblick bezüglich auf des Medizinproduktegesetz. Herr Gieße versteht die Situation und erwartet kontinuierlich über neue Erkenntnisse unterrichtet zu werden. Herr Gieße steht im Kontakt mit dem Geräteverantwortlichen Funktionsleiter ZNA Herr Miehe. | Herrn Giese geraten keine Treiber zu installieren und auf Rückruf des zuständigen Technikers zu warten. | SWO-183062 | No | Mr. Gieße is informed about the current situation with the installation of Zebra's label printer. Informed about the fact that currently ( according to official information from the last SEKO meeting ) Radiometer assumes no support and responsibility for an external printer and its installation by means of a printer driver. The situation is currently being discussed with TS and Koppenhagen and then decided how to continue to act on this issue with regard to the Medical Devices Act. Mr. Gieße understands the situation and expects to be continuously informed about new findings. Mr. Gieße is in contact with the device manager function manager ZNA Mr. Miehe. | Mr Giese does not advise to install any drivers and to wait for the technician responsible to be called back. | Revisit required | SW | Analyzer Control Module | Printer Driver
ABL80 FLEX BASIC | 更换903-334后修复(by tim) | 更换903-334后修复(by tim) | SWO-182189 | No | Repair after replacement 903-334 (by tim) | Repair after replacement 903-334 (by tim)
ABL80 FLEX BASIC | 更换902-985后修复(by tim) | 更换902-985后修复(by tim) | SWO-182116 | No | Repair after replacement 902-985 (by tim) | Repair after replacement 902-985 (by tim)
ABL80 FLEX BASIC | Done zeroing calibration, all the values are fine, replaced the sensor cassette, after calibration, machine is working fine. | Done zeroing calibration, all the values are fine, replaced the sensor cassette, after calibration, machine is working fine. | SWO-184741 | No | Done zeroing calibration, all the values are fine, replaced the sensor cassette, after calibration, machine is working fine. | Done zeroing calibration, all the values are fine, replaced the sensor cassette, after calibration, machine is working fine.
ABL80 CO-OX | Normally, the Pco2 fails system cycle QC Level 2 low after a couple of weeks of the SC been installed. I had checked the Zero cal and it was fine. Customer tried another SC and it also failed. | I replaced the Connecter Analog Pcb. Tried the New SC and it came in. I am ordering a manifold unit just in case problem persist It could be air contamination. | D25A3472F45E96AA | No | Normally, the Pco2 fails system cycle QC Level 2 low after a couple of weeks of the SC been installed. I had checked the Zero cal and it was fine. Customer tried another SC and it also failed. | I replaced the Connecter Analog Pcb. Tried the New SC and it came in. I am ordering a manifold unit just in case problem persist It could be air contamination.
ABL90 |  | B | D25D3BA3E02A35F1 | Yes |  |  | Re classified RMED | No input
ABL90 FLEX | Überprüfung, KAL und QC | Überprüfung, KAL und QC | SWO-183037 | No | Review, KAL and QC | Review, KAL and QC | Assisted Replacement | SC | SC | pCO2/pH
ABL90 FLEX | Inlet Gasket with Holder erneuert, Einlassdichtungsgummi hat sich beim erneuern durch Kunde nach vorne raus gedrückt. Gerät und Sensorkassette gespült. Lösungspack neu installiert. Einlaufverfahren durchgeführt, alles wieder gut. | Inlet Gasket with Holder erneuert, Einlassdichtungsgummi hat sich beim erneuern durch Kunde nach vorne raus gedrückt. Gerät und Sensorkassette gespült. Lösungspack neu installiert. Einlaufverfahren durchgeführt, alles wieder gut. | SWO-183052 | No | Inlet Gasket with Holder renewed, inlet seal rubber has pushed itself forward when renewed by customer. device and sensor cartridge. Solution pack reinstalled. Run-in procedures carried out, all well again. | Inlet Gasket with Holder renewed, inlet seal rubber has pushed itself forward when renewed by customer. device and sensor cartridge. Solution pack reinstalled. Run-in procedures carried out, all well again. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Gerät überprüft, JW durchgeführt | Gerät überprüft, JW durchgeführt | SWO-183042 | No | Device checked, JW performed | Device checked, JW performed | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D25A3B9DF4424325 | No |  |  | Setup | SW | Time setting | Time setting | Daylight
ABL90 FLEX PLUS | Nieuwe batterij geplaatst. Nu wel energie 89% --> 100%. Ook normale schakeling tussen battery voeding en netspanning. | Nieuwe batterij geplaatst. Nu wel energie 89% --> 100%. Ook normale schakeling tussen battery voeding en netspanning. | SWO-182939 | Yes | New battery inserted. Now energy is 89%--> 100%. Also normal switching between battery power and mains voltage. | New battery inserted. Now energy is 89%--> 100%. Also normal switching between battery power and mains voltage. | Re classified RMED | Same day upSWO | Battery | Battery
ABL90 FLEX PLUS | Inlet mechanisme gedemonteerd en gecontroleerd en weer gemonteerd (spring wind). Alle tubings vervangen. Analyzer start direct zonder problemen op. Vijf bloedmonsters bepaald. | Inlet mechanisme gedemonteerd en gecontroleerd en weer gemonteerd (spring wind). Alle tubings vervangen. Analyzer start direct zonder problemen op. Vijf bloedmonsters bepaald. | SWO-182885 | No | Inlet mechanism disassembled and controlled and reassembled (Springwind). Replace All tubings. Analyzer starts immediately without any problems. Five blood samples determined. | Inlet mechanism disassembled and controlled and reassembled (Springwind). Replace All tubings. Analyzer starts immediately without any problems. Five blood samples determined. | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Nieuwe voedingsuinit gemonteerd. Nu wel overgang batterij status naar netspanning. | Nieuwe voedingsuinit gemonteerd. Nu wel overgang batterij status naar netspanning. | SWO-182954 | No | New feed suinit mounted. Now it does transition battery status to mains voltage. | New feed suinit mounted. Now it does transition battery status to mains voltage. | Defect | HW | Power Supply | Power Supply
AQT90 FLEX | 更换903-364后修复(by tim) | 更换903-364后修复(by tim) | SWO-182436 | Yes | Repair after replacement 903-364 (by tim) | Repair after replacement 903-364 (by tim) | Defect | HW | Process Module | Optical Unit | 131
ABL835 FLEX |  | B | D25A67FCF4050119 | No |  |  | User error | Sample volume | Sample volume | Sample volume
ABL815 FLEX | Lampenspannung überprüft - Ergebnisse OK. Es wurde schon öfters eine Proteinentfernung und eine tHb Kalibrierung durchgeführt. Der Fehler tritt dennoch weiterhin auf. Oximetrie Modul erneuert und reklamiert. Startphase durchgeführt, Lampenspannung überprüft, Kalibriert, service tHb kal und tHb kal durchgeführt. QC's durchgeführt. | Lampenspannung überprüft - Ergebnisse OK. Es wurde schon öfters eine Proteinentfernung und eine tHb Kalibrierung durchgeführt. Der Fehler tritt dennoch weiterhin auf. Oximetrie Modul erneuert und reklamiert. Startphase durchgeführt, Lampenspannung überprüft, Kalibriert, service tHb kal und tHb kal durchgeführt. QC's durchgeführt. | SWO-183009 | No | Lamp voltage checked - results OK. A protein removal and tHb calibration has been performed several times. However, the error still occurs. Oximetry module renewed and complained. Start phase carried out, lamp voltage checked, calibrated, service tHb kal and tHb kal carried out. QC's. | Lamp voltage checked - results OK. A protein removal and tHb calibration has been performed several times. However, the error still occurs. Oximetry module renewed and complained. Start phase carried out, lamp voltage checked, calibrated, service tHb kal and tHb kal carried out. QC's. | Defect | HW | Oxi module | Oxi module
ABL825 FLEX | Tausch der pCO2 Elektrode. Nach mehrmaliger Kalibration liegt die Empfindlichkeit bei über 95%, und der Staus liegt etwas über 20 mmHg, Werte in Ordnung. | Tausch der pCO2 Elektrode. Nach mehrmaliger Kalibration liegt die Empfindlichkeit bei über 95%, und der Staus liegt etwas über 20 mmHg, Werte in Ordnung. | SWO-183026 | No | Replacement of the pCO2 electrode. After several calibrations, the sensitivity is over 95%, and the traffic jams are just over 20 mmHg, values are fine. | Replacement of the pCO2 electrode. After several calibrations, the sensitivity is over 95%, and the traffic jams are just over 20 mmHg, values are fine. | Assisted Replacement | Electrode | Electrode | pCO2
ABL815 FLEX | Frontalen Barcodescanner gereinigt und justiert. Barcodescanner Flex-Q-Modul gereinigt. Funktionskontrolle beider Scanner i.O. | Frontalen Barcodescanner gereinigt und justiert. Barcodescanner Flex-Q-Modul gereinigt. Funktionskontrolle beider Scanner i.O. | SWO-182948 | No | Frontal barcode scanner cleaned and adjusted. barcode scanner Flex-Q module. Functional control of both scanners in the | Frontal barcode scanner cleaned and adjusted. barcode scanner Flex-Q module. Functional control of both scanners in the | Cleaned | HW | Basic Unit | Barcode reader, Front
ABL825 FLEX | Kabelverbindungen getestet, kein Erfolg. PC Modul getestet. Bei abgesteckten Anschlüssen 1 und 3 Funktioniert der Drucker, werden diese wieder angesteckt ( blauer und schwarzer Stecker ) keine Reaktion des Druckers. CPU Unit complete DMS910 erneuert. Danach wieder vollständige Funktion gegeben. | Kabelverbindungen getestet, kein Erfolg. PC Modul getestet. Bei abgesteckten Anschlüssen 1 und 3 Funktioniert der Drucker, werden diese wieder angesteckt ( blauer und schwarzer Stecker ) keine Reaktion des Druckers. CPU Unit complete DMS910 erneuert. Danach wieder vollständige Funktion gegeben. | SWO-183051 | No | Cable connections tested, no success. PC module tested. With plugged-in connectors 1 and 3 If the printer works, these are plugged back on (blue and black plug) no reaction of the printer. CPU Unit complete DMS910 renewed. Then full function is given again. | Cable connections tested, no success. PC module tested. With plugged-in connectors 1 and 3 If the printer works, these are plugged back on (blue and black plug) no reaction of the printer. CPU Unit complete DMS910 renewed. Then full function is given again. | Defect | HW | DMS_PC | CPU
AQT90 FLEX | Revision de mantenimiento. Quitar cubeta de la process wheel Realizacion de test de servicio. Realizacion de limpieza. Comprobacion con qc. | B | D25CA3A2E043793A | No | Maintenance review. Remove bucket from the process wheel Performing service test. Cleaning. Check with qc. |  | Defect | Consumables | SP | Waste lid
AQT90 FLEX | Beim Eintreffen Gerät prinzipiell betriebsbereit (Testkapazität niedrig). AQT geöffnet und mechanisch beschädigtes Cup entfernt, Anwender hat zuvor weitere 4 Cups entfernt bei Entnahme der Reagenzienkassette. Alle mechanischen Komponenten überprüft, keine Fehler oder Fehlfunktionen feststellbar. Fehler wurde wahrscheinlich verursacht durch fehlerhaften Abtransport verwendeter Cups und deren Blockade in Reag. Kassette. Mit aktueller, gestern neu eingesetzter Reag. Kassette Cup Transport von 10 Cups erfolgreich absolviert. Neustart, Initialisierung, Systemüberprüfung OK QC Messung wird später vom Anwender erledigt! | Beim Eintreffen Gerät prinzipiell betriebsbereit (Testkapazität niedrig). AQT geöffnet und mechanisch beschädigtes Cup entfernt; Anwender hat zuvor weitere 4 Cups entfernt bei Entnahme der Reagenzienkassette. Alle mechanischen Komponenten überprüft, keine Fehler oder Fehlfunktionen feststellbar. Fehler wurde wahrscheinlich verursacht durch fehlerhaften Abtransport verwendeter Cups und deren Blockade in Reag. Kassette. Mit aktueller, gestern neu eingesetzter Reag. Kassette Cup Transport von 10 Cups erfolgreich absolviert. Neustart, Initialisierung, Systemüberprüfung OK QC Messung wird später vom Anwender erledigt! | SWO-182883 | No | When arriving device is basically ready for operation (test capacity low). AQT opened and mechanically damaged cup removed, user has previously removed another 4 cups when removing the reagent cartridge. All mechanical components checked, no errors or malfunctions detectable. Error was probably caused by incorrect removal of used cups and their blocking in Reag. Cassette. With current, yesterday newly installed Reag. Cassette Cup Transport of 10 Cups successfully completed. Restart, Initialization, System Check OK QC measurement is done later by the user! | When arriving device is basically ready for operation (test capacity low). AQT opened and mechanically damaged cup removed; User has previously removed a further 4 cups when removing the reagent cartridge. All mechanical components checked, no errors or malfunctions detectable. Error was probably caused by incorrect removal of used cups and their blocking in Reag. Cassette. With current, yesterday newly installed Reag. Cassette Cup Transport of 10 Cups successfully completed. Restart, Initialization, System Check OK QC measurement is done later by the user! | Defect | Consumables | SP | Waste lid
ABL80 FLEX BASIC | 更换902-972后修复(tim) | 更换902-972后修复(tim) | SWO-182835 | No | Repair after replacement 902-972 (tim) | Repair after replacement 902-972 (tim)
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-183129 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX PLUS |  | B | D25BE7706A3FE6B8 | Yes |  |  | Re classified RMED | Phone call
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-183075 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Glu/Lac
ABL800 BASIC | Replaced Lactate membrane | Replaced Lactate membrane | SWO-184551 | No | Replaced Lactate membrane | Replaced Lactate membrane | Assisted Replacement | Electrode | Membrane | Lac
ABL80 FLEX BASIC |  | B | SWO-183148 | No
AQT90 FLEX | 更换903-198后修复(by tim) | 更换903-198后修复(by tim) | SWO-182906 | No | Repair after replacement 903-198 (by tim) | Repair after replacement 903-198 (by tim) | Defect | HW | Analyzer Control Module | PCB, Main Board
AQT90 FLEX |  | B | SWO-183143 | No |  |  | Defect | HW | Incubator module | Punch up | 514
AQT90 FLEX | 更换902-924后修复 | 更换902-924后修复 | SWO-183099 | No | Repair after replacement 902-924 | Repair after replacement 902-924 | Defect | HW | Analyzer Control Module | Barcode reader, display | BR2
ABL80 FLEX BASIC | 更换试剂包后修复 | 更换试剂包后修复 | SWO-184153 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit
ABL80 OSM | COMPLETED CUST. QC RUN AND CUSTOMER RUN WITH ALL RESULTS OUTSIDE NORMALS. EVALUATION OF SYSTEM RESULTS COMPLETE. | COMPLETED CLEANING PACK. RERAN CUSTOMERS QC. | D25E476656378A6D | Yes | COMPLETED CUST. QC RUN AND CUSTOMER RUN WITH ALL RESULTS OUTSIDE NORMALS. EVALUATION OF SYSTEM RESULTS COMPLETE. | COMPLETED CLEANING PACK. RERAN CUSTOMERS QC.
ABL80 CO-OX | Customer had a major leak so he had to replace the SP Then Po2 was out on level 8302 too high 165. | I replaced the Spine Pcb Manifold. If problems continue will open a techline case. | D25E33EC565696B6 | No | Customer had a major leak so he had to replace the SP Then Po2 was out on level 8302 too high 165. | I replaced the Spine Pcb Manifold. If problems continue will open a techline case.
ABL80 FLEX | 自行更换备用机上的902-972后修复（By Ran） | 自行更换备用机上的902-972后修复（By Ran） | SWO-183213 | No | Post-902-972 repair on standby (By Ran) | Post-902-972 repair on standby (By Ran)
ABL90 | VERIFIED UNIT NOT TRANSMITTING. | WORKED WITH ALERE AND OBTAINED NEW ASTM FILE. INSTALLED FILE. UNIT NOW TRANSMITTING CORRECTLY. | D25DFC49560D57FA | No | VERIFIED UNIT NOT TRANSMITTING. | WORKED WITH ALERE AND OBTAINED NEW ASTM FILE. INSTALLED FILE. UNIT NOW TRANSMITTING CORRECTLY. | Setup | His/lis | His/lis | His/lis | ASTM file
ABL90 FLEX PLUS |  | B | D25DC8275638CC7B | Yes
ABL90 FLEX PLUS | Zeitzone auf GTB (Athen) gesetzt. Zeit mit Zeitserver synchronisiert, Daten auf CF-Karte gesichert. | Zeitzone auf GTB (Athen) gesetzt. Zeit mit Zeitserver synchronisiert, Daten auf CF-Karte gesichert. | SWO-183237 | No | time zone set to GTB (Athens). Time synchronized with time server, data backed up to CF card. | time zone set to GTB (Athens). Time synchronized with time server, data backed up to CF card. | Setup | SW | Time setting | Time setting | Summer time
ABL90 FLEX PLUS | Entered service program and done fluid transport test and failed on pressure test. Investigated and straightened bent probe and reran test and passed, | Entered service program and done fluid transport test and failed on pressure test. Investigated and straightened bent probe and reran test and passed, | SWO-184297 | No | Entered service program and done fluid transport test and failed on pressure test. Investigated and straightened bent probe and reran test and passed, | Entered service program and done fluid transport test and failed on pressure test. Investigated and straightened bent probe and reran test and passed, | Unknown | HW | Auto Inlet module | Probe bend | ext. O2 Leaking
ABL90 FLEX PLUS | Einlasssonde und Dichtung ersetzt. | Startphase QC und Kalibration. | SWO-183622 | No | inlet probe and seal. | Start phase QC and calibration. | Defect | HW | Auto Inlet module | Inlet Position Cal. | Revisit - days (MRM)
AQT90 FLEX |  | B | SWO-183202 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX |  | B | D25DEB1D56260548 | No |  |  | Assisted Replacement | Consumables | Gas cylinder | Gas2
ABL825 FLEX | Fehlersuche: Fluidic- sowie AutoCheck Modul ausgebaut und CANBUS Kabelführung zu beiden Bottle Illum. PCBs und zum Fluidic Modul überprüft. Am linken Fl.Mod. Steckplatz war ein Kontaktpin seitlich abgebogen => Pin gerichtet, Stecker mehrfach testweise gesteckt und überprüft, OK. Alles wieder montiert und sämtliche Funktionen überprüft, Alles OK. Kalibrationen, QC Kontrolle, OK. | Fehlersuche: Fluidic- sowie AutoCheck Modul ausgebaut und CANBUS Kabelführung zu beiden Bottle Illum. PCBs und zum Fluidic Modul überprüft. Am linken Fl.Mod. Steckplatz war ein Kontaktpin seitlich abgebogen => Pin gerichtet, Stecker mehrfach testweise gesteckt und überprüft, OK. Alles wieder montiert und sämtliche Funktionen überprüft, Alles OK. Kalibrationen, QC Kontrolle, OK. | SWO-183501 | No | Troubleshooting: Fluidic and AutoCheck module removed and CANBUS cable routing to both Bottle Illum. PCBs and to the Fluidic module. On the left Fl.Mod. Slot was a contact pin bent sideways => Pin directed, plug plug plug plug repeatedly plugged and checked, OK. Everything reassembled and all functions checked, All OK. Calibrations, QC control, OK. | Troubleshooting: Fluidic and AutoCheck module removed and CANBUS cable routing to both Bottle Illum. PCBs and to the Fluidic module. On the left Fl.Mod. Slot was a contact pin bent sideways => Pin directed, plug plug plug plug repeatedly plugged and checked, OK. Everything reassembled and all functions checked, All OK. Calibrations, QC control, OK. | Defect | HW | Basic Unit | PCB, Illumination
ABL817 FLEX |  | B | D25EFDCDCC0FF9E9 | No |  |  | Defect | HW | El/Met Module | Valve
AQT90 FLEX |  | B | D25E2822560A4A8E | No |  |  | Cleaned | HW | Aspiration Module | Needle Wash Unit | 813
ABL80 OSM | Checked data logs for errors. Possible blockage error message intermittently. | Ran cleaning procedure for fault isolation. Also ran liquid sensor calibration with pass results. Configured remote support on all on-site OSM units. | D25F7B12CC14578C | No | Checked data logs for errors. Possible blockage error message intermittently. | Ran cleaning procedure for fault isolation. Also ran liquid sensor calibration with pass results. Configured remote support on all on-site OSM units.
ABL80 OSM | Checked data logs for errors. Intermittent Possible blockage message during QC runs. | Ran cleaning procedure. Replaced all parts from maintenance kit for fault isolation. | D25F7721CC0E5787 | No | Checked data logs for errors. Intermittent Possible blockage message during QC runs. | Ran cleaning procedure. Replaced all parts from maintenance kit for fault isolation.
ABL80 OSM | PREVIOUS SWO ESTABLISHED HEMOLYZER AND SPECTROMETER NECESSARY TO RESTORE RESULTS | REPLACED HEMOLYZER AND SPECTROMETER. COMPLETED THB CAL AND CONTROL RUN | D25F291FCC4FCEF1 | Yes | PREVIOUS SWO ESTABLISHED HEMOLYZER AND SPECTROMETER NECESSARY TO RESTORE RESULTS | REPLACED HEMOLYZER AND SPECTROMETER. COMPLETED THB CAL AND CONTROL RUN
ABL80 OSM | SAME AS SWO BABV3JAA7 FOR COHB HIGH ON CONTROL BLOOD | REPLACED HEMOLYZER AND SPECTRROMETER WITH COMPLETED THB CAL AND CONTROL RUN | D25F2A90CC5CCEF6 | Yes | SAME AS SWO BABV3JAA7 FOR COHB HIGH ON CONTROL BLOOD | REPLACED HEMOLYZER AND SPECTRROMETER WITH COMPLETED THB CAL AND CONTROL RUN
ABL80 FLEX | QC Failure for O2 and CL, Replaced Signal Data Board, Sensor Board and perform zeroing, Passed QC and two point calibration report, Working ok, problem resolve. | QC Failure for O2 and CL, Replaced Signal Data Board, Sensor Board and perform zeroing, Passed QC and two point calibration report, Working ok, problem resolve. | SWO-184534 | No | QC Failure for O2 and CL, Replaced Signal Data Board, Sensor Board and perform zeroing, Passed QC and two point calibration report, Working ok, problem resolve. | QC Failure for O2 and CL, Replaced Signal Data Board, Sensor Board and perform zeroing, Passed QC and two point calibration report, Working ok, problem resolve.
ABL80 OSM | Upon arrival confirmed fluid blockage errors. | Replaced solution pack. Found waste line disconnected reconnected. System operating normally as of 4 11 2017. | D25F3A1BCC2FFEAF | No | Upon arrival confirmed fluid blockage errors. | Replaced solution pack. Found waste line disconnected reconnected. System operating normally as of 4 11 2017.
ABL90 FLEX | SC交換。 SC交換後はQC3レベルとも範囲内に収まる。 | B | SWO-184155 | No | SC replacement. After SC exchange, both QC3 levels fit within the range. |  | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX | JW durchgeführt, Sensorkassette ausgetauscht | JW durchgeführt, Sensorkassette ausgetauscht | SWO-183920 | No | JW performed, sensor cartridge replaced | JW performed, sensor cartridge replaced | Assisted Replacement | SC | SC | Glu
ABL90 FLEX | Einlassarm (R1102N010) eingebaut, System und Funktion i.O. | Einlassarm (R1102N010) eingebaut, System und Funktion i.O. | SWO-183833 | No | Inlet arm (R1102N010) built-in, system and function in the | Inlet arm (R1102N010) built-in, system and function in the | Defect | HW | Inlet module | Inlet Handle | Broken
ABL90 | The customer replaced the printer paper and then had general printer errors. The customer shut down the instrument but had liquid sensor errors after restarting the instrument. Checked all internal connections and found printer power connection was loose which may have caused the instrument electronics to malfunction. | Reseated printer connection. Verified instrument operation. | D25F7E0CCC386F62 | No | The customer replaced the printer paper and then had general printer errors. The customer shut down the instrument but had liquid sensor errors after restarting the instrument. Checked all internal connections and found printer power connection was loose which may have caused the instrument electronics to malfunction. | Reseated printer connection. Verified instrument operation. | Adjustment | HW | Printer Module | Cable | Loose
ABL90 FLEX PLUS |  | B | D25F8630CC03900B | Yes
ABL90 FLEX PLUS | manually turned curved wheel assembly to the closed resting position. Tested inlet in service mode and functioned well in all positions. | manually turned curved wheel assembly to the closed resting position. Tested inlet in service mode and functioned well in all positions. | SWO-184318 | No | manually turned curved wheel assembly to the closed resting position. Tested inlet in service mode and functioned well in all positions. | manually turned curved wheel assembly to the closed resting position. Tested inlet in service mode and functioned well in all positions. | Defect | HW | Auto Inlet module | Inlet Position Cal.
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-183631 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Glu
AQT90 FLEX | 更换903-138后修复(By Ran) | 更换903-138后修复(By Ran) | SWO-183378 | No | Repair after replacement 903-138 (By Ran) | Repair after replacement 903-138 (By Ran) | Defect | HW | Incubator module | PCB, Incubator | 5685
ABL826 FLEX |  | B | D260E430424DDBE0 | No |  |  | Assisted Replacement | Electrode | Electrode | K+
ABL825 FLEX | found tubing crimped #5 line, hole in solution pump tubing,glu missing o ring, ph not sealing, had impeded probe,rinse errors. blood sample tested with clotted instrument | replaced tubings,inlet gasket, o rings ph and glu, remembraned glu,replaced inlet nipples | D25F4B32CC0F05A8 | Yes | found tubing crimped #5 line, hole in solution pump tubing,glu missing o ring, ph not sealing, had impeded probe,rinse errors. blood sample tested with clotted instrument | replaced tubings,inlet gasket, o rings ph and glu, remembraned glu,replaced inlet nipples
ABL837 FLEX |  | B | D260EA0F423DDBE2 | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | pCO2
ABL825 FLEX | Checked all logs to verify past performance and found 581 errors on QC and patient samples. Last THB cal 1.0277 | Performed complete OXI test with passing results. Checked Verified lamp adjustment - PASS. Replaced Hemolyzer, adjusted lamp (3257)and service THB (0.97) user THB (0.9772) | D25F2390CC4C70B2 | No | Checked all logs to verify past performance and found 581 errors on QC and patient samples. Last THB cal 1.0277 | Performed complete OXI test with passing results. Checked Verified lamp adjustment - PASS. Replaced Hemolyzer, adjusted lamp (3257)and service THB (0.97) user THB (0.9772) | Defect | HW | Oxi module | Hemolyzer
AQT90 FLEX | Proteinreinigung nach Service Manuel durchgeführt, Systemreinigung durchgeführt. Reinigung des Washblocks, der Pipette und des Pipetteblocks. Reinigung des vorderen Filters. Reinigung der Cup Wash unit. Schläuche Pumpen P2 und P4 unterschiedlich lang und Abrieb an Achsen. Pumpenköpfe getauscht. Blankmessungen durchgeführt, Messwerte innerhalb Grenzwerte, in Ordnung. Flowwerte überprüft, sehr gut, Werte liegen zwischen 1100 und 1200. | Proteinreinigung nach Service Manuel durchgeführt, Systemreinigung durchgeführt. Reinigung des Washblocks, der Pipette und des Pipetteblocks. Reinigung des vorderen Filters. Reinigung der Cup Wash unit. Schläuche Pumpen P2 und P4 unterschiedlich lang und Abrieb an Achsen. Pumpenköpfe getauscht. Blankmessungen durchgeführt, Messwerte innerhalb Grenzwerte, in Ordnung. Flowwerte überprüft, sehr gut, Werte liegen zwischen 1100 und 1200. | SWO-183837 | No | Protein cleaning carried out according to Service Manuel, system cleaning performed. Cleaning of the washblock, the pipette and the pipette block. Cleaning the front filter. Cleaning the Cup Wash unit. Hoses pumps P2 and P4 different length and abrasion on axes. pump heads. Blank measurements performed, measurements within limit values, fine. Flow values checked, very well, values are between 1100 and 1200. | Protein cleaning carried out according to Service Manuel, system cleaning performed. Cleaning of the washblock, the pipette and the pipette block. Cleaning the front filter. Cleaning the Cup Wash unit. Hoses pumps P2 and P4 different length and abrasion on axes. pump heads. Blank measurements performed, measurements within limit values, fine. Flow values checked, very well, values are between 1100 and 1200. | Cleaned | HW | Aspiration Module | Needle Wash Unit | 1223
AQT90 FLEX | Prozeßrad nicht drehbar, optische Einheit demontiert, Verklemmtes Testcup entfernt, Montage, Neustart ok, SafePositions ok, Startphase, Systemüberprüfung ok, QK TnI ok | Reinigung, Funktionen von Storage-, Inkubator- und Prozesswheel kontrolliert, Testcups aus Prozessrad und Inkubatorrad entfernt, Startphase, Systemtest ok, TestQK, Setsave 22, erneut Testcup entfernt sowie Solutionpack kontrolliert und ersetzt, Startphase, Systemtest ok, TestQK ok | SWO-183887 | No | Process wheel not rotatable, optical unit disassembled, clamped test cup removed, assembly, restart ok, SafePositions ok, start phase, system check ok, QK TnI ok | Cleaning, functions of storage, incubator and process wheel controlled, test cups removed from process wheel and incubator wheel, start phase, system test ok, TestQK, Setsave 22, re-removed test cup as well as solutionpack controlled and replaced, start phase, system test ok , TestQK ok | Defect | Consumables | SP | Cup Stack
ABL80 CO-OX | Normally the level one QC for Pco2 runs high on both units. When internal Qc fails it is either the Pco2 low on 8302 or the Po2 high on 8303 internal Qc. Ex: Pco2 68.6 too low or P02 92 too high. This happens with many different lots of both SC SP. The room temperatures run between 66-68 F-No major air draft noticed. On this unit I have replaced Sensor connector Analog Spine Manifold with no change in its operation. | I tried many things different SC Same problem Verified zero cal several times ok finally replaced the clear tubings from a spare kit and that seems to have solved the problem I am opening a Techline case, because they have experienced on off problems on both units specially on Pco2. | D2605FF9422C7EAF | No | Normally the level one QC for Pco2 runs high on both units. When internal Qc fails it is either the Pco2 low on 8302 or the Po2 high on 8303 internal Qc. Ex: Pco2 68.6 too low or P02 92 too high. This happens with many different lots of both SC SP. The room temperatures run between 66-68 F-No major air draft noticed. On this unit I have replaced Sensor connector Analog Spine Manifold with no change in its operation. | I tried many things different SC Same problem Verified zero cal several times ok finally replaced the clear tubings from a spare kit and that seems to have solved the problem I am opening a Techline case, because they have experienced on off problems on both units specially on Pco2.
ABL90 | Checked all logs to verify performance and found activities when staff had printer issues and restarted and or opened and closed the printer lid. | Replaced the printer and tested with passing results. Performed liquid transport programs for preventative maintenance. customer will continue to monitor. | D260A6D34209BDF5 | No | Checked all logs to verify performance and found activities when staff had printer issues and restarted and or opened and closed the printer lid. | Replaced the printer and tested with passing results. Performed liquid transport programs for preventative maintenance. customer will continue to monitor. | Defect | HW | Printer Module | Printer Module
ABL90 | Kathy is running qc studies and, pO2 is intermittently failing on each one of the external qc. She replaced the sensor cassette (Part #944-005, Lot #966) on 4 8 17 for the same problem. | 4 12 17 Upon arrival, I ran 5 blood samples to condition the SC and, ran 3 levels of qc, all passing. Ran 3 calibrations, ran qc again and, qc failied on level 1 for pO2 and, level 2 for pCO2. Back-flush the analyzer, ran qc, failed again on level 3. Last replaced the inlet gasket, the inlet connector gasket and, the sensor interface PCB. Interchanged the sensor cassette with another ABL90. 4 13 17 Follow up visit to check analyzer performance. No problems found today. | D260BA6642513143 | No | Kathy is running qc studies and, pO2 is intermittently failing on each one of the external qc. She replaced the sensor cassette (Part #944-005, Lot #966) on 4 8 17 for the same problem. | 4 12 17 Upon arrival, I ran 5 blood samples to condition the SC and, ran 3 levels of qc, all passing. Ran 3 calibrations, ran qc again and, qc failied on level 1 for pO2 and, level 2 for pCO2. Back-flush the analyzer, ran qc, failed again on level 3. Last replaced the inlet gasket, the inlet connector gasket and, the sensor interface PCB. Interchanged the sensor cassette with another ABL90. 4 13 17 Follow up visit to check analyzer performance. No problems found today. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX PLUS | Startphase QC und Kalibration. | Startphase QC und Kalibration. | SWO-183915 | No | Start phase QC and calibration. | Start phase QC and calibration. | Defect | HW | Auto Inlet module | Auto Inlet module
ABL90 FLEX PLUS | E-Teil eingetragen | Startphase QC und Kalibration. | SWO-187361 | Yes | E-part entered | Start phase QC and calibration. | Re classified RMED | Double entry
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-184017 | Yes | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-184154 | Yes | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-186650 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Reference | 1361
ABL90 FLEX | 连续更换2张测试卡后修复 | 连续更换2张测试卡后修复 | SWO-184015 | Yes | Repair after 2 consecutive test cards | Repair after 2 consecutive test cards
ABL800 BASIC | 配管チューブ交換(リンスライン) タンパク除去、LS調整・ポンプCAL確認 エアライン清掃 CAL GAS2交換 CAL・QC測定データ確認 | 配管チューブ交換(リンスライン) タンパク除去、LS調整・ポンプCAL確認 エアライン清掃 CAL GAS2交換 CAL・QC測定データ確認 | SWO-184147 | No | Pipe tube replacement (rinse line) Protein removal, LS adjustment and pump CAL confirmation Airline Cleaning CAL GAS2 replacement Cal/QC measurement data confirmation | Pipe tube replacement (rinse line) Protein removal, LS adjustment and pump CAL confirmation Airline Cleaning CAL GAS2 replacement Cal/QC measurement data confirmation | Assisted Replacement | Consumables | Gas cylinder | Gas2
ABL800 BASIC | 配管チューブ交換(リンスライン) コンタミ除去、LS調整・ポンプCAL確認 | 配管チューブ交換(リンスライン) コンタミ除去、LS調整・ポンプCAL確認 | SWO-184161 | No | Pipe tube replacement (rinse line) Contamination removal, LS adjustment and pump CAL confirmation | Pipe tube replacement (rinse line) Contamination removal, LS adjustment and pump CAL confirmation | Assisted Replacement | Consumables | Gas cylinder | Gas2
ABL825 FLEX |  | Der Kunde bittet um einen Service. | SWO-184094 | Yes |  | The customer asks for a service. | Re classified RMED | PM
ABL825 FLEX | Checked data cal logs for drift errors, found several lactate cal errors. Found a crack in the ca++ measuring well. Checked grounding strap for modules. Following power up all calibration cleared following the third cal. QC and samples were run, the ca++ results for the samples read lower when compared to the other analyzer.Performed a cal2, the sens on ca++ went from 99% to 86% and lactate drift was at .3. After several cals wasn't able to clear errors untill cleaning connectors. | Replaced measuring chamber, cleaned contacts on reference and ca++ connectors. | D266FF80901B58B8 | No | Checked data cal logs for drift errors, found several lactate cal errors. Found a crack in the ca++ measuring well. Checked grounding strap for modules. Following power up all calibration cleared following the third cal. QC and samples were run, the ca++ results for the samples read lower when compared to the other analyzer. Performed a cal2, the sens on ca++ went from 99% to 86% and lactate drift was at .3. After several cals wasn't able to clear errors untill cleaning connectors. | Replaced measuring chamber, cleaned contacts on reference and ca++ connectors. | Defect | HW | El/Met Module | Measuring chamber | Cracked
ABL825 FLEX | Neon Intensity too low. | Completed a oxi tests. Lamp failed neon intensity. Warranty replacement Replaced lamp unit. Completed oxi tests again. Completed all service tests. | D260A5D54210CD73 | No | Neon Intensity too low. | Completed a oxi tests. Lamp failed neon intensity. Warranty replacement Replaced lamp unit. Completed oxi tests again. Completed all service tests. | Defect | HW | Oxi module | Lamp unit | Neon
ABL80 FLEX BASIC | 更换902-972一个正常。(By Ran) | 更换902-972一个正常。(By Ran) | SWO-183938 | No | Replace 902-972 one normal. (By Ran) | Replace 902-972 one normal. (By Ran)
ABL80 FLEX BASIC | 更换902-972后修复(By Ran) | 更换902-972后修复(By Ran) | SWO-183864 | No | Repair after replacement 902-972 (By Ran) | Repair after replacement 902-972 (By Ran)
ABL80 FLEX BASIC | Found Temp A, B, D errors, Sensor cassette got leaked solution, sensor connector got damaged, replaced the standby Sensor connector PCB, performed the Zeroing, installed the new SC found all the parameters are working, performed the test found results are Ok, instrument working satisfactorily. Defective SC details: Lot: 234692, Sn: 1562492. New SC details: Lot: 235032, Sn: 1566297 | B | SWO-184179 | Yes | Found Temp A, B, D errors, Sensor cassette got leaked solution, sensor connector got damaged, replaced the standby Sensor connector PCB, performed the Zeroing, installed the new SC found all the parameters are working, performed the test found results are Ok, instrument working satisfactorily. Defective SC details: Lot: 234692, Sn: 1562492. New SC details: Lot: 235032, Sn: 1566297
ABL90 FLEX | Der Time-Server stellt nicht automatisch von Winter- auf Sommerzeit um, daher entsprechende Einstellung im Betriebssystem vorgenommen. Nach Geräteneustart bleibt die korrekte Zeit hinterlegt. Einstellung analog am AQT S N 838R0418N014 vorgenommen. | Der Time-Server stellt nicht automatisch von Winter- auf Sommerzeit um, daher entsprechende Einstellung im Betriebssystem vorgenommen. Nach Geräteneustart bleibt die korrekte Zeit hinterlegt. Einstellung analog am AQT S N 838R0418N014 vorgenommen. | SWO-184446 | No | The Time server does not automatically switch from winter to daylight saving time, so the appropriate setting is made in the operating system. After the device is restarted, the correct time is stored. setting analogously to the AQT S N 838R0418N014. | The Time server does not automatically switch from winter to daylight saving time, so the appropriate setting is made in the operating system. After the device is restarted, the correct time is stored. setting analogously to the AQT S N 838R0418N014. | Setup | SW | Time setting | Time setting | Summer time
ABL90 FLEX | Hemolysator erneuert, Ventilschlauch getauscht, tHb Kalibration, Systemkontrolle. | Fehlersuche, laut Frau Paul QC grenzwertig seit letzten Serviceeinsatz und jetzt außer Bereich bei pO2 Level C und erneut Leckage Fehler->Überprüfung des internen Flüssigkeitstransportweges von Abfall bis Einlass, Sensor Bay geöffnet und LP herausgenommen->erneutes Einsetzen des LP sehr schwergängig->mittlere Kunststoffführung des LP am Analysator sitzt nicht korrekt->diese Führung wieder festgezogen, Leckagetest bestanden Startphase bestanden->QC bei pO2 in allen Leveln Zielwert | SWO-184343 | No | Hemolysator renewed, valve hose exchanged, tHb calibration, system control. | Troubleshooting, according to Ms Paul QC borderline since last service use and now out of range at pO2 Level C and again leakage error->checking of the internal liquid transport path from waste to inlet, Sensor Bay opened and LP removed->again Inserting the LP very difficult-to-use->medium plastic guide of the LP on the analyzer sits incorrectly->this guide reattached, leakage test passed Start phase passed->QC at pO2 at all levels target value | Revisit required | HW | Oxi module | Hemolyzer | Leakage
ABL90 FLEX | LP erneuert. Startphase, Kal, QC und Blutmessung iO. | Fehlersuche, laut Frau Paul QC grenzwertig seit letzten Serviceeinsatz und jetzt außer Bereich bei pO2 Level C und erneut Leckage Fehler->Überprüfung des internen Flüssigkeitstransportweges von Abfall bis Einlass, Sensor Bay geöffnet und LP herausgenommen->erneutes Einsetzen des LP sehr schwergängig->mittlere Kunststoffführung des LP am Analysator sitzt nicht korrekt->diese Führung wieder festgezogen, Leckagetest bestanden Startphase bestanden->QC bei pO2 in allen Leveln Zielwert | SWO-184415 | No | LP renewed. Start phase, Kal, QC and blood measurement iO. | Troubleshooting, according to Ms Paul QC borderline since last service use and now out of range at pO2 Level C and again leakage error->checking of the internal liquid transport path from waste to inlet, Sensor Bay opened and LP removed->again Inserting the LP very difficult-to-use->medium plastic guide of the LP on the analyzer sits incorrectly->this guide reattached, leakage test passed Start phase passed->QC at pO2 at all levels target value | Assisted Replacement | SP | SP | FT | Leakage
ABL90 | Liquid sensors 1 and 2 are failing between the sensor cassette which are part of the bay unit. The values for LS1 and LS2 are 245 high out of range which could indicate and electrical malfunction of the sensors. | Replaced sensor cassette bay unit. Verified instrument operation. | D261A402B8046F75 | No | Liquid sensors 1 and 2 are failing between the sensor cassette which are part of the bay unit. The values for LS1 and LS2 are 245 high out of range which could indicate and electrical malfunction of the sensors. | Replaced sensor cassette bay unit. Verified instrument operation. | Defect | HW | Basic Unit | Sensor Cassette Bay
ABL90 | Found the analyzer is asking for back-flush. It is suspecting a clot. | I checked the fluid transport in service mode and, found all the tests had failed. When I tried back-flushing the analyzer, I found a heavy back-pressure. I inspected the fluid path by individually flushing each line and, the inlet probe. Found the sensor bay unit had a major blockage. Replaced anew sensor bay unit. Performed the fluid transport checks and, they all passed within specifications. | D261EC01B85457BB | No | Found the analyzer is asking for back-flush. It is suspecting a clot. | I checked the fluid transport in service mode and, found all the tests had failed. When I tried back-flushing the analyzer, I found a heavy back-pressure. I inspected the fluid path by individually flushing each line and, the inlet probe. Found the sensor bay unit had a major blockage. Replaced anew sensor bay unit. Performed the fluid transport checks and, they all passed within specifications. | Defect | HW | Basic Unit | Sensor Cassette Bay
ABL90 FLEX | Kal, QK fehlerhaft, SP90 vor zwei Tagen neu, Sensorkasette ersetzt, Startphase | Kal, QK fehlerhaft, SP90 vor zwei Tagen neu, Sensorkasette ersetzt, Startphase | SWO-184368 | Yes | Kal, QK faulty, SP90 new two days ago, sensor calet replaced, start-up phase | Kal, QK faulty, SP90 new two days ago, sensor calet replaced, start-up phase | Re classified RMED | SC | SC | Reference
ABL90 FLEX | Siehe SWO 184368 Beide Unterschriften zu SWO 184368 & zu SWO 184404 | Kal, QK fehlerhaft, SP90 vor zwei Tagen neu, Sensorkasette ersetzt, Startphase | SWO-184367 | No | See SWO 184368 Both signatures for SWO 184368 & SWO 184404 | Kal, QK faulty, SP90 new two days ago, sensor calet replaced, start-up phase | Assisted Replacement | SC | SC | Reference
ABL90 FLEX PLUS | Im Datei Archive nachgeschaut. Die Fehlermeldungen sind keine Probe erkannt. Den Analysator im Service Menü durchgecheckt. Keine Meldungen vorhanden. | Im Datei Archive nachgeschaut. Die Fehlermeldungen sind keine Probe erkannt. Den Analysator im Service Menü durchgecheckt. Keine Meldungen vorhanden. | SWO-184258 | No | Looked in the Archive file. The error messages are not detected as a sample. Checked the analyzer in the service menu. There are no messages. | Looked in the Archive file. The error messages are not detected as a sample. Checked the analyzer in the service menu. There are no messages. | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX PLUS | Mit MED. Techniker Herr Leistner Lösungspack erneut installiert. Lösungspack war nicht vollständig nach unten gedrückt. Einlassdichtung streifte an LP und blockierte die korrekte Positionierung. Einlassdichtung gewechselt und gleichen LP wieder installiert. Gerät durchgespült, tHb Kal. fällig und durchgeführt . Reklamationsverfahren erklärt. | B | SWO-184422 | No | With MED. Technician Mr. Leistner solution pack reinstalled. Solution pack was not fully pushed down. Inlet seal swiped at LP and blocked the correct positioning. inlet seal changed and reinstalled the same LP. device flushed through, tHb Kal. due and carried out . complaint procedure. |  | User error | SP | SP Activation | SP Activation
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-184470 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | FT | 184
AQT90 FLEX | 更换备用机备件903-584后修复(By Ran) | 更换备用机备件903-584后修复(By Ran) | SWO-184207 | No | Repair after replacement of spare parts 903-584 (By Ran) | Repair after replacement of spare parts 903-584 (By Ran) | Defect | HW | Aspiration Module | Tube holder
ABL800 FLEX | Replace the Crea reagent & calibrate the system. | Replace the Crea reagent & calibrate the system. | SWO-184571 | No | Replace the Crea reagent & calibrate the system. | Replace the Crea reagent & calibrate the system. | Assisted Replacement | Consumables | Consumables | Crea expired
ABL825 FLEX |  | B | D261B080B83C525C | Yes
ABL825 FLEX |  | B | D261E9BBB83EA2FC | No |  |  | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL837 FLEX |  | B | D261F3D1B85DD03F | No |  |  | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Allgemeine Kontrolle i.o Membranenwartung Lac gleiche Lot R3306 durchgeführt i.o Kal. i.o QC Level 3 i.o Erneut Kal. i.o Alle QC i.o Erneut Kal. Empfindlichkeit schwankt QC Level 3 sinkend Lac Elektrode ersetzt mit neuer Lot Membrane aus Thalwil Alte Lac Elektrode YH03 = Garantiersatz i.o Kal. und QC i.o Letzte 2. Kal. wurde abgebrochen Gleicher Fehler wie in der Nacht Ventile pH BG-Modul ausgebaut und gereinigt i.o Spüllösung ersetzt i.o Schlausch Spüllösung ersetzt i.o Zwei 2. Kal. i.o | Allgemeine Kontrolle i.o Membranenwartung Lac gleiche Lot R3306 durchgeführt i.o Kal. i.o QC Level 3 i.o Erneut Kal. i.o Alle QC i.o Erneut Kal. Empfindlichkeit schwankt QC Level 3 sinkend Lac Elektrode ersetzt mit neuer Lot Membrane aus Thalwil Alte Lac Elektrode YH03 = Garantiersatz i.o Kal. und QC i.o Letzte 2. Kal. wurde abgebrochen Gleicher Fehler wie in der Nacht Ventile pH BG-Modul ausgebaut und gereinigt i.o Spüllösung ersetzt i.o Schlausch Spüllösung ersetzt i.o Zwei 2. Kal. i.o | SWO-184384 | No | General control i.o Membrane Maintenance Lac Same Lot R3306 Performed i.o Kal. Ok QC Level 3 i.o Once again Kal. Ok All QC i.o Once again Kal. Sensitivity fluctuates QC Level 3 sinking Lac electrode replaced with new solder membrane from Thalwil Old Lac Electrode YH03 = Warranty Kit i.o Kal. and QC i.o Last 2nd Kal. has been canceled Same mistake as at night Valves pH BG module removed and cleaned i.o Rinsing solution replaces i.o Schlausch rinsing solution replaces i.o Two 2nd Kal. Ok | General control i.o Membrane Maintenance Lac Same Lot R3306 Performed i.o Kal. Ok QC Level 3 i.o Once again Kal. Ok All QC i.o Once again Kal. Sensitivity fluctuates QC Level 3 sinking Lac electrode replaced with new solder membrane from Thalwil Old Lac Electrode YH03 = Warranty Kit i.o Kal. and QC i.o Last 2nd Kal. has been canceled Same mistake as at night Valves pH BG module removed and cleaned i.o Rinsing solution replaces i.o Schlausch rinsing solution replaces i.o Two 2nd Kal. Ok | Cleaned | HW | pH/BG Module | Valve
AQT90 FLEX | 機械修理後セットアップ、ラディアンスとの接続確認、測定確認 | B | SWO-184453 | No | Machine repair setup, connection confirmation with radiance, measurement confirmation |  | Setup | His/lis | His/lis | His/lis
AQT90 FLEX | Reinigung, Funktionen von Storage-, Inkubator- und Prozesswheel kontrolliert, Testcups aus Prozessrad und Inkubatorrad entfernt, Startphase, Systemtest ok, TestQK, Setsave 22, erneut Testcup entfernt sowie Solutionpack kontrolliert und ersetzt, Startphase, Systemtest ok, TestQK ok | Reinigung, Funktionen von Storage-, Inkubator- und Prozesswheel kontrolliert, Testcups aus Prozessrad und Inkubatorrad entfernt, Startphase, Systemtest ok, TestQK, Setsave 22, erneut Testcup entfernt sowie Solutionpack kontrolliert und ersetzt, Startphase, Systemtest ok, TestQK ok | SWO-184404 | No | Cleaning, functions of storage, incubator and process wheel controlled, test cups removed from process wheel and incubator wheel, start phase, system test ok, TestQK, Setsave 22, re-removed test cup as well as solutionpack controlled and replaced, start phase, system test ok , TestQK ok | Cleaning, functions of storage, incubator and process wheel controlled, test cups removed from process wheel and incubator wheel, start phase, system test ok, TestQK, Setsave 22, re-removed test cup as well as solutionpack controlled and replaced, start phase, system test ok , TestQK ok | Defect | Consumables | SP | Cup Stack | Revisit
ABL90 |  | B | D262D5332E332E4D | No |  |  | Revisit required | HW | DMS_PC | CPU
ABL90 |  | B | D268181F06340DC1 | No |  |  | Unknown | Unknown | Unknown | Unknown
ABL90 FLEX PLUS |  | B | D2715BF54041D70C | No |  |  | User error | SP Expired | SP Expired | SP Expired
ABL90 FLEX PLUS |  | B | D262F4622E0CE2F3 | No |  |  | Defect | HW | Auto Inlet module | Inlet Position Cal.
AQT90 FLEX | 自行更换备用机备件636-550后修复(By Ran) | 自行更换备用机备件636-550后修复(By Ran) | SWO-184325 | No | Self-replacement of spare parts 636-550 after repair (By Ran) | Self-replacement of spare parts 636-550 after repair (By Ran) | Defect | HW | Process Module | Process Unit sensor | 5257
AQT90 FLEX | 更换903-364正常。(By Ran) | 更换903-364正常。(By Ran) | SWO-184478 | No | Replacement of 903-364 is normal. (By Ran) | Replacement of 903-364 is normal. (By Ran) | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX | 更换903-329一个正常。(By Ran) | 更换903-329一个正常。(By Ran) | SWO-184292 | No | Replace the 903-329 one normal. (By Ran) | Replace the 903-329 one normal. (By Ran) | Defect | HW | Aspiration Module | Needle Wash Unit
ABL800 BASIC | pH電極交換 | pH電極交換 | SWO-184455 | No | pH electrode replacement | pH electrode replacement | Assisted Replacement | Electrode | Electrode | pH
AQT90 FLEX |  | B | D262E0822E4A03E6 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-184635 | No | Replace the test card | Replace the test card
ABL90 FLEX | Fehlersuche, laut Frau Paul QC grenzwertig seit letzten Serviceeinsatz und jetzt außer Bereich bei pO2 Level C und erneut Leckage Fehler->Überprüfung des internen Flüssigkeitstransportweges von Abfall bis Einlass, Sensor Bay geöffnet und LP herausgenommen->erneutes Einsetzen des LP sehr schwergängig->mittlere Kunststoffführung des LP am Analysator sitzt nicht korrekt->diese Führung wieder festgezogen, Leckagetest bestanden Startphase bestanden->QC bei pO2 in allen Leveln Zielwert | Fehlersuche, laut Frau Paul QC grenzwertig seit letzten Serviceeinsatz und jetzt außer Bereich bei pO2 Level C und erneut Leckage Fehler->Überprüfung des internen Flüssigkeitstransportweges von Abfall bis Einlass, Sensor Bay geöffnet und LP herausgenommen->erneutes Einsetzen des LP sehr schwergängig->mittlere Kunststoffführung des LP am Analysator sitzt nicht korrekt->diese Führung wieder festgezogen, Leckagetest bestanden Startphase bestanden->QC bei pO2 in allen Leveln Zielwert | SWO-184685 | No | Troubleshooting, according to Ms Paul QC borderline since last service use and now out of range at pO2 Level C and again leakage error->checking of the internal liquid transport path from waste to inlet, Sensor Bay opened and LP removed->again Inserting the LP very difficult-to-use->medium plastic guide of the LP on the analyzer sits incorrectly->this guide reattached, leakage test passed Start phase passed->QC at pO2 at all levels target value | Troubleshooting, according to Ms Paul QC borderline since last service use and now out of range at pO2 Level C and again leakage error->checking of the internal liquid transport path from waste to inlet, Sensor Bay opened and LP removed->again Inserting the LP very difficult-to-use->medium plastic guide of the LP on the analyzer sits incorrectly->this guide reattached, leakage test passed Start phase passed->QC at pO2 at all levels target value | Adjustment | HW | Basic Unit | SP Guide
ABL90 | Upon arrival, found the analyzer presenting error 964 Leak current in relation to Solution Pack detected. Also found, pO2 calibrations failing with error 1002- pO2 dark data is out of range. | Back-flush the analyzer and, replaced a new solution pack and, sensor cassette. Performed the complete pO2 check, found the pO2 dark value high outside of range 738.71 (Range: 10-200). Replaced a new sensor interface pcb and, ran the pO2 check, passing within specifications. 4 17 Followup visit. Found analyzer presenting calibration drift errors on lactate. The customer is requesting the analyzer replaced in light of the problems experienced. I will informa my manager and sales. | D264398EA45392B0 | No | Upon arrival, found the analyzer presenting error 964 Leak current in relation to Solution Pack detected. Also found, pO2 calibrations failing with error 1002- pO2 dark data is out of range. | Back-flush the analyzer and, replaced a new solution pack and, sensor cassette. Performed the complete pO2 check, found the pO2 dark value high outside of range 738.71 (Range: 10-200). Replaced a new sensor interface pcb and, ran the pO2 check, passing within specifications. 4 17 Followup visit. Found analyzer presenting calibration drift errors on lactate. The customer is requesting the analyzer replaced in light of the problems experienced. I will informa my manager and sales. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX PLUS | Replaced bent probe. Analyzer calibrated and passed all Qc's | Replaced bent probe. Analyzer calibrated and passed all Qc's | SWO-184844 | No | Replaced bent probe. Analyzer calibrated and passed all Qc's | Replaced bent probe. Analyzer calibrated and passed all Qc's | Unknown | HW | Auto Inlet module | Probe bend | ext. O2 Leaking
ABL90 FLEX PLUS | ZVBM aus Kl. Detmold geholt. Analysator geprüft, keinen Fehler festgestellt. Kapillar Modus standardmäßig auf 65ul gestellt, Kundin hat bei Probemessung die 65ul Kapillare nicht komplett gefüllt. Testmessung im 45ul Modus durchgeführt, Messung i.O. 45ul als Standart Modus eingestellt. Ersatz SC90 noch im Kühlschrank, Lösungspack und Einlassdichtung als Ersatz vor Ort gelassen. | ZVBM aus Kl. Detmold geholt. Analysator geprüft, keinen Fehler festgestellt. Kapillar Modus standardmäßig auf 65ul gestellt, Kundin hat bei Probemessung die 65ul Kapillare nicht komplett gefüllt. Testmessung im 45ul Modus durchgeführt, Messung i.O. 45ul als Standart Modus eingestellt. Ersatz SC90 noch im Kühlschrank, Lösungspack und Einlassdichtung als Ersatz vor Ort gelassen. | SWO-184768 | No | ZVBM from Kl. Detmold. Analyzer checked, no error detected. Capillary mode set to 65ul by default, customer did not completely fill the 65ul capillary when sampled. Test measurement performed in 45ul mode, measurement i.O. 45ul set as standard mode. Replacement SC90 still left in the refrigerator, solution pack and inlet seal as a replacement on site. | ZVBM from Kl. Detmold. Analyzer checked, no error detected. Capillary mode set to 65ul by default, customer did not completely fill the 65ul capillary when sampled. Test measurement performed in 45ul mode, measurement i.O. 45ul set as standard mode. Replacement SC90 still left in the refrigerator, solution pack and inlet seal as a replacement on site. | User error | Sample volume | Sample volume | Sample volume
AQT90 FLEX | 更换903-364后修复(By Ran) | 更换903-364后修复(By Ran) | SWO-184481 | No | Repair after replacement 903-364 (By Ran) | Repair after replacement 903-364 (By Ran) | Defect | HW | Process Module | Optical Unit | 131
ABL90 FLEX | The inlet gasket was lifted slightly and this then allowed the inlet to open. The inlet gasket was pressed along the probe, (as if introducing a sample) and an audible click was heard as the gasket sat home. The inlet now opened and closed acoordingly. The fluid transport tests were ran in the service program and all passed. The analyser was allowed to restart. The Sensor cassette was replaced due to the calibration and QC failures especially as solution has not been pumped through the analyser since the 13th April. The SC was left to start-up and stabilise. | The inlet gasket was lifted slightly and this then allowed the inlet to open. The inlet gasket was pressed along the probe, (as if introducing a sample) and an audible click was heard as the gasket sat home. The inlet now opened and closed acoordingly. The fluid transport tests were ran in the service program and all passed. The analyser was allowed to restart. The Sensor cassette was replaced due to the calibration and QC failures especially as solution has not been pumped through the analyser since the 13th April. The SC was left to start-up and stabilise. | SWO-184607 | No | The inlet gasket was lifted slightly and this then allowed the inlet to open. The inlet gasket was pressed along the probe, (as if introducing a sample) and an audible click was heard as the gasket sat home. The inlet now opened and closed acoordingly. The fluid transport tests were ran in the service program and all passed. The analyser was allowed to restart. The Sensor cassette was replaced due to the calibration and QC failures especially as solution has not been pumped through the analyser since the 13th April. The SC was left to start-up and stabilise. | The inlet gasket was lifted slightly and this then allowed the inlet to open. The inlet gasket was pressed along the probe, (as if introducing a sample) and an audible click was heard as the gasket sat home. The inlet now opened and closed acoordingly. The fluid transport tests were ran in the service program and all passed. The analyser was allowed to restart. The Sensor cassette was replaced due to the calibration and QC failures especially as solution has not been pumped through the analyser since the 13th April. The SC was left to start-up and stabilise. | Adjustment | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Inlet gasket removed and inspected, the gasket appeared to be deformed. New inlet gasket fitted to the analyser and the analyser restarted passing calibrations and QC. | Inlet gasket removed and inspected, the gasket appeared to be deformed. New inlet gasket fitted to the analyser and the analyser restarted passing calibrations and QC. | SWO-184608 | No | Inlet gasket removed and inspected, the gasket appeared to be deformed. New inlet gasket fitted to the analyser and the analyser restarted passing calibrations and QC. | Inlet gasket removed and inspected, the gasket appeared to be deformed. New inlet gasket fitted to the analyser and the analyser restarted passing calibrations and QC. | User error | HW | Inlet module | Inlet Gasket
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-184618 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-184636 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Lac
AQT90 FLEX | 更换636-549一个正常。 by andrew | 更换636-549一个正常。 by andrew | SWO-184603 | No | Replace 636-549 one normal. by andrew | Replace 636-549 one normal. by andrew | Defect | HW | Aspiration Module | Needle Sensor | 4794
ABL825 FLEX | Cals, QC and tHb check are all at or near mean values. | Cleaned and aligned oximeter. Set cal verification sample size to 195ul. | D266E46E9055D1C3 | No | Cals, QC and tHb check are all at or near mean values. | Cleaned and aligned oximeter. Set cal verification sample size to 195ul. | Setup | Sample volume | Sample volume | Sample volume
ABL825 FLEX | QC and tHb calibration and cal are all on mean values. | Cleaned and aligned oximeter. Results still low. Set cal verification sample size to 195ul. | D266B0E3905AD1BF | No | QC and tHb calibration and cal are all on mean values. | Cleaned and aligned oximeter. Results still low. Set cal verification sample size to 195ul. | Setup | Sample volume | Sample volume | Sample volume
ABL825 FLEX | Checked data logs. Found 581 failures and high MetHb values on QC. Calibration also failing with tHB drifting intermittently. | Reseat and adjust lamp unit. Replaced tube for heater. Adjusted halogen intensity. Reseat connectors on oxy control pcb. Ran tHB calibration in service mode with pass results. | D266CC849024CE00 | No | Checked data logs. Found 581 failures and high MetHb values on QC. Calibration also failing with tHB drifting intermittently. | Reseat and adjust lamp unit. Replaced tube for heater. Adjusted halogen intensity. Reseat connectors on oxy control pcb. Ran tHB calibration in service mode with pass results. | Adjustment | HW | Oxi module | Lamp unit | Halogen
ABL800 BASIC | Done touch screen calibration, Restart the analyzer, problem resolve. | Done touch screen calibration, Restart the analyzer, problem resolve. | SWO-186894 | No | Done touch screen calibration, Restart the analyzer, problem resolve. | Done touch screen calibration, Restart the analyzer, problem resolve. | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX | Remove incubator wheel and discard used cup from wheel, cleaned interior dust and calibrate incubator wheel, process wheel and punch out, problem resolve. | Remove incubator wheel and discard used cup from wheel, cleaned interior dust and calibrate incubator wheel, process wheel and punch out, problem resolve. | SWO-186895 | No | Remove incubator wheel and discard used cup from wheel, cleaned interior dust and calibrate incubator wheel, process wheel and punch out, problem resolve. | Remove incubator wheel and discard used cup from wheel, cleaned interior dust and calibrate incubator wheel, process wheel and punch out, problem resolve. | Cleaned | HW | Incubator module | Cup Sensor | calibrate39
AQT90 FLEX | Found storage wheel stuck and not moving. Analyser opened and cleaned for storage wheel, Punch-out sensor and incubatore wheel. Analyser serviced and calibrated properly. handed over in afine working condition. | Found storage wheel stuck and not moving. Analyser opened and cleaned for storage wheel, Punch-out sensor and incubatore wheel. Analyser serviced and calibrated properly. handed over in afine working condition. | SWO-184922 | No | Found storage wheel stuck and not moving. Analyser opened and cleaned for storage wheel, Punch-out sensor and incubatore wheel. Analyser serviced and calibrated properly. handed over in afine working condition. | Found storage wheel stuck and not moving. Analyser opened and cleaned for storage wheel, Punch-out sensor and incubatore wheel. Analyser serviced and calibrated properly. handed over in afine working condition. | Cleaned | HW | Storage Module | Storage wheel sensor
ABL80 CO-OX | The tHb sensitivity is failing causing internal QC to be Not Calculated. | Replaced hemolyzer. | D2681F13064DC0EF | No | The tHb sensitivity is failing causing internal QC to be Not Calculated. | Replaced hemolyzer.
ABL80 FLEX BASIC |  | B | SWO-184729 | No
ABL80 FLEX BASIC | Replace the Sensor Cassette with the new one,Electrolyte parameter got passed but PO2,PCO2,pH failed,Run the Zeroing check ,O2 was not stabilized,Then replace the sensor board,again run the zeroing check ,it got passed successfully,then install the sensor cassette successfully. | Replace the Sensor Cassette with the new one,Electrolyte parameter got passed but PO2,PCO2,pH failed,Run the Zeroing check ,O2 was not stabilized,Then replace the sensor board,again run the zeroing check ,it got passed successfully,then install the sensor cassette successfully. | SWO-185463 | No | Replace the Sensor Cassette with the new one,Electrolyte parameter got passed but PO2,PCO2,pH failed,Run the Zeroing check ,O2 was not stabilized,Then replace the sensor board,again run the zeroing check ,it got passed successfully,then install the sensor cassette successfully. | Replace the Sensor Cassette with the new one,Electrolyte parameter got passed but PO2,PCO2,pH failed,Run the Zeroing check ,O2 was not stabilized,Then replace the sensor board,again run the zeroing check ,it got passed successfully,then install the sensor cassette successfully.
ABL90 | ANNUAL PM DUE | ANNUAL PM COMPLETE | D267F97F063DF378 | Yes | ANNUAL PM DUE | ANNUAL PM COMPLETE
ABL90 FLEX PLUS |  | B | D2688B25060CEA3E | Yes
ABL90 FLEX PLUS | Split in pinch valve tubing, replaced tubing, now ok. | Split in pinch valve tubing, replaced tubing, now ok. | SWO-184793 | No | Split in pinch valve tubing, replaced tubing, now ok. | Split in pinch valve tubing, replaced tubing, now ok. | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX PLUS | Correctly inserted inter connecting Gasket and straightened slightly bent probe. | Correctly inserted inter connecting Gasket and straightened slightly bent probe. | SWO-185159 | No | Correctly inserted inter connecting Gasket and straightened slightly bent probe. | Correctly inserted inter connecting Gasket and straightened slightly bent probe. | Defect | HW | Auto Inlet module | Connector Gasket
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-185015 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | Leaking
AQT90 FLEX | 更换903-364一个正常。 by andrew | 更换903-364一个正常。 by andrew | SWO-184633 | No | Replace 903-364 one normal. by andrew | Replace 903-364 one normal. by andrew | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-198及903-215后修复 by andrew | 更换903-198及903-215后修复 by andrew | SWO-184616 | No | Repair by Andrew after replacing 903-198 and 903-215 | Repair by Andrew after replacing 903-198 and 903-215 | Defect | HW | Analyzer Control Module | PCB, Main Board | 1999
ABL825 FLEX | Checked carousel for debris. Checked to make sure customer had proper lot numbers installed etc.. | Replaced 14 pole cable. Restarted analyzer and reseated the carousel. Performed startup. | D267DE9606224D7B | No | Checked carousel for debris. Checked to make sure customer had proper lot numbers installed etc.. | Replaced 14 pole cable. Restarted analyzer and reseated the carousel. Performed startup. | Defect | HW | Autocheck module | FPC Cable
AQT90 FLEX | ウィルススキャン | B | SWO-188701 | Yes | Virus Scan |  | Re classified RMED | Virus scan
ABL80 FLEX Basic version | Verificação dos registos de dados. Passagem de 2 niveis de QC com resultados OK. Realização de gasimetria, o resultado de Na continua baixo. Execução de limpeza conforme procedimento. Verificação da impressora. Calibração e verificação com QC. | B | D2694AF67C21D4B8 | No | Verification of data records. Passage of 2 levels of QC with results OK. Achievement of Gasimetry, the result of Na is still low. Cleaning execution as per procedure. Check the printer. Calibration and verification with QC.
ABL80 FLEX BASIC | 更换902-972一个正常。 by andrew | 更换902-972一个正常。 by andrew | SWO-184658 | No | Replace 902-972 one normal. by andrew | Replace 902-972 one normal. by andrew
ABL80 FLEX BASIC |  | B | SWO-185042 | No
ABL80 FLEX BASIC | 更换试剂包后修复 | 更换试剂包后修复 | SWO-185361 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit
ABL80 | (18-4-2016) Verificación del fallo (se cambia la F.A, comprobación de la pantalla, bateria y módulo PC (siendo este la causa de la avería). Se solicita un módulo nuevo. (19-4-2016) Desmontaje y sustitución del módulo PC, verificación del fungible reutilizable, calibración y verifación con ciclos de calidad. | Cambio del módulo PC. | D26947277C4B1A7B | No | (18-4-2016) Fault check (changes F.A., screen check, battery and PC module (this being the cause of the breakdown). A new module is requested. (19-4-2016) Disassembly and replacement of the PC module, reusable fungible verification, calibration and verification with quality cycles. | Changing the PC module.
ABL90 FLEX | Netzteil wird bestellt | Netzteil wird bestellt | SWO-185285 | No | Power supply is ordered | Power supply is ordered | Revisit required | HW | Basic Unit | Power Supply
ABL90 FLEX PLUS | Fehlersuche, Fließwege gespült, Einlaßdichtung und -halter gew., Kal. + QC i.O. | Fehlersuche; Fließwege gespült; Einlaßdichtung und -halter gew.; Kal. + QC i.O. | SWO-185185 | No | Troubleshooting, flow paths rinsed, inlet seal and holder, Kal. + QC in the first place | Troubleshooting; Flow paths rinsed; inlet seal and holder. Kal. + QC in the first place | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | Remove inlet module. Remove probe and holder. Check for obstruction .Assemble inlet module, Fit inlet module and perform inlet calibration and analyzer calibration. All OK. | Remove inlet module. Remove probe and holder. Check for obstruction .Assemble inlet module, Fit inlet module and perform inlet calibration and analyzer calibration. All OK. | SWO-185479 | No | Remove inlet module. Remove probe and holder. Check for obstruction . Assemble inlet module, Fit inlet module and perform inlet calibration and analyzer calibration. All OK. | Remove inlet module. Remove probe and holder. Check for obstruction . Assemble inlet module, Fit inlet module and perform inlet calibration and analyzer calibration. All OK. | Defect | HW | Auto Inlet module | Auto Inlet module | User error?
ABL90 FLEX PLUS | Replaced all tubing's and inlet gasket but still fails Qc, Noticed heard oxygen leaking from the cubical next to the analyzer and the entrance to the cubical in less than a meter away, located leak and solved the issue of oxygen leaking. Left analyzer to stabilize and informed ward sister in charge and lab tech to check on the analyzer if it corrected the pO2 Qc issue, spoke to Paul the tech from lab on 19 04 2017 at 10H00 to move analyzer to were there is no oxygen contamination in the air. Awaiting his feedback with regards to the move if it corrected the pO2 issue. | Replaced all tubing's and inlet gasket but still fails Qc, Noticed heard oxygen leaking from the cubical next to the analyzer and the entrance to the cubical in less than a meter away, located leak and solved the issue of oxygen leaking. Left analyzer to stabilize and informed ward sister in charge and lab tech to check on the analyzer if it corrected the pO2 Qc issue, spoke to Paul the tech from lab on 19 04 2017 at 10H00 to move analyzer to were there is no oxygen contamination in the air. Awaiting his feedback with regards to the move if it corrected the pO2 issue. | SWO-185221 | No | Replaced all tubing's and inlet gasket but still fails Qc, Noticed heard oxygen leaking from the cubical next to the analyzer and the entrance to the cubical in less than a meter away, located leak and solved the issue of oxygen leaking. Left analyzer to stabilize and informed ward sister in charge and lab tech to check on the analyzer if it corrected the pO2 Qc issue, spoke to Paul the tech from lab on 19 04 2017 at 10H00 to move analyzer to were there is no oxygen contamination in the air. Awaiting his feedback with regards to the move if it corrected the pO2 issue. | Replaced all tubing's and inlet gasket but still fails Qc, Noticed heard oxygen leaking from the cubical next to the analyzer and the entrance to the cubical in less than a meter away, located leak and solved the issue of oxygen leaking. Left analyzer to stabilize and informed ward sister in charge and lab tech to check on the analyzer if it corrected the pO2 Qc issue, spoke to Paul the tech from lab on 19 04 2017 at 10H00 to move analyzer to were there is no oxygen contamination in the air. Awaiting his feedback with regards to the move if it corrected the pO2 issue. | User error | Inviroment | Inviroment | ext. O2 Leaking | ext. O2 Leaking
ABL90 FLEX | 更换910-342一个正常。by andrew | 更换910-342一个正常。by andrew | SWO-184707 | No | Replace 910-342 one normal. by andrew | Replace 910-342 one normal. by andrew | Defect | HW | DMS_PC | CPU
ABL90 FLEX | 更改903-581后修复 by andrew | 更改903-581后修复 by andrew | SWO-184662 | No | Fix by andrew after changing 903-581 | Fix by andrew after changing 903-581 | Defect | HW | Basic Unit | Sensor Interface PCB | 964
AQT90 FLEX | 更换备件924-591重新定标后修复 by andrew | 更换备件924-591重新定标后修复 by andrew | SWO-184671 | No | Repair by andrew after replacement of spare parts 924-591 re-calibration | Repair by andrew after replacement of spare parts 924-591 re-calibration | Defect | HW | Aspiration Module | Wet Section, Manifold | 3414
ABL800 BASIC | pH電極交換 | pH電極交換 | SWO-188681 | No | pH electrode replacement | pH electrode replacement | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Einlass tropft jeweils in letzter Sequenz der Spülzyklen erheblich. Sonde während Spülung untersucht, sie fährt offenbar etwas weiter heraus als üblich, Nadel hat längenmittig einen zurückgebogenen Knick => Nadel und Einlassdichtung ersetzt, mehrere Kalibrationen und Spülzyklen beobachtet, kein weiterer Flüssigkeitsaustritt feststellbar. Noch installierte Cl Elektrode durch Dummy ersetzt. Endkontrolle: QC Messung OK | Einlass tropft jeweils in letzter Sequenz der Spülzyklen erheblich. Sonde während Spülung untersucht, sie fährt offenbar etwas weiter heraus als üblich; Nadel hat längenmittig einen zurückgebogenen Knick => Nadel und Einlassdichtung ersetzt, mehrere Kalibrationen und Spülzyklen beobachtet, kein weiterer Flüssigkeitsaustritt feststellbar. Noch installierte Cl Elektrode durch Dummy ersetzt. Endkontrolle: QC Messung OK | SWO-185168 | No | Inlet drips significantly in the last sequence of the rinsing cycles. Probe examined during rinsing, it apparently goes a little further out than usual, needle has replaced a bent needle and inlet seal in the middle of the length, observed several calibrations and rinsing cycles, no further liquid leak can be detected. Still installed Cl electrode replaced by dummy. Final control: QC measurement OK | Inlet drips significantly in the last sequence of the rinsing cycles. Probe examined during flushing, it apparently travels a little further than usual; Needle has replaced a bent crease => needle and inlet seal in the middle of the length, observed several calibrations and rinsing cycles, no further liquid leak can be detected. Still installed Cl electrode replaced by dummy. Final control: QC measurement OK | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Alle relevanten Komponenten untersucht, Ventile im Service überprüft, Alles OK. Einlassnadel leicht gebogen - gerichtet. Messkammer auf Dichtheit geprüft, OK. pH Elektrode und Referenz ausgebaut, überprüft und klebenden pH Dichtring ersetzt. ServiceDump erstellt sowie Datenextrakt gezogen zur Weiterleitung an RDEU. Kalibrationen und QC Messung OK | Alle relevanten Komponenten untersucht, Ventile im Service überprüft, Alles OK. Einlassnadel leicht gebogen - gerichtet. Messkammer auf Dichtheit geprüft, OK. pH Elektrode und Referenz ausgebaut, überprüft und klebenden pH Dichtring ersetzt. ServiceDump erstellt sowie Datenextrakt gezogen zur Weiterleitung an RDEU. Kalibrationen und QC Messung OK | SWO-185192 | No | All relevant components examined, valves checked in service, all OK. Inlet needle slightly bent - directed. Measuring chamber tested for tightness, OK. pH electrode and reference removed, checked and adhesive pH sealing ring replaced. ServiceDump created and data extract drawn for forwarding to RDEU. Calibrations and QC measurement OK | All relevant components examined, valves checked in service, all OK. Inlet needle slightly bent - directed. Measuring chamber tested for tightness, OK. pH electrode and reference removed, checked and adhesive pH sealing ring replaced. ServiceDump created and data extract drawn for forwarding to RDEU. Calibrations and QC measurement OK | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL827 FLEX | CONFIRMED ISSUE WHEN ONSITE. | ADJUSTED pH BG MODULE SETTINGS AND PERFORMED MULTIPLE 2-PT CALS. PO2 IS NOW STABLE FOR 0.1mmHg LEVEL | D2696F207C4C43F1 | No | CONFIRMED ISSUE WHEN ONSITE. | ADJUSTED pH BG MODULE SETTINGS AND PERFORMED MULTIPLE 2-PT CALS. PO2 IS NOW STABLE FOR 0.1mmHg LEVEL | Adjustment | HW | pH/BG Module | pO2 | Zero adjustment
ABL815 FLEX |  | B | D26BACDB6819FA1E | No |  |  | Defect | Electrode | Electrode | Sealing ring
ABL835 FLEX | Messkammer, Konnektor überprüft, Elektrode neu membraniert, gesäubert und einkalibriert, Empf. bleibt wiederholt zu hoch, testweise gebrauchte Elektrode membraniert und eingesetzt, Verlauf wird beobachtet | B | SWO-185266 | No | Measuring chamber, connector checked, electrode re-membraneed, cleaned and calibrated, empf. remains repeatedly too high, test-used electrode membranes and used, course is observed |  | Assisted Replacement | Electrode | Membrane | pCO2
ABL80 FLEX BASIC |  | B | SWO-185320 | No
ABL80 FLEX BASIC | QCにてClが外れる件で、シグナルデータボード、センサーボード交換実施。ゼロチェック実施。様子見。 | QCにてClが外れる件で、シグナルデータボード、センサーボード交換実施。ゼロチェック実施。様子見。 | SWO-185292 | No | In the case where cl is removed in QC, the signal data board and the sensor board exchange were carried out. Zero check ed. Wait. | In the case where cl is removed in QC, the signal data board and the sensor board exchange were carried out. Zero check ed. Wait.
ABL90 FLEX | Einlassdichtung mit Halter defekt. Schwarze Dichtung hat sich nach vorn herausgeschossen. Garantie. | Einlassdichtung mit Halter defekt. Schwarze Dichtung hat sich nach vorn herausgeschossen. Garantie. | SWO-186620 | No | Inlet seal with holder defective. Black poetry has shot forward. Guarantee. | Inlet seal with holder defective. Black poetry has shot forward. Guarantee. | Defect | HW | Inlet module | Inlet Gasket
ABL90 | Decontamination and Planned Maintenance. | CONDUCTED DECONTAMINATION ON ANALYZER. REPLACED SENSOR CASSETTE AND SOLUTION PACK. Submitted Credit Claim forms. | D26ABB96F252A18E | No | Decontamination and Planned Maintenance. | CONDUCTED DECONTAMINATION ON ANALYZER. REPLACED SENSOR CASSETTE AND SOLUTION PACK. Submitted Credit Claim forms. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS | IP und Netzwerk konfiguriert | IP und Netzwerk konfiguriert | SWO-185582 | No | IP and network configured | IP and network configured | Setup | His/lis | His/lis | His/lis
ABL90 FLEX PLUS |  | B | D26AC2F6F237EB8B | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
TCM5 FLEX |  | B | D26A93D7F25DB96A | Yes |  |  | Re classified RMED | setup new analyzer
TCM5 FLEX |  | B | D26A9398F22DB969 | Yes |  |  | Re classified RMED | setup new analyzer
ABL825 FLEX | SUSTITUCIÓN DE HDD SATA CONFIGURACIÓN COMPLETA REVISIÓN PREVENTIVA SEGÚN PROCEDIMIENTO | SUSTITUCIÓN DE HDD SATA CONFIGURACIÓN COMPLETA REVISIÓN PREVENTIVA SEGÚN PROCEDIMIENTO | D26AD5A2F206919F | No | SATA HDD SUBSTITUTE FULL CONFIGURATION PREEMPTIVE REVIEW AS PROCEDURE | SATA HDD SUBSTITUTE FULL CONFIGURATION PREEMPTIVE REVIEW AS PROCEDURE | Defect | HW | DMS_PC | HDD
ABL835 FLEX | Revisión preventiva según procedimiento | Revisión preventiva según procedimiento | D26A8908F23B9198 | Yes | Preventive review as procedure | Preventive review as procedure | Re classified RMED | PM
ABL835 FLEX |  | B | D26BC404682ACD30 | No |  |  | Revisit required | HW | Auto Check Module | Fan
AQT90 FLEX | Einbau der Sicherungsringe an den Antriebsrädchen der Inlet Doors. Bestellnummer der Sicherungsringe:048-290. Sie sind nicht in CRM enthalten. | Einbau der Sicherungsringe an den Antriebsrädchen der Inlet Doors. Bestellnummer der Sicherungsringe:048-290. Sie sind nicht in CRM enthalten. | SWO-185574 | No | Installation of the safety rings on the drive wheel of the Inlet Doors. Order number of the fuse rings:048-290. They are not included in CRM. | Installation of the safety rings on the drive wheel of the Inlet Doors. Order number of the fuse rings:048-290. They are not included in CRM. | Defect | HW | Cover and Chassis | Inlet Door
AQT90 FLEX |  | B | D26AD054F21D80EB | No |  |  | Setup | His/lis | His/lis | His/lis
AQT90 FLEX | Nieuwe naald gemonteerd en afgesteld. Teller gereset. Systeemcontrole, OK. | Nieuwe naald gemonteerd en afgesteld. Teller gereset. Systeemcontrole, OK. | SWO-185409 | No | New needle mounted and adjusted. Counter reset. System Check, OK. | New needle mounted and adjusted. Counter reset. System Check, OK. | Defect | HW | Aspiration Module | Bend needle
AQT90 FLEX | neue Adressen eingegeben, Adressen abgespeichert, Datenübertragung kontrolliert, | neue Adressen eingegeben, Adressen abgespeichert, Datenübertragung kontrolliert, | SWO-185545 | No | new addresses, stored addresses, data transmission controlled, | new addresses, stored addresses, data transmission controlled, | Setup | His/lis | His/lis | His/lis
ABL80 FLEX BASIC | センサーボード交換。SC交換。 | B | SWO-186689 | No | Sensor board replacement. SC replacement.
ABL90 |  | B | D26BC2FD6812BCB2 | Yes
ABL90 FLEX | Glaskörper gereinigt und LP erneuert. Ohne Erfolg. SC erneuert. | Mehrfach kalibriert und QC Kontrollen durchgeführt. Alle Werte wieder grün. Gerät ist einsatzbereit. | SWO-186659 | Yes | Vitreous cleaned and LP renewed. Without success. SC renewed. | Multiple calibrated and QC checks performed. All values again green. Device is ready for use. | Re classified RMED | SC | SC | pCO2
ABL90 FLEX | Mehrfach kalibriert und QC Kontrollen durchgeführt. Alle Werte wieder grün. Gerät ist einsatzbereit. | Mehrfach kalibriert und QC Kontrollen durchgeführt. Alle Werte wieder grün. Gerät ist einsatzbereit. | SWO-186701 | No | Multiple calibrated and QC checks performed. All values again green. Device is ready for use. | Multiple calibrated and QC checks performed. All values again green. Device is ready for use. | Assisted Replacement | SC | SC | pCO2
ABL90 | Rebooted anlyzer and does not help | Reloade software, ran blood sample, the screen does not lock up | D26BD4A2685FCF68 | No | Rebooted anlyzer and does not help | Reloade software, ran blood sample, the screen does not lock up | Defect | SW | SW | Freeze
ABL90 FLEX | Fehlersuche, Lösungen können nicht gepumpt werden, ELD, Einlasssonde und Anschlussdichtung überprüft ok, SC und LP ausgeworfen, von Abfall zu linken Anschluss der SC gespült ok, SC gespült ok, von rechten Anschluss der SC zu Einlass gespült (schwergängig->nach mehrmaligem Spülen ok), nochmals komplette Spülung des Analysators durchgeführt, auf Serviceprogrammebene werden jetzt alle Testprogramme bestanden Startphase bestanden, System-und Funktionskontrolle i.O. | Fehlersuche, Lösungen können nicht gepumpt werden, ELD, Einlasssonde und Anschlussdichtung überprüft ok, SC und LP ausgeworfen, von Abfall zu linken Anschluss der SC gespült ok, SC gespült ok, von rechten Anschluss der SC zu Einlass gespült (schwergängig->nach mehrmaligem Spülen ok), nochmals komplette Spülung des Analysators durchgeführt, auf Serviceprogrammebene werden jetzt alle Testprogramme bestanden Startphase bestanden, System-und Funktionskontrolle i.O. | SWO-186855 | No | Troubleshooting, solutions cannot be pumped, ELD, inlet probe and connection seal checked ok, SC and LP ejected, from waste to left connection the SC flushed ok, SC flushed ok, from right connection of the SC to inlet flushed (difficult-to-go->after repeated Rinse ok), again complete flushing of the analyzer is carried out, at the service program level all test programs are now passed start-up phase passed, system and function control i.O. | Troubleshooting, solutions cannot be pumped, ELD, inlet probe and connection seal checked ok, SC and LP ejected, from waste to left connection the SC flushed ok, SC flushed ok, from right connection of the SC to inlet flushed (difficult-to-go->after repeated Rinse ok), again complete flushing of the analyzer is carried out, at the service program level all test programs are now passed start-up phase passed, system and function control i.O. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 | The customer is running qc study and, pO2 is failing on all qc levels. Recommended the customer to run waste blood every day and every time during the qc study so the sensor cassette is conditioned at all times for qc. Customer is requesting my assistance with this task. i will speak to Don Masters about the customer's request. | 4 19 17 Replaced a new sensor interface PCB. Found sensor cassette expired, replaced from customer's stock. 4 20 17 Revisit to check sensor cassette performance and, test with qc. Ran 7 blood samples and, ran qc. After 3 qc ampoules of each level, level 1 is low and 2 is low failing out of qc range. 4 21 17 Ran qc on both the old and new analyzers. The old analyzer failed again on levels 1 and 2. The new analyzer did not present any problems. | D26BFD3E6832D4B9 | No | The customer is running qc study and, pO2 is failing on all qc levels. Recommended the customer to run waste blood every day and every time during the qc study so the sensor cassette is conditioned at all times for qc. Customer is requesting my assistance with this task. i will speak to Don Masters about the customer's request. | 4 19 17 Replaced a new sensor interface PCB. Found sensor cassette expired, replaced from customer's stock. 4 20 17 Revisit to check sensor cassette performance and, test with qc. Ran 7 blood samples and, ran qc. After 3 qc ampoules of each level, level 1 is low and 2 is low failing out of qc range. 4 21 17 Ran qc on both the old and new analyzers. The old analyzer failed again on levels 1 and 2. The new analyzer did not present any problems. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX PLUS | Reloaded software now ok. | B | SWO-186796 | No | Reloaded software now ok. |  | Defect | SW | RMED SW | RMED SW | Service tip available ???
ABL90 FLEX PLUS | Gerät verlangt immer wieder Einlassdichtungswechsel. Mit Kunden das Wechseln der Einlassdichtung durchgeführt , Aktion bestätigt und Wartungsintervalle neu eingestellt. Systemtest durchgeführt, alles o.k. | Gerät verlangt immer wieder Einlassdichtungswechsel. Mit Kunden das Wechseln der Einlassdichtung durchgeführt , Aktion bestätigt und Wartungsintervalle neu eingestellt. Systemtest durchgeführt, alles o.k. | SWO-186836 | No | device repeatedly requires inlet seal changes. With customers changing the inlet seal performed, action confirmed and maintenance intervals re-set. System test performed, all ok. | device repeatedly requires inlet seal changes. With customers changing the inlet seal performed, action confirmed and maintenance intervals re-set. System test performed, all ok. | Assisted Replacement | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D26BE601683AF5B1 | No |  |  | Revisit required | SC | SC | Lac | Sensitivity
ABL815 FLEX | Klärung des Standortes des Gerätes nach Umzug Neubau, Klärung Zugang zum Gerät wg. Neubau, DemoOptionKey (gültig bis 21.7.2017) installiert, KAL & QK ok | Klärung des Standortes des Gerätes nach Umzug Neubau, Klärung Zugang zum Gerät wg. Neubau, DemoOptionKey (gültig bis 21.7.2017) installiert, KAL & QK ok | SWO-186780 | No | Clarification of the location of the device after moving new building, clarification access to the device wg. New building, DemoOptionKey (valid until 21.7.2017) installed, KAL & QK ok | Clarification of the location of the device after moving new building, clarification access to the device wg. New building, DemoOptionKey (valid until 21.7.2017) installed, KAL & QK ok | Setup | Demo key exp | Demo key exp | Demo key exp
ABL825 FLEX | 比較電極交換。 | 比較電極交換。 | SWO-186880 | No | Comparison electrode replacement. | Comparison electrode replacement. | Assisted Replacement | Electrode | Electrode | Reference
ABL835 FLEX |  | B | D26BCB9C6838CD33 | No |  |  | Defect | HW | Autocheck module | Autocheck module
ABL800 BASIC | 測定ソフトウェア再インストール | B | SWO-188425 | No | Re-install measurement software |  | Defect | SW | RMED SW | Start button
ABL800 BASIC | 測定ソフトウェア再インストール実施 | B | SWO-188418 | No | Reinstallation of measurement software |  | Defect | SW | RMED SW | Start button
ABL815 FLEX | Monitor komplett erneuert. Einlaufverfahren durchgeführt. | Monitor komplett erneuert. Einlaufverfahren durchgeführt. | SWO-186839 | No | Monitor completely renewed. run-in procedures. | Monitor completely renewed. run-in procedures. | Defect | HW | DMS_PC | Touch screen
ABL835 FLEX | Found a pin hole in one of the tubes. | Replaced all of the tubing withing the upper analyzer section. Ran a THb calibration, LS sensor adjust and pump calibration. All passed. Rouella will continue to monitor analyzer. The COHb is running low compared to peer values obtained from technical support for QC level2 (DEC 2016). Will contact support specialist if the resulst continue to trend downward. | D26BD4DA6853A080 | No | Found a pin hole in one of the tubes. | Replaced all of the tubing withing the upper analyzer section. Ran a THb calibration, LS sensor adjust and pump calibration. All passed. Rouella will continue to monitor analyzer. The COHb is running low compared to peer values obtained from technical support for QC level2 (DEC 2016). Will contact support specialist if the resulst continue to trend downward. | Defect | HW | Oxi module | Tubings
AQT90 FLEX | Performed the service programs, wet section problem with P4 and P5 was failed, replaced the wash block and cleaned the filter, restarted instrument but wash block was not responding, found problem with inlet PCB need to replace with new. Waiting for spares. | B | SWO-186931 | No | Performed the service programs, wet section problem with P4 and P5 was failed, replaced the wash block and cleaned the filter, restarted instrument but wash block was not responding, found problem with inlet PCB need to replace with new. Waiting for spares. |  | Defect | HW | Aspiration Module | PCB, Inlet | FTC 13
ABL80 FLEX BASIC | 更换903-381一个正常。by andrew | 更换903-381一个正常。by andrew | SWO-185609 | No | Replace 903-381 one normal. by andrew | Replace 903-381 one normal. by andrew
ABL80 FLEX BASIC | 更换903-001一个正常。 by andrew | 更换903-001一个正常。 by andrew | SWO-185421 | No | Replace 903-001 one normal. by andrew | Replace 903-001 one normal. by andrew
ABL90 FLEX PLUS |  | B | D26F50E4CA627EE9 | No |  |  | Replaced | Analyzer replaced | Analyzer replaced | Lac | Sensitivity
AQT90 FLEX | 重新定标正常。by andrew | 重新定标正常。by andrew | SWO-186621 | No | Recalibration is normal. by andrew | Recalibration is normal. by andrew | Assisted Replacement | Consumables | Test Kit | D-Dimer | 154
ABL90 FLEX | 更换903-199后修复(by tim) | B | SWO-186908 | No | Repair after replacement 903-199 (by tim) |  | Defect | HW | Oxi module | Hemolyzer | Clot
AQT90 FLEX | 更换903-210一个正常。by andrew | 更换903-210一个正常。by andrew | SWO-186886 | No | Replace the 903-210 one normal. by andrew | Replace the 903-210 one normal. by andrew | Defect | HW | Process Module | PCB, Cup dryer | 168
ABL80 CO-OX | Phone support. Customer reporting leaking from to solution packs. | Sent customer another solution pack overnight to see if it is leaking. It was the solution packs that were leaking. Problem went away with new solution pack. | D26F9C9CCA0E3B33 | Yes | Phone support. Customer reporting leaking from to solution packs. | Sent customer another solution pack overnight to see if it is leaking. It was the solution packs that were leaking. Problem went away with new solution pack.
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-187642 | Yes | Replace the test card | Replace the test card
ABL90 | Verified. | Replaced barcode reader. | D26FA2E2CA594A37 | No | Verified. | Replaced barcode reader. | Defect | HW | Basic Unit | Barcode reader
ABL90 | Verified. | Replaced barcode reader. | D26FA489CA4E4A39 | No | Verified. | Replaced barcode reader. | Defect | HW | Basic Unit | Barcode reader
ABL90 FLEX PLUS | Oxi - Schlauch ausgetauscht. Einlassdichtung aus dem Kundenbestand ausgetauscht. Neustart. | Oxi - Schlauch ausgetauscht. Einlassdichtung aus dem Kundenbestand ausgetauscht. Neustart. | SWO-187548 | No | Oxi - Hose replaced. Inlet seal from the customer stock. Restart. | Oxi - Hose replaced. Inlet seal from the customer stock. Restart. | Defect | HW | Oxi module | Pinch valve tube
ABL800 FLEX | Try to calibrate. Reinstall the membrane and try again. But not calibrating lactate parameter. Need to replace Lactate membrane. Asked from indore will be expecting tomorrow. | Lactate membrane replaced. Analyser serviced and calibrated properly. Handed over in a fine working condition. Electrolyte and Crea Glucose profile created as requested by staff. | SWO-188576 | No | Try to calibrate. Reinstall the membrane and try again. But not calibrating lactate parameter. Need to replace Lactate membrane. Asked from indore will be expecting tomorrow. | Lactate membrane replaced. Analyser serviced and calibrated properly. Handed over in a fine working condition. Electrolyte and Crea Glucose profile created as requested by staff. | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | Research data logs. Checked inlet flow. | Reseated inlet probe and tubing. Cleaned electrode fluid from pH measuring chamber and pH electrode. Performed Zero current adjust in diagnostics Restart calibrations. Ran a few patient samples with no errors. | D2702597CA62B21D | No | Research data logs. Checked inlet flow. | Reseated inlet probe and tubing. Cleaned electrode fluid from pH measuring chamber and pH electrode. Performed Zero current adjust in diagnostics Restart calibrations. Ran a few patient samples with no errors. | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL825 FLEX | Lichtfilter El Met ist nicht eingebaut. Provisorisch abgedunkelt. Lichtfilter bei RMED bestellt. | Lichtfilter El Met ist nicht eingebaut. Provisorisch abgedunkelt. Lichtfilter bei RMED bestellt. | SWO-187180 | No | Light filter El Met is not installed. Provisionally darkened. light filter ordered from RMED. | Light filter El Met is not installed. Provisionally darkened. light filter ordered from RMED. | Defect | HW | El/Met Module | Kapton filter | missing
ABL80 OSM | Checked data logs. Found results with high COHGB. Reran whole blood control, it is still running 4.5. The range is 2.1 to 3.9 | Replaced hemolyzer. Ran tHB calibration. Reran whole blood control with a result of 1.9. It is below the range established by running multiple samples on the other OSM on April 19th. | D27119CE4045F04F | Yes | Checked data logs. Found results with high COHGB. Reran whole blood control, it is still running 4.5. The range is 2.1 to 3.9 | Replaced hemolyzer. Ran tHB calibration. Reran whole blood control with a result of 1.9. It is below the range established by running multiple samples on the other OSM on April 19th.
ABL80 CO-OX | The latch that holds the SP door hinge assemby closed is broken. | Replaced molded door hinge assembly. | D2709FA0401AC107 | No | The latch that holds the SP door hinge assemby closed is broken. | Replaced molded door hinge assembly.
ABL90 FLEX | Einlassdichtung überprüft.ist neu.. Nadel nicht Verbogen. Gerät überprüft gespült. Stenose Gerinnsel freigespült. Pump..LS ...Pressure Test wird bestanden. Initialisierung. Gerät überprüft arbeitet, aber PCO2 Glu.Lac sind durch stillstand nicht ok. Neue SC ist nicht vorhanden, kein Wechsel möglich. | Einlassdichtung überprüft.ist neu.. Nadel nicht Verbogen. Gerät überprüft gespült. Stenose Gerinnsel freigespült. Pump..LS ...Pressure Test wird bestanden. Initialisierung. Gerät überprüft arbeitet, aber PCO2 Glu.Lac sind durch stillstand nicht ok. Neue SC ist nicht vorhanden, kein Wechsel möglich. | SWO-187290 | No | Inlet seal checked.is new.. Needle not bent. Device checked flushed. Stenosis clots rinsed free. Pump.. Ls... Pressure test is passed. Initialization. Device checked works, but PCO2 Glu.Lac are not ok by standstill. New SC is not available, no change is possible. | Inlet seal checked.is new.. Needle not bent. Device checked flushed. Stenosis clots rinsed free. Pump.. Ls... Pressure test is passed. Initialization. Device checked works, but PCO2 Glu.Lac are not ok by standstill. New SC is not available, no change is possible. | Revisit required | SC | SC | Glu/Lac
ABL90 | RAN SEVERAL SAMPLES WITH REOCCURRING RESULTS | REPLACED HEMOLYZER AND COMPLETED ADJUSTMENTS | D2722C1AB61A2BC4 | Yes | RAN SEVERAL SAMPLES WITH REOCCURRING RESULTS | REPLACED HEMOLYZER AND COMPLETED ADJUSTMENTS
ABL90 FLEX PLUS | De afvoer Inlet naar SC is verstopt. De blokkage heb verwijderd, OK. | De afvoer Inlet naar SC is verstopt. De blokkage heb verwijderd, OK. | SWO-187480 | No | The drain Inlet to SC is clogged. Have removed the block, OK. | The drain Inlet to SC is clogged. Have removed the block, OK. | Cleaned | Fluidic pathway | Fluidic pathway | Clot
ABL90 FLEX PLUS |  | B | D270DB09400FE37D | No |  |  | Defect | HW | Auto Inlet module | Auto Inlet module
AQT90 FLEX | 更换测试卡 | 更换测试卡 | SWO-187394 | No | Replace the test card | Replace the test card | Defect | Consumables | Test Kit | Cartridge | 5784
ABL800 FLEX | Lactate membrane replaced. Analyser serviced and calibrated properly. Handed over in a fine working condition. Electrolyte and Crea Glucose profile created as requested by staff. | Lactate membrane replaced. Analyser serviced and calibrated properly. Handed over in a fine working condition. Electrolyte and Crea Glucose profile created as requested by staff. | SWO-188581 | No | Lactate membrane replaced. Analyser serviced and calibrated properly. Handed over in a fine working condition. Electrolyte and Crea Glucose profile created as requested by staff. | Lactate membrane replaced. Analyser serviced and calibrated properly. Handed over in a fine working condition. Electrolyte and Crea Glucose profile created as requested by staff. | Assisted Replacement | Electrode | Membrane | Lac
ABL800 BASIC | Cleaned gasket with hypo chloride solution and also cleaned waste path. Then given pump calibration. Unit is working ok. | Cleaned gasket with hypo chloride solution and also cleaned waste path. Then given pump calibration. Unit is working ok. | SWO-188196 | No | Cleaned gasket with hypo chloride solution and also cleaned waste path. Then given pump calibration. Unit is working ok. | Cleaned gasket with hypo chloride solution and also cleaned waste path. Then given pump calibration. Unit is working ok. | Cleaned | HW | Inlet module | Inlet Gasket
ABL800 BASIC | インレットプローブ曲がり修正 エアライン清掃 LS調整実施 CAL・QC測定データ確認 | インレットプローブ曲がり修正 エアライン清掃 LS調整実施 CAL・QC測定データ確認 | SWO-187947 | No | Inlet probe bending fix Airline Cleaning Ls adjustment implemented Cal/QC measurement data confirmation | Inlet probe bending fix Airline Cleaning Ls adjustment implemented Cal/QC measurement data confirmation | Defect | HW | Inlet module | Probe bend | 663
ABL825 FLEX | No sample detected errors. Also some QC logging as unknowns. | FPC cable misrouted. Replaced FPC cable. Checked autocheck module alignment. Replaced pump tubings. | D270CFF1403222DB | No | No sample detected errors. Also some QC logging as unknowns. | FPC cable misrouted. Replaced FPC cable. Checked autocheck module alignment. Replaced pump tubings. | Defect | HW | Autocheck module | FPC Cable
ABL825 FLEX | oxi spectrum mismatch on samples | ran complete oxi test cavitaion current failed replaced hemolyzer unit | D270C8FA400FF6E4 | No | oxi spectrum mismatch on samples | Ran complete Oxi test cavitaion current failed replaced Hemolyzer unit | Defect | HW | Oxi module | Hemolyzer
AQT90 FLEX | ServiceDump erstellt und analysiert. Letzte 5Flows unter 1600. Analysator geöffnet. Pumpe P2 und P4 ausgetauscht. Schläuche verschließen. Verunreinigte Cup Wash Unit bei Halbjahreswartung gereinigt. QC durchgefallen. Parameter neu kalibriert. Weiterhin fehlerhafte QC. Kunde wird erneut angefahren Im Anschluss Halbjahreswartung | Durchführung einer Jahreswartung Maßnahme. Zusätzlich Erneuerung des Pumpenkopfes P2 und P5. Durchführung von Blankmessungen. Messwerte in Ordnung innerhalb der Grenzwerte. Überprüfung der Flowwerte, Werte liegen nun unter dem Wert von 1300. Sind nun niedriger, vorher bei über 1400. Qualitätskontrollen in Ordnung, zwischen oberem Grenzwert und Zielwert. Service Dump gezogen zur Ausarbeitung an Technischen Support. Entscheidung, ob noch Hardware getauscht wird. | SWO-187574 | No | ServiceDump creates and analyzes. Last 5Flows under 1600. Analyzer open. pump P2 and P4. Hoses. Contaminated Cup Wash Unit cleaned at half-year maintenance. QC. parameter recalibrated. Further faulty QC. Customer is approached again Following half-year maintenance | Implementation of an annual maintenance measure. In addition, renewal of the pump head P2 and P5. Execution of blank measurements. Measured values in order within the limit values. Checking the flow values, values are now below the value of 1300. Are now lower, previously at over 1400. Quality controls are fine, between the upper limit value and the target value. Service Dump pulled to technical support for elaboration. decide whether to replace hardware. | Performance | HW | Process Module | Tubings
ABL80 OSM | Checked data logs, observed the out of range values. Customers whole blood control is running 4.1. | Replaced hemolyzer. Ran tHb calibration. Customers whole blood control result is now running 1.6. | D272512DB607791B | Yes | Checked data logs, observed the out of range values. Customers whole blood control is running 4.1. | Replaced hemolyzer. Ran tHb calibration. Customers whole blood control result is now running 1.6.
ABL80 OSM | The analyzer could not find the remote support key. | I cleared the remote support key out and reloaded it. | D272368AB661C10E | No | The analyzer could not find the remote support key. | I cleared the remote support key out and reloaded it.
ABL80 | Checked the data logs and verrified the problem. The customer used 24 ampuoles of level 1 QC over past few days in trying to get the analyzer working. | Performed a zero cal. Replaced the solution pack. | D271FCFDB632A3B4 | No | Checked the data logs and verrified the problem. The customer used 24 ampuoles of level 1 QC over past few days in trying to get the analyzer working. | Performed a zero cal. Replaced the solution pack.
ABL90 FLEX | Ventil schwergängig, erneuert Testmessung durchgeführt i.O | Ventil schwergängig, erneuert Testmessung durchgeführt i.O | SWO-187720 | No | Valve heavy-duty, renewed Test measurement carried out in i.O | Valve heavy-duty, renewed Test measurement carried out in i.O | Defect | HW | Oxi module | Pinch valve
ABL90 | REPLACED ANDF TESTED WITH NEW ANNUAL TUBING, AND SPECTROMETER. TRY WET SECTION PCB | REPLACED WET SECTION PCB, SPECTROMETER AND COMPLETED ALL CALIBRATIONS AND UPGRADED TO SOFTWARE 3.2 | D2722F2FB6152BCB | Yes | REPLACED ANDF TESTED WITH NEW ANNUAL TUBING, AND SPECTROMETER. TRY WET SECTION PCB | REPLACED WET SECTION PCB, SPECTROMETER AND COMPLETED ALL CALIBRATIONS AND UPGRADED TO SOFTWARE 3.2
ABL90 FLEX | Removed SP and activated and re inserted into analyser Carried out fluid transport checks within service program to get analyser back in ready mode | Removed SP and activated and re inserted into analyser Carried out fluid transport checks within service program to get analyser back in ready mode | SWO-187916 | No | Removed SP and activated and re inserted into analyser Carried out fluid transport checks within service program to get analyser back in ready mode | Removed SP and activated and re inserted into analyser Carried out fluid transport checks within service program to get analyser back in ready mode | User error | SP | SP Activation | SP Activation
ABL90 FLEX | power supply erneuert, Teile zurücckgetauscht | power supply erneuert, Teile zurücckgetauscht | SWO-187650 | No | power supply renewed, parts exchanged | power supply renewed, parts exchanged | Defect | HW | Basic Unit | Power Supply
ABL90 FLEX | Gerät zerlegt, Austausch aller Schlauchverbindungen, Quetschventil im Oxy Unit geprüft i.O., Sensor Bay Unit, Hemolyzer Unit & Einslassnadel gespült, Befestigung an unterteil des Oxy Unit abgebrochen ->ersetzt div. Serviceprogramme durchgeführt (Fluidtransport, kpl. Oxy & pO2 Test) fehlerfrei, Kal & QC | Gerät zerlegt, Austausch aller Schlauchverbindungen, Quetschventil im Oxy Unit geprüft i.O., Sensor Bay Unit, Hemolyzer Unit & Einslassnadel gespült, Befestigung an unterteil des Oxy Unit abgebrochen ->ersetzt div. Serviceprogramme durchgeführt (Fluidtransport, kpl. Oxy & pO2 Test) fehlerfrei, Kal & QC | SWO-187812 | No | Device disassembled, replacement of all hose connections, squeezing valve in the Oxy Unit tested i.O., Sensor Bay Unit, Hemolyzer Unit & inlet needle flushed, attachment to lower part of the Oxy Unit aborted ->replaced various service programs (fluid transport, cpl. Oxy & pO2 Test) error-free, Kal & QC | Device disassembled, replacement of all hose connections, squeezing valve in the Oxy Unit tested i.O., Sensor Bay Unit, Hemolyzer Unit & inlet needle flushed, attachment to lower part of the Oxy Unit aborted ->replaced various service programs (fluid transport, cpl. Oxy & pO2 Test) error-free, Kal & QC | Defect | HW | Oxi module | Hemolyzer
ABL90 | lights and alarms black screen | na | D272263FB638865B | No | lights and alarms black screen | In | Revisit required | HW | DMS_PC | CPU
ABL90 FLEX | Datensicherung durchgeführt, Software neu installiert, Datensicherung wieder eingespielt, Patientenlayout nach Kundenwunsch angepasst, Referenzbereiche eingetragen, 1 std Startphase durchgeführt, Einstellungen nach Kundenwunsch geändert. | Datensicherung durchgeführt, Software neu installiert, Datensicherung wieder eingespielt, Patientenlayout nach Kundenwunsch angepasst, Referenzbereiche eingetragen, 1 std Startphase durchgeführt, Einstellungen nach Kundenwunsch geändert. | SWO-187826 | No | Data backup performed, software reinstalled, data backup re-recorded, patient layout adapted to customer requirements, reference areas entered, 1 std start phase performed, settings changed according to customer request. | Data backup performed, software reinstalled, data backup re-recorded, patient layout adapted to customer requirements, reference areas entered, 1 std start phase performed, settings changed according to customer request. | Defect | SW | Patient layout | Patient layout
ABL90 FLEX PLUS |  | B | D2722DB5B65E2F5D | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
AQT90 FLEX | 更换903-198后修复(by tin) | 更换903-198后修复(by tin) | SWO-187058 | No | Repair after replacement 903-198 (by tin) | Repair after replacement 903-198 (by tin) | Defect | HW | Analyzer Control Module | PCB, Main Board | 5766
ABL800 BASIC | 更换902-833后修复(by tim) | 更换902-833后修复(by tim) | SWO-187279 | No | Repair after replacement 902-833 (by tim) | Repair after replacement 902-833 (by tim) | Defect | HW | Oxi module | Oxi module | 682
ABL825 FLEX | 更换902-822一个正常。(by tim) | 更换902-822一个正常。(by tim) | SWO-187324 | No | Replace 902-822 one normal. (by tim) | Replace 902-822 one normal. (by tim) | Defect | HW | Basic Unit | Barcode reader, Basic Unit
ABL837 FLEX | pH eelctrode chamber has fluid build up intermittently. | Cleaned pH electrode chamber. Checked all valves. Replaced inlet nipple and pH o-ring. Replaced pump tubings. 6 Month PM done. | D2723B50B6428B15 | No | pH eelctrode chamber has fluid build up intermittently. | Cleaned pH electrode chamber. Checked all valves. Replaced inlet nipple and pH o-ring. Replaced pump tubings. 6 Month PM done. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
AQT90 FLEX | Found All Sample measurements are Positive, Cleaned the process wheel, and cup wash unit but problem was remaining same need to replace with process wheel cup wash unit and OU. | Found All Sample measurements are Positive, Cleaned the process wheel, and cup wash unit but problem was remaining same need to replace with process wheel cup wash unit and OU. | SWO-187793 | No | Found All Sample measurements are Positive, Cleaned the process wheel, and cup wash unit but problem was remaining same need to replace with process wheel cup wash unit and OU. | Found All Sample measurements are Positive, Cleaned the process wheel, and cup wash unit but problem was remaining same need to replace with process wheel cup wash unit and OU. | Cleaned | HW | Process Module | Process wheel
AQT90 FLEX | Abarbeitung des Data Collection Guide Analysatortechnisch. ServiceDump analysiert, Flow ohne Beanstandung Cup Wash Unit versehen mit leichter Kratzspur, CWU zerlegt und gereinigt, freier Lauf und einwandfreie Funktion. Nadelkolben gereinigt, leichte Verzögerung im zurückgleiten. Nadel ungewöhnlich stumpf (Counter 3000 pierces), erneuert und Counter resetet. Weitere Überprüfungen des DCG ohne Beanstandung. Blankmessungen und Flowanalyse i.O. DataLogs gesichert. Vorfall: Kist am 20.04.2017 12:47 | Abarbeitung des Data Collection Guide Analysatortechnisch. ServiceDump analysiert, Flow ohne Beanstandung Cup Wash Unit versehen mit leichter Kratzspur, CWU zerlegt und gereinigt, freier Lauf und einwandfreie Funktion. Nadelkolben gereinigt, leichte Verzögerung im zurückgleiten. Nadel ungewöhnlich stumpf (Counter 3000 pierces), erneuert und Counter resetet. Weitere Überprüfungen des DCG ohne Beanstandung. Blankmessungen und Flowanalyse i.O. DataLogs gesichert. Vorfall: Kist am 20.04.2017 12:47 | SWO-187879 | No | Processing of the Data Collection Guide analytical technical. ServiceDump analyzed, flow without complaint Cup Wash Unit provided with light scratch track, CWU disassembled and cleaned, free running and flawless function. Needle piston cleaned, slight delay in sliding back. Needle unusually blunt (Counter 3000 pierces), renewed and counter resets. Further checks of the DCG without complaint. Blank measurements and flow analysis in the first place DataLogs. Incident: Kist on 20.04.2017 12:47 | Processing of the Data Collection Guide analytical technical. ServiceDump analyzed, flow without complaint Cup Wash Unit provided with light scratch track, CWU disassembled and cleaned, free running and flawless function. Needle piston cleaned, slight delay in sliding back. Needle unusually blunt (Counter 3000 pierces), renewed and counter resets. Further checks of the DCG without complaint. Blank measurements and flow analysis in the first place DataLogs. Incident: Kist on 20.04.2017 12:47 | Cleaned | HW | Aspiration Module | Needle assembly
AQT90 FLEX |  | B | D27267EBB6442F61 | No |  |  | Defect | HW | Aspiration Module | Tube holder
AQT90 FLEX | Durchführung einer Jahreswartung Maßnahme. Zusätzlich Erneuerung des Pumpenkopfes P2 und P5. Durchführung von Blankmessungen. Messwerte in Ordnung innerhalb der Grenzwerte. Überprüfung der Flowwerte, Werte liegen nun unter dem Wert von 1300. Sind nun niedriger, vorher bei über 1400. Qualitätskontrollen in Ordnung, zwischen oberem Grenzwert und Zielwert. Service Dump gezogen zur Ausarbeitung an Technischen Support. Entscheidung, ob noch Hardware getauscht wird. | Durchführung einer Jahreswartung Maßnahme. Zusätzlich Erneuerung des Pumpenkopfes P2 und P5. Durchführung von Blankmessungen. Messwerte in Ordnung innerhalb der Grenzwerte. Überprüfung der Flowwerte, Werte liegen nun unter dem Wert von 1300. Sind nun niedriger, vorher bei über 1400. Qualitätskontrollen in Ordnung, zwischen oberem Grenzwert und Zielwert. Service Dump gezogen zur Ausarbeitung an Technischen Support. Entscheidung, ob noch Hardware getauscht wird. | SWO-187911 | Yes | Implementation of an annual maintenance measure. In addition, renewal of the pump head P2 and P5. Execution of blank measurements. Measured values in order within the limit values. Checking the flow values, values are now below the value of 1300. Are now lower, previously at over 1400. Quality controls are fine, between the upper limit value and the target value. Service Dump pulled to technical support for elaboration. decide whether to replace hardware. | Implementation of an annual maintenance measure. In addition, renewal of the pump head P2 and P5. Execution of blank measurements. Measured values in order within the limit values. Checking the flow values, values are now below the value of 1300. Are now lower, previously at over 1400. Quality controls are fine, between the upper limit value and the target value. Service Dump pulled to technical support for elaboration. decide whether to replace hardware. | Re classified RMED | Double entry
ABL80 FLEX BASIC | 更换933-148后修复(by tim) | 更换933-148后修复(by tim) | SWO-187638 | No | Repair after replacement 933-148 (by tim) | Repair after replacement 933-148 (by tim)
ABL80 CO-OX |  | B | D2738F952C5E009B | No
ABL80 FLEX BASIC | 更换902-972后修复(by tim) | 更换902-972后修复(by tim) | SWO-187430 | No | Repair after replacement 902-972 (by tim) | Repair after replacement 902-972 (by tim)
ABL90 FLEX | Backed up setup files in explorer Installed fresh database Loaded setup files | Backed up setup files in explorer Installed fresh database Loaded setup files | SWO-188204 | No | Backed up setup files in explorer Installed fresh database Loaded setup files | Backed up setup files in explorer Installed fresh database Loaded setup files | Defect | SW | Windows | Windows
ABL90 | at power up speaker beeps barcode active, lights on | replaced cpu board,or cassette reads exceeds life in instrument yet put on 10 days prior to computer crash | D27351282C111658 | No | at power up speaker beeps barcode active, lights on | replaced cpu board,or cassette reads exceeds life in instrument yet put on 10 days prior to computer crash | Unknown | HW | DMS_PC | CPU
ABL90 FLEX | 更换580-083后修复(by tim) | 更换580-083后修复(by tim) | SWO-187748 | No | Repair after replacement 580-083 (by tim) | Repair after replacement 580-083 (by tim) | Defect | HW | Basic Unit | SC Door
ABL800 FLEX | Replace the Crea reagent. | Replace the Crea reagent. | SWO-188573 | No | Replace the Crea reagent. | Replace the Crea reagent. | User error | Consumables | Consumables | Crea expired
ABL800 BASIC | Removed the inlet casket, dipped it in sodium hypochloride solution in order to remove the blood, it was found to be blocked. Run protein removal after connecting the inlet casket. Told the technicians about the protein removal process. Run 2 point calibration, run few samples machine is found working fine. | Removed the inlet casket, dipped it in sodium hypochloride solution in order to remove the blood, it was found to be blocked. Run protein removal after connecting the inlet casket. Told the technicians about the protein removal process. Run 2 point calibration, run few samples machine is found working fine. | SWO-188676 | No | Removed the inlet casket, dipped it in sodium hypochloride solution in order to remove the blood, it was found to be blocked. Run protein removal after connecting the inlet casket. Told the technicians about the protein removal process. Run 2 point calibration, run few samples machine is found working fine. | Removed the inlet casket, dipped it in sodium hypochloride solution in order to remove the blood, it was found to be blocked. Run protein removal after connecting the inlet casket. Told the technicians about the protein removal process. Run 2 point calibration, run few samples machine is found working fine. | Cleaned | HW | Inlet module | Inlet Gasket | Clot
ABL820 FLEX | pCO2 high on low QC compared to peer group. | Replaced pH O-Ring. Replaced pCO2 membrane and cleaned measuring chamber. | D273665C2C10D536 | No | pCO2 high on low QC compared to peer group. | Replaced pH O-Ring. Replaced pCO2 membrane and cleaned measuring chamber. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL800 BASIC | Found Lamp intensity low, performed the complete Oximetry test, found Lamp adjustment was not OK, performed the Lamp adjustment performed the calibration found ctHb values are OK, performed the test found ctHb and SO2 values are OK. Found Instrument working satisfactorily. | Found Lamp intensity low, performed the complete Oximetry test, found Lamp adjustment was not OK, performed the Lamp adjustment performed the calibration found ctHb values are OK, performed the test found ctHb and SO2 values are OK. Found Instrument working satisfactorily. | SWO-188542 | No | Found Lamp intensity low, performed the complete Oximetry test, found Lamp adjustment was not OK, performed the Lamp adjustment performed the calibration found ctHb values are OK, performed the test found ctHb and SO2 values are OK. Found Instrument working satisfactorily. | Found Lamp intensity low, performed the complete Oximetry test, found Lamp adjustment was not OK, performed the Lamp adjustment performed the calibration found ctHb values are OK, performed the test found ctHb and SO2 values are OK. Found Instrument working satisfactorily. | Adjustment | HW | Oxi module | Lamp unit | Low intensity
ABL825 FLEX | Examined QC data logs. | Replaced inlet and pH BG module, provided by depot (Andy Siwerka) inlet module S N: 783R0792N002, pH BG module: 823R0967N006. Verified operation of ABL825. | D2735F392C0EE404 | No | Examined QC data logs. | Replaced inlet and pH BG module, provided by depot (Andy Siwerka) inlet module S N: 783R0792N002, pH BG module: 823R0967N006. Verified operation of ABL825. | Defect | HW | Inlet module | Inlet module
ABL80 OSM | QC leaking from front of unit during system cycles. | Reposition waste tubing. Replaced solution pack. Cleaned QC off of external covers. | D2752CEFA203CBC8 | No | QC leaking from front of unit during system cycles. | Reposition waste tubing. Replaced solution pack. Cleaned QC off of external covers.
ABL90 FLEX | Replaced bent inlet probe and damaged split inlet gasket. | Replaced bent inlet probe and damaged split inlet gasket. | SWO-188537 | No | Replaced bent inlet probe and damaged split inlet gasket. | Replaced bent inlet probe and damaged split inlet gasket. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Replaced printer. | Replaced printer. | SWO-188465 | No | Replaced printer. | Replaced printer. | Defect | HW | Printer Module | Printer
ABL90 FLEX PLUS | Clutch Schrauben nachgezogen, Probelauf ok | Clutch Schrauben nachgezogen, Probelauf ok | SWO-188443 | No | Clutch screws tightened, test run ok | Clutch screws tightened, test run ok | Adjustment | HW | Auto Inlet module | Inlet drive gear
AQT90 FLEX | 更换903-371后修复(by tim) | 更换903-371后修复(by tim) | SWO-187955 | No | Repair after replacement 903-371 (by tim) | Repair after replacement 903-371 (by tim) | Defect | HW | Incubator module | Punch up
ABL835 FLEX | Comprobar módulo de gases, limpiar transductor y tubos capilares de paso de gases, calibrar y probar con controles. | Limpiar transductor y tubos de paso de gases. | D2749480A24847DB | No | Check gas module, clean transducer and gas passing capillary tubes, calibrate and test with controls. | Clean transducer and gas pass-through tubes. | Cleaned | HW | Gas module | Gas module | Replace Transducer expected
ABL825 FLEX | ABL überprüft, fließtechnisch alle Funktionen OK, keine Auffälligkeiten. Weiterhin Driften hoch und runter bei Kal und Mismatch Meldung bei QC Messung => Oximodul ersetzt, Abgleich, OxiCompleteTest, Kalibrationen, Endkontrolle: drei QC Messungen OK | ABL überprüft, fließtechnisch alle Funktionen OK, keine Auffälligkeiten. Weiterhin Driften hoch und runter bei Kal und Mismatch Meldung bei QC Messung => Oximodul ersetzt, Abgleich, OxiCompleteTest, Kalibrationen, Endkontrolle: drei QC Messungen OK | SWO-188483 | No | ABL checked, flowed all functions OK, no abnormalities. Further drifting up and down at Kal and mismatch message at QC measurement => Oximodul replaced, matching, OxiCompleteTest, calibrations, final control: three QC measurements OK | ABL checked, flowed all functions OK, no abnormalities. Further drifting up and down at Kal and mismatch message at QC measurement => Oximodul replaced, matching, OxiCompleteTest, calibrations, final control: three QC measurements OK | Defect | HW | Oxi module | Oxi module | MetHb and HHb
ABL825 FLEX | Elyte Messkammer geprüft und gereinigt, Funktionskontrolle ok | Elyte Messkammer geprüft und gereinigt, Funktionskontrolle ok | SWO-188474 | No | Elyte measuring chamber tested and cleaned, function control ok | Elyte measuring chamber tested and cleaned, function control ok | Cleaned | HW | El/Met Module | Measuring chamber | Residue
ABL825 FLEX | Allgemeine Kontrolle i.o pCO2 Elektrode ersetzt und Membranenwartung durchgeführt i.o Membrane aus Thalwil Danach Status instabil Erneut Membranenwartung durchgeführt i.o Garantieaustausch pCO2 Elektrode XT04 Kal. und QC i.o | Allgemeine Kontrolle i.o pCO2 Elektrode ersetzt und Membranenwartung durchgeführt i.o Membrane aus Thalwil Danach Status instabil Erneut Membranenwartung durchgeführt i.o Garantieaustausch pCO2 Elektrode XT04 Kal. und QC i.o | SWO-188383 | No | General control i.o pCO2 electrode replaced and membrane maintenance carried out i.o Membrane from Thalwil After that status unstable Again membrane maintenance carried out i.o Warranty replacement pCO2 electrode XT04 Kal. and QC i.o | General control i.o pCO2 electrode replaced and membrane maintenance carried out i.o Membrane from Thalwil After that status unstable Again membrane maintenance carried out i.o Warranty replacement pCO2 electrode XT04 Kal. and QC i.o | Assisted Replacement | Electrode | Electrode | pCO2 | Sensitivity high
AQT90 FLEX | Analysator geöffnet und eingeschaltet. ServiceDump analysiert, i.O. Pumpentest P1-P5 durchgeführt, Änderungen des Value an LS und Htc Zelle ohne Beanstandung. Schlauchlage optimiert, leichte knicke im Schlauchsystem. Metallklammern gelockert. Initialisierungsphase 5 Blankmessungen durchgeführt, ServiceDump erstellt und analysiert. Flow zwischen 1249-1309 Analysator auf Messplatz installiert und gestartet. Analysator messbereit MCC2 9162 | Analysator geöffnet und eingeschaltet. ServiceDump analysiert, i.O. Pumpentest P1-P5 durchgeführt, Änderungen des Value an LS und Htc Zelle ohne Beanstandung. Schlauchlage optimiert, leichte knicke im Schlauchsystem. Metallklammern gelockert. Initialisierungsphase 5 Blankmessungen durchgeführt, ServiceDump erstellt und analysiert. Flow zwischen 1249-1309 Analysator auf Messplatz installiert und gestartet. Analysator messbereit MCC2 9162 | SWO-188461 | No | Analyzer opened and turned on. ServiceDump analyzed, i.O. Pump test P1-P5 performed, changes in value to LS and Htc cell without complaint. Hose position optimized, slight bends in the hose system. Metal clamps loosened. Initialization phase 5 blank measurements performed, ServiceDump created and analyzed. Flow between 1249-1309 Analyzer installed and started on measuring station. Analyzer ready for measurement MCC2 9162 | Analyzer opened and turned on. ServiceDump analyzed, i.O. Pump test P1-P5 performed, changes in value to LS and Htc cell without complaint. Hose position optimized, slight bends in the hose system. Metal clamps loosened. Initialization phase 5 blank measurements performed, ServiceDump created and analyzed. Flow between 1249-1309 Analyzer installed and started on measuring station. Analyzer ready for measurement MCC2 9162 | Adjustment | HW | Aspiration Module | Tubings
AQT90 FLEX | 重新定标后修复(By Ran) | 重新定标后修复(By Ran) | SWO-187959 | No | Repair after re-calibration (By Ran) | Repair after re-calibration (By Ran) | Defect | SW | RMED SW | SW - Upgrade | LEH SW Version
ABL90 FLEX |  | B | SWO-188637 | No |  |  | Defect | HW | Power Supply | Power Supply
AQT90 FLEX |  | B | SWO-188658 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-584后修复(by tim) | 更换903-584后修复(by tim) | SWO-188604 | No | Repair after replacement 903-584 (by tim) | Repair after replacement 903-584 (by tim) | Defect | HW | Aspiration Module | Tube holder
ABL90 | Customer reported that the P02 would not pass cals or QC, they have replaced the SC and SP twice as well as flushed the analyzer. | Arrived on site at 10:30 AM and found the analyzer operating without issues, all cals and qc had passed sometime over the weekend. Tested p02 complete and OXI complete and found ok. Ran all fluid transport tests and found no iisues there as well. Also made changes for the code 39 check digit, the analyzer now ignores the check digit. | D28DDD4CDA53C699 | Yes | Customer reported that the P02 would not pass cals or QC, they have replaced the SC and SP twice as well as flushed the analyzer. | Arrived on site at 10:30 AM and found the analyzer operating without issues, all cals and qc had passed sometime over the weekend. Tested p02 complete and OXI complete and found ok. Ran all fluid transport tests and found no iisues there as well. Also made changes for the code 39 check digit, the analyzer now ignores the check digit.
ABL90 | ISSUES WITH BARCODE SCANNING WHILE SETTING UP INTERFACE WITH OMNITECH | Reconfigured Barcode configuarion in roma.ini to match working unit. | D2785714045A3288 | No | ISSUES WITH BARCODE SCANNING WHILE SETTING UP INTERFACE WITH OMNITECH | Reconfigured Barcode configuarion in roma.ini to match working unit. | Setup | SW | Barcode reader | Barcode reader
ABL90 FLEX PLUS |  | B | D2814D943E4FE80F | Yes |  |  | Re classified RMED | Delivering Manual
ABL90 FLEX |  | B | SWO-188728 | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB
ABL800 BASIC | ABL715⇒ABL800Basic（825GL）へ入れ替え | ABL715⇒ABL800Basic（825GL）へ入れ替え | SWO-194986 | Yes | Replaced by ABL715⇒ABL800Basic(825GL) | Replaced by ABL715⇒ABL800Basic(825GL) | Re classified RMED | Installation
ABL827 FLEX | Checked all tubings and inlet area | Replaced broken nipple. | D27904197A0AD1C6 | Yes | Checked all tubings and inlet area | Replaced broken nipple.
ABL837 FLEX |  | B | D2825BCCB41308A4 | No |  |  | Cleaned | HW | pH/BG Module | Measuring chamber | pCO2
ABL835 FLEX | tHb 1 point calibration value was below specifications. QC ranges are have been changed multiple times to accommodate QC result trend or high CV's. Sample mixing from tubes is not optimum for blood sample comparison and may contribute to spectrum mis-match errors when not mixed properly. pO2 QC results are most likely from running QC manually accross the two instruments as L-3 will be most affected by room air contamination. | Adjusted lamp unit per Radiometer specifications and matched tHb 1point calibration value to peer. Calibrated both instruments for tHb with S7770. All COOX values matched within .1 Recommended that blood comparison samples be pulled into syringe with sample size to exceed 2ml emphasize prior mixing and mixing between samples. Running a cleaning program after every cord gas is recommended. pO2 L-3 QC performance would be optimized by diligence in sample handling to minimize air contamination and alternating instruments for better comparison. | D278ACB90435DBAE | No | tHb 1 point calibration value was below specifications. QC ranges are have been changed multiple times to accommodate QC result trend or high CV's. Sample mixing from tubes is not optimum for blood sample comparison and may contribute to spectrum mis-match errors when not mixed properly. pO2 QC results are most likely from running QC manually accross the two instruments as L-3 will be most affected by room air contamination. | Adjusted lamp unit per Radiometer specifications and matched tHb 1point calibration value to peer. Calibrated both instruments for tHb with S7770. All COOX values matched within .1 Recommended that blood comparison samples be pulled into syringe with sample size to exceed 2ml emphasize prior mixing and mixing between samples. Running a cleaning program after every cord gas is recommended. pO2 L-3 QC performance would be optimized by diligence in sample handling to minimize air contamination and alternating instruments for better comparison. | Adjustment | HW | Oxi module | Lamp unit
ABL825 FLEX | Touch positions not working correctly. | Recalibrate touch screen in Windows 7. Test operation. | D278B2790402AFEA | No | Touch positions not working correctly. | Recalibrate touch screen in Windows 7. Test operation. | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX |  | B | D2829FB9B4468E34 | Yes
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-188890 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC | i Button replace. Solution pack replace. 2-point calibration check. | i Button replace. Solution pack replace. 2-point calibration check. | SWO-195187 | No | i Button replace. Solution pack replace. 2-point calibration check. | i Button replace. Solution pack replace. 2-point calibration check.
ABL90 FLEX | Materiallieferung SC90 946-008 (R0516) | B | SWO-189163 | Yes | Materiallieferung SC90 946-008 (R0516) |  | Re classified RMED | Logistical upSWO
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-188915 | Yes | Repair after replacing test card | Repair after replacing test card
AQT90 FLEX | 更换903-364后修复 by jesson | 更换903-364后修复 by jesson | SWO-188729 | No | Repair by jesson after replacing 903-364 | Repair by jesson after replacing 903-364 | Defect | HW | Process Module | Optical Unit | 131
ABL815 FLEX | Erneut ist die Membrane des BG Modul Gas-Ventils mittig ausgestanzt => Ventil komplett erneuert. Startverfahren etc., Kal und QC Messung OK | Erneut ist die Membrane des BG Modul Gas-Ventils mittig ausgestanzt => Ventil komplett erneuert. Startverfahren etc., Kal und QC Messung OK | SWO-189061 | No | Once again, the membrane of the BG Module gas valve is punched out in the middle => valve is completely renewed. Start method etc., Kal and QC measurement OK | Once again, the membrane of the BG Module gas valve is punched out in the middle => valve is completely renewed. Start method etc., Kal and QC measurement OK | Defect | HW | pH/BG Module | Valve, air | Membrane
ABL825 FLEX | Examined data logs. | Replaced complete inlet module, provided by depot. Verified operation of ABL825. | D279C70E7A31C66C | No | Examined data logs. | Replaced complete inlet module, provided by depot. Verified operation of ABL825. | Defect | HW | Inlet module | Inlet module | clot?
ABL800 BASIC | Cleaned The PH Chamber, Run The Cleaning Mode, Machine Passed 2 Point Calibration & QC Level 1. Machine Working Fine | Cleaned The PH Chamber, Run The Cleaning Mode, Machine Passed 2 Point Calibration & QC Level 1. Machine Working Fine | SWO-190127 | No | Cleaned The PH Chamber, Run The Cleaning Mode, Machine Passed 2 Point Calibration & QC Level 1. Machine Working Fine | Cleaned The PH Chamber, Run The Cleaning Mode, Machine Passed 2 Point Calibration & QC Level 1. Machine Working Fine | Cleaned | HW | pH/BG Module | Measuring chamber
ABL800 BASIC | replaced the touch screen with new, did calibration for touch screen ,found touch working properly.instrument kept under observation. | B | SWO-188987 | No | replaced the touch screen with new, did calibration for touch screen ,found touch working properly.instrument kept under observation. |  | Defect | HW | DMS_PC | Touch screen
ABL835 FLEX | Kein Fehler festgestellt Friction Detector Unit ausgetauscht Einlass Nadel und Einlass Dichtung ausgetauscht Ventile überprüft Kapillare Testmessung durchgeführt Funktionskontrolle i.o | Kein Fehler festgestellt Friction Detector Unit ausgetauscht Einlass Nadel und Einlass Dichtung ausgetauscht Ventile überprüft Kapillare Testmessung durchgeführt Funktionskontrolle i.o | SWO-188956 | No | No error detected Detector Unit replaced Inlet needle and inlet seal replaced Valves checked Capillary test measurement performed Functional control i.o | No error detected Detector Unit replaced Inlet needle and inlet seal replaced Valves checked Capillary test measurement performed Functional control i.o | Revisit required | HW | Inlet module | Inlet Gasket
ABL835 FLEX | Friction Detector zu SWO188956 | Kein Fehler festgestellt Friction Detector Unit ausgetauscht Einlass Nadel und Einlass Dichtung ausgetauscht Ventile überprüft Kapillare Testmessung durchgeführt Funktionskontrolle i.o | SWO-189153 | No | Detector Friction to SWO188956 | No error detected Detector Unit replaced Inlet needle and inlet seal replaced Valves checked Capillary test measurement performed Functional control i.o | Defect | HW | Inlet module | Friction detector unit
AQT90 FLEX | Found and replaced with Process wheel and OU, Cup wash, Calibrated instrument found Calibration results are Ok, performed the test results are OK. Found instrument is working satisfactorily. | Found and replaced with Process wheel and OU, Cup wash, Calibrated instrument found Calibration results are Ok, performed the test results are OK. Found instrument is working satisfactorily. | SWO-190116 | No | Found and replaced with Process wheel and OU, Cup wash, Calibrated instrument found Calibration results are Ok, performed the test results are OK. Found instrument is working satisfactorily. | Found and replaced with Process wheel and OU, Cup wash, Calibrated instrument found Calibration results are Ok, performed the test results are OK. Found instrument is working satisfactorily. | Defect | HW | Process Module | Process wheel
AQT90 FLEX | Storage wheel check. reagent cartrige check. Service mode check. analyzer reset. | Storage wheel check. reagent cartrige check. Service mode check. analyzer reset. | SWO-195190 | No | Storage wheel check. reagent cartrige check. Service mode check. analyzer reset. | Storage wheel check. reagent cartrige check. Service mode check. analyzer reset. | Adjustment | HW | Storage Module | Storage wheel | 5689
ABL80 FLEX CO-OX | Fr. Ütmü hat wie gebeten SC und LP erneuert, zuvor das Datum+Uhrzeit richtig gestellt. Die Uhrzeit verstellt sich weiterhin nach Neustart um 1 Std. Vermutlich Sommer Winterzeit Problem. Das Datum ändert sich wohl im Betrieb auf Morgen. Fr. Ütmü bittet um einen Service Einsatz. | Fr. Ütmü hat wie gebeten SC und LP erneuert, zuvor das Datum+Uhrzeit richtig gestellt. Die Uhrzeit verstellt sich weiterhin nach Neustart um 1 Std. Vermutlich Sommer Winterzeit Problem. Das Datum ändert sich wohl im Betrieb auf Morgen. Fr. Ütmü bittet um einen Service Einsatz. | SWO-189365 | No | Ms. Ütmü has renewed SC and LP as requested, previously set the date+time correctly. The time still adjusts by 1 hour after reboot. Probably summer winter time problem. The date probably changes to tomorrow. Ms. Ütmü asks for a service. | Ms. Ütmü has renewed SC and LP as requested, previously set the date+time correctly. The time still adjusts by 1 hour after reboot. Probably summer winter time problem. The date probably changes to tomorrow. Ms. Ütmü asks for a service.
ABL80 FLEX BASIC | Found SC 80 air detection failure, Inactivated the Air detection reinstalled the SC but problem was remaining same, replaced with new SC 80 found instrument is functional. | Found SC 80 air detection failure, Inactivated the Air detection reinstalled the SC but problem was remaining same, replaced with new SC 80 found instrument is functional. | SWO-190119 | Yes | Found SC 80 air detection failure, Inactivated the Air detection reinstalled the SC but problem was remaining same, replaced with new SC 80 found instrument is functional. | Found SC 80 air detection failure, Inactivated the Air detection reinstalled the SC but problem was remaining same, replaced with new SC 80 found instrument is functional.
ABL80 CO-OX |  | B | D27BFF906642EC8B | Yes
ABL90 |  | B | D27ACFE9F05AD6A6 | No |  |  | Defect | HW | Barcode reader | Barcode reader, display
ABL90 FLEX | Uhrzeit eigestellt und Verfahren für die automatische Zeiteinstellung angewandt, System und Funktion i.O. | Einstellungen auf Standard zurückgestellt, da die Analysatoren an einem Zeitserver hängen!!!! Anschließend Analysator mehrfach neu gestartet, um zu überprüfen, dass die Einstellungen passen System und Funktion i.O. | SWO-189528 | No | time and procedures for automatic time setting applied, system and function in the order | Settings are set back to default because the analyzers are attached to a time server!!!! Then analyzer restarted several times to verify that the settings fit System and function in the | Setup | SW | Time setting | Time setting
ABL90 | SCHEDULED D ECONTAMINATION. | CONDUCTED DECONTAMINATION ON ANALYZER. REPLACED SENSOR CASSETTE AND SOLUTION PACK. S UBMITTED C REDIT C LAIM FORMS . | D27B0020F03BA6CD | Yes | SCHEDULED D ECONTAMINATION. | CONDUCTED DECONTAMINATION ON ANALYZER. REPLACED SENSOR CASSETTE AND SOLUTION PACK. S UBMITTED C REDIT C LAIM FORMS . | Re classified RMED | PM
ABL90 FLEX PLUS |  | B | D27ACD80F05DCA51 | No |  |  | Adjustment | HW | Basic Unit | Touch screen
ABL90 FLEX PLUS | Replaced the inlet probe with a new one. | Replaced the inlet probe with a new one. | SWO-189414 | No | Replaced the inlet probe with a new one. | Replaced the inlet probe with a new one. | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX PLUS | Präventiv Einweisung durchgeführt. | Sr. Anja bittet um eine Überprüfung des ABL. | SWO-189814 | No | Preventive instruction. | Sr. Anja asks for a review of the ABL. | Defect | HW | PM Kit | PM Kit
ABL90 FLEX PLUS | Roma.ini heb aangepast. | Roma.ini heb aangepast. | SWO-189326 | Yes | Roma. INI have adjusted. | Roma. INI have adjusted. | Re classified RMED | SW | His/lis | His/lis | FSE : On request of the client, Roma.ini file has to be adjusted in order to change the patient ID to AccesionNumber.
ABL90 FLEX PLUS | Roma.ini heb aangepast. | Roma.ini heb aangepast. | SWO-189327 | Yes | Roma. INI have adjusted. | Roma. INI have adjusted. | Re classified RMED | SW | His/lis | His/lis | See SWO-189326
ABL90 FLEX PLUS |  | B | D27B1188F0055958 | No |  |  | Revisit required | Unknown | Unknown | Unknown | Techline
ABL90 FLEX | 连续更换试剂包和测试卡后修复 | 连续更换试剂包和测试卡后修复 | SWO-189520 | Yes | Repair after continuous replacement of reagent kits and test cards | Repair after continuous replacement of reagent kits and test cards
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-189492 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX | 更换903-364，重新定标。by jesson | 更换903-364，重新定标。by jesson | SWO-188848 | No | Replace the 903-364 and re-mark the bid. by jesson | Replace the 903-364 and re-mark the bid. by jesson | Defect | HW | Process Module | Optical Unit | 131
ABL815 FLEX | Geräte gemäß Absprache mit Frau Bolk und Dr. Jedzek - Zader konfiguriert und getestet. | B | SWO-189400 | No | Devices configured and tested in consultation with Ms. Bolk and Dr. Jedzek - Zader. |  | Setup | SW | His/lis | His/lis
ABL825 FLEX | Customer reported that after scanning the accession number during a sample then pressing enter on the external keyboard would cause the analyzer to shut down. | Tested power supply and found the -12vdc out of tolerance (-12.84vdc). Replaced power supply and tested the analyzer. It no longer reboots when entering demographics during samples. | D28DE495DA2EC6A6 | No | Customer reported that after scanning the accession number during a sample then pressing enter on the external keyboard would cause the analyzer to shut down. | Tested power supply and found the -12vdc out of tolerance (-12.84vdc). Replaced power supply and tested the analyzer. It no longer reboots when entering demographics during samples. | Defect | HW | Power Supply | Power Supply | Same day
ABL825 FLEX | Performed flex Q calibration in service mode, found position 1 was calibrated OK but positions 2 and 3 were off and needed adjusting. | After performing calibration tested samples from each position of flex module and found all operations normal. | D28DE06DDA1FC69E | No | Performed flex Q calibration in service mode, found position 1 was calibrated OK but positions 2 and 3 were off and needed adjusting. | After performing calibration tested samples from each position of flex module and found all operations normal. | Adjustment | HW | FlexQ Module | Positioning | Same day
ABL825 FLEX | BG Messkammer geprüft, demontiert und gereinigt, Funktionskontrolle ok | BG Messkammer geprüft, demontiert und gereinigt, Funktionskontrolle ok | SWO-189688 | No | BG measuring chamber tested, dismantled and cleaned, function control ok | BG measuring chamber tested, dismantled and cleaned, function control ok | Cleaned | HW | El/Met Module | Measuring chamber
ABL825 FLEX | 0331 errors are present, indicating no sample detected. Ampoule breaker appears to break ampoules. However, some are broken out of alignment or access to a different ampoule than the one broken is observed. | Replaced the 14 pole autocheck cable. Verified that all QC ampoules are properly scanned, broken and accessed during QC running. Analyzer is within acceptable operating specifications. | D27AE573F02DC117 | No | 0331 errors are present, indicating no sample detected. Ampoule breaker appears to break ampoules. However, some are broken out of alignment or access to a different ampoule than the one broken is observed. | Replaced the 14 pole autocheck cable. Verified that all QC ampoules are properly scanned, broken and accessed during QC running. Analyzer is within acceptable operating specifications. | Defect | HW | Autocheck module | FPC Cable
ABL837 FLEX | Client advised of failing QCs for inhomogeneous soltiions intermittently. | Adjusted Autocheck Carousel alignment, replaced inlet gasket. | D27ADB60F0516A4D | No | Client advised of failing QCs for inhomogeneous soltiions intermittently. | Adjusted Autocheck Carousel alignment, replaced inlet gasket. | Adjustment | HW | Autocheck module | Carroussel
AQT90 FLEX | Mit Fr. Weber das Scannerfenster gereinigt, der erste Barcode wurde erkannt, beim zweiten bleibt der Fehler, zweiten Barcode von Hand eingeben lassen. Kundin bittet um Service Einsatz am Montag oder Mittwoch nächste Woche. | Mit Fr. Weber das Scannerfenster gereinigt, der erste Barcode wurde erkannt, beim zweiten bleibt der Fehler, zweiten Barcode von Hand eingeben lassen. Kundin bittet um Service Einsatz am Montag oder Mittwoch nächste Woche. | SWO-189341 | No | With Ms. Weber cleaned the scanner window, the first barcode was detected, the second one remains the error, the second barcode is entered by hand. Customer asks for service on Monday or Wednesday next week. | With Ms. Weber cleaned the scanner window, the first barcode was detected, the second one remains the error, the second barcode is entered by hand. Customer asks for service on Monday or Wednesday next week. | Cleaned | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX |  | B | D27AD44EF0077B2B | No |  |  | Cleaned | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-189755 | No | Repair after replacing test card | Repair after replacing test card
ABL80 FLEX BASIC |  | B | SWO-189554 | No
ABL90 FLEX | Reinigung Fließwege, Service-Programme ausgeführt, tHb-Kal wiederholt, Kal i.O., QC i.O. | B | SWO-189950 | No | Cleaning flow paths, service programs executed, tHb-Kal repeated, Kal i.O., QC i.O. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | tHb
ABL90 | Reviewed circumstances leading to reported issue. Customer indicates a SP was recently replaced, having difficulty removing it to re-install. Suspect SP was not activated completely as cause of issue. | Had customer remove SP, determined the SP was not activated completely. Customer pushed down on top of SP completely to engage, reinstalled, SP initialized and completed calibration and QC with no further problems. | D28086803E5B8768 | No | Reviewed circumstances leading to reported issue. Customer indicates a SP was recently replaced, having difficulty removing it to re-install. Suspect SP was not activated completely as cause of issue. | Had customer remove SP, determined the SP was not activated completely. Customer pushed down on top of SP completely to engage, reinstalled, SP initialized and completed calibration and QC with no further problems. | User error | SP | SP Activation | SP Activation
ABL90 FLEX | Einstellungen auf Standard zurückgestellt, da die Analysatoren an einem Zeitserver hängen!!!! Anschließend Analysator mehrfach neu gestartet, um zu überprüfen, dass die Einstellungen passen System und Funktion i.O. | Einstellungen auf Standard zurückgestellt, da die Analysatoren an einem Zeitserver hängen!!!! Anschließend Analysator mehrfach neu gestartet, um zu überprüfen, dass die Einstellungen passen System und Funktion i.O. | SWO-189663 | No | Settings are set back to default because the analyzers are attached to a time server!!!! Then analyzer restarted several times to verify that the settings fit System and function in the | Settings are set back to default because the analyzers are attached to a time server!!!! Then analyzer restarted several times to verify that the settings fit System and function in the | Setup | SW | Time setting | Time setting
ABL90 FLEX PLUS | Datensicherung auf allen Ebenen durchgeführt. Roma wurde beendet und nicht neu gestartet. Überprüfung der Module und Neustart des Analysators. Initialisierungsphase 1Stunde. Kalibrierfehler pCO2, Glu und Lac Fällige tHb Kal durchgeführt Kal und QC in diversen Parametern nicht i.O. SC wurde getauscht und reklamiert. | Datensicherung auf allen Ebenen durchgeführt. Roma wurde beendet und nicht neu gestartet. Überprüfung der Module und Neustart des Analysators. Initialisierungsphase 1Stunde. Kalibrierfehler pCO2, Glu und Lac Fällige tHb Kal durchgeführt Kal und QC in diversen Parametern nicht i.O. SC wurde getauscht und reklamiert. | SWO-189667 | No | data backup at all levels. Roma has been terminated and not restarted. Check the modules and restart the analyzer. Initialization phase 1hour. Calibration error pCO2, Glu and Lac Due tHb Kal performed Kal and QC in various parameters not in the SC has been exchanged and complained about. | data backup at all levels. Roma has been terminated and not restarted. Check the modules and restart the analyzer. Initialization phase 1hour. Calibration error pCO2, Glu and Lac Due tHb Kal performed Kal and QC in various parameters not in the SC has been exchanged and complained about. | Defect | SW | Power off/on | Power off/on
ABL90 FLEX PLUS | Copie Roma Ini de l'équipement ABL90 + du Labo e.o Test informatique ok e.o Calibration e.o contrôle de qualité e.o | Copie Roma Ini de l'équipement ABL90 + du Labo e.o Test informatique ok e.o Calibration e.o contrôle de qualité e.o | SWO-189611 | No | Roma Ini copy of the ABL90 equipment - from the Lab e.o Computer test ok e.o Calibration e.o quality control e.o | Roma Ini copy of the ABL90 equipment - from the Lab e.o Computer test ok e.o Calibration e.o quality control e.o | Setup | SW | Roma.ini | Roma.ini
ABL90 FLEX PLUS | Roma.ini heb aangepast. | Roma.ini heb aangepast. | SWO-189720 | Yes | Roma. INI have adjusted. | Roma. INI have adjusted. | Re classified RMED | SW | His/lis | His/lis | See SWO-189326
ABL90 FLEX PLUS |  | B | D27CABB86630FFD2 | Yes |  |  | Re classified RMED | Logistical upSWO | Supply SC
ABL90 FLEX |  | B | SWO-189484 | No |  |  | Defect | HW | Inlet module | Inlet module
ABL90 | REVISIÓN DE HISTÓRICOS. REALIZAR PURGAS DE LAVADO DEL CIRCUITO DE FLUÍDICA CON NaCl (10%) Y AGUA DESTILADA. LIMPIEZA DE LA SONDA DE ENTRADA. CALIBRACIONES Y CONTROLES | REVISIÓN DE HISTÓRICOS. REALIZAR PURGAS DE LAVADO DEL CIRCUITO DE FLUÍDICA CON NaCl (10%) Y AGUA DESTILADA. LIMPIEZA DE LA SONDA DE ENTRADA. | D27C1366660C68A2 | No | HISTORICAL REVIEW. MAKE WASH PURGES OF THE FLUDIC CIRCUIT WITH NaCl (10%) AND DESTILATED WATER. CLEANING THE INPUT PROBE. CALIBRATIONS AND CONTROLS | HISTORICAL REVIEW. MAKE WASH PURGES OF THE FLUDIC CIRCUIT WITH NaCl (10%) AND DESTILATED WATER. CLEANING THE INPUT PROBE. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
AQT90 FLEX | 更换910-304一个正常。by jesson | 更换910-304一个正常。by jesson | SWO-189162 | No | Replace 910-304 one normal. by jesson | Replace 910-304 one normal. by jesson | Defect | HW | Storage Module | Barcode reader, Storage wheel | BR1
ABL827 FLEX | QC samples not always aspirating, random inlet errors. | Replaced inlet probe. | D27C9CC466422867 | No | QC samples not always aspirating, random inlet errors. | Replaced inlet probe. | Defect | HW | Inlet module | Probe
ABL825 FLEX | Gerät überprüft. Empf.Status unauffällig. Messkammer BG Modul gereinigt. Blut an pH Elektrode entfernt. Einlass durchgespült. Air Valve V7 gespült. Nach mehreren Kal pH ok. Na Grenzwertig hoch bei Q.C. nach erneuter Kal und Q.C alles ok. | Ein Vorortservice wird gewünscht. | SWO-189725 | No | device checked. Empf.Status inconspicuous. Measuring chamber BG module cleaned. Blood removed from pH electrode. Inlet flushed through. Air Valve V7 rinsed. After several cal pH ok. Well high at Q.C. after re-cal and Q.C all ok. | A suburban service is desired. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL835 FLEX | USB port testet med andre usb stick og der er godt nok forlænget ventetid, men det kan kan lade sig gøre at benytte porten. | USB port testet med andre usb stick og der er godt nok forlænget ventetid, men det kan kan lade sig gøre at benytte porten. | SWO-189965 | No | USB Port tested with other USB stick and there is well enough prolonged latency, but it can be possible to use the port. | USB Port tested with other USB stick and there is well enough prolonged latency, but it can be possible to use the port. | No fault found | SW | DMS_PC | USB | Slow comm.
ABL835 FLEX | Udskiftning af Cl elektroden | Udskiftning af Cl elektroden | SWO-189976 | No | Replacing the Cl electrode | Replacing the Cl electrode | Assisted Replacement | Electrode | Electrode | Cl
ABL825 FLEX | Checked data logs for errors. Unit getting errors code 643 DMS disconnected from AU. | Troubleshoot unit. Found first analyzer control PCB defective. Replaced with second Analyzer Control PCB. Re-installed option keys. | D27C33D16636954C | No | Checked data logs for errors. Unit getting errors code 643 DMS disconnected from AU. | Troubleshoot unit. Found first analyzer control PCB defective. Replaced with second Analyzer Control PCB. Re-installed option keys. | Defect | HW | Basic Unit | Analyzer control PCB
AQT90 FLEX |  | B | D27BF31F666172A2 | No |  |  | Unknown | Unknown | Unknown | Unknown | Techline
AQT90 FLEX |  | B | D27C2511662B044D | No |  |  | Defect | HW | Aspiration Module | Tube holder
TCM/Tosca | Found Comm, PCB was defective. | Replaced Comm, PCB 903-101 | D27C3EBC66201E6A | Yes | Found Comm, PCB was defective. | Replaced Comm, PCB 903-101
ABL80 OSM | Computer unit fan is making excessive noise and battery has no life at all. | Replaced fan & battery. Calibrated and ran QC. Instrument operates normally. | D27D7061DC27D518 | No | Computer unit fan is making excessive noise and battery has no life at all. | Replaced fan & battery. Calibrated and ran QC. Instrument operates normally.
ABL80 FLEX BASIC |  | B | SWO-189877 | No
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-189866 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX | ガスケット洗浄→回路洗浄 | B | SWO-190212 | No | Gasket cleaning → Circuit cleaning |  | Assisted Replacement | SP | SP | FT | 181
ABL90 FLEX | Nach Erneuerung der Sensorkassette, hat sich das Gerät einkalibriert. Alle Service Tests durchgeführt. Kal. und QCs sind i.O | Herr K. wünscht für morgen Früh (5.5.) einen Vorortservice. | SWO-190494 | No | After renewing the sensor cartridge, the device has calibrated itself. All service tests performed. Kal. and QCs are i.O. | Mr. K. wants a suburban service for tomorrow morning (5.5.). | User error | SC | SC | SC expired
ABL90 FLEX | Einweisung durchgeführt. | Einweisung durchgeführt. | SWO-190911 | Yes | instruction carried out. | instruction carried out. | Re classified RMED | Double entry
ABL90 FLEX | Konfiguration aller 3 Geräte überprüft, neu eingegeben und abgespeichert. Radiance Verbindung wieder verfügbar und Daten werden übermittelt. Status der Gerät OK. | Konfiguration aller 3 Geräte überprüft, neu eingegeben und abgespeichert. Radiance Verbindung wieder verfügbar und Daten werden übermittelt. Status der Gerät OK. | SWO-190146 | No | Configuration of all 3 devices checked, re-entered and saved. Radiance connection is available again and data is transmitted. Status of the device OK. | Configuration of all 3 devices checked, re-entered and saved. Radiance connection is available again and data is transmitted. Status of the device OK. | Setup | SW | His/lis | His/lis
ABL90 FLEX PLUS | SC wird vom Kunden erneuert und erhält von mir (TobEh) gegen 12 Uhr einen Rückruf!!!! | SC wird vom Kunden erneuert und erhält von mir (TobEh) gegen 12 Uhr einen Rückruf!!!! | SWO-189861 | No | SC is renewed by the customer and receives a call back from me (TobEh) around 12 o'clock!!!! | SC is renewed by the customer and receives a call back from me (TobEh) around 12 o'clock!!!! | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS |  | B | D2822741B431F144 | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL5 | Austausch des defekten Druckers aus Kundengerät Kal- und QC | Austausch des defekten Druckers aus Kundengerät Kal- und QC | SWO-190045 | No | Replacing the defective printer from customer device Kal- and QC | Replacing the defective printer from customer device Kal- and QC
AQT90 FLEX | 更换903-364后未修复，要重新借903-236一个 by jesson | 更换903-236后修复(By Ran) | SWO-189218 | No | Replace 903-364 after not repaired, to borrow 903-236 a by jesson | Repair after replacement 903-236 (By Ran) | Revisit required | HW | Process Module | Optical Unit | 129
AQT90 FLEX |  | B | SWO-189858 | Yes |  |  | Re classified RMED | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364一个正常。by jesson | 更换903-364一个正常。by jesson | SWO-189563 | No | Replace 903-364 one normal. by jesson | Replace 903-364 one normal. by jesson | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364后修复 | 更换903-364后修复 | SWO-189546 | No | Repair after replacement 903-364 | Repair after replacement 903-364 | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | 更换903-364后修复 | SWO-189971 | No |  | Repair after replacement 903-364 | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX |  | B | SWO-189846 | No |  |  | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | 更换902-847后修复 by jesson | 更换902-847后修复 by jesson | SWO-189513 | Yes | Repair by jesson after replacement 902-847 | Repair by jesson after replacement 902-847 | Re classified RMED | OOB
ABL835 FLEX | Målekammer renset Glu elektrode skiftet | Målekammer renset Glu elektrode skiftet | SWO-189972 | No | Measuring chamber cleaned The Glu electrode switch | Measuring chamber cleaned The Glu electrode switch | Cleaned | HW | El/Met Module | Measuring chamber
ABL827 FLEX | Kal2-Lösung niedriger Stand --> Kal-Fehler bei Empf. 2und Empf. 3, alle Lösungen erneuert, Sitz der Elektroden geprüft --> Kal. --> Kal- Fehler Empf.3 --> Neumembranierung nicht möglich, da keine Membranen auf Station verfügbar sind --> Station bestellt Material und führt die Membranierung in Verbindung mit einer Einweisung ( Radiometer ) durch MedTech wird informiert | Kal2-Lösung niedriger Stand --> Kal-Fehler bei Empf. 2und Empf. 3, alle Lösungen erneuert, Sitz der Elektroden geprüft --> Kal. --> Kal- Fehler Empf.3 --> Neumembranierung nicht möglich, da keine Membranen auf Station verfügbar sind --> Station bestellt Material und führt die Membranierung in Verbindung mit einer Einweisung ( Radiometer ) durch MedTech wird informiert | SWO-189906 | No | Kal2 solution low-level --> Kal error at Empf. 2and Empf. 3, all solutions renewed, seat of electrodes tested --> Kal. --> Kal- Error Empf.3 --> Re-membraneing not possible, as no membranes are available on station --> Station orders material and performs the membraneing in conjunction with a briefing ( radiometer ) MedTech is informed | Kal2 solution low-level --> Kal error at Empf. 2and Empf. 3, all solutions renewed, seat of electrodes tested --> Kal. --> Kal- Error Empf.3 --> Re-membraneing not possible, as no membranes are available on station --> Station orders material and performs the membraneing in conjunction with a briefing ( radiometer ) MedTech is informed | Assisted Replacement | Electrode | Membrane | Glu
ABL835 FLEX | Performed complete oxi test. All results in range. Problem seems to be intermittent. | Replaced Spectrometer. After replacing Lamp unit failed complete oxi test. Replaced lamp unit. | D27D69ACDC121963 | No | Performed Complete oxy test. ALL results in range. Problem seems to be intermittent. | Replaced Spectrometer. After replacing Lamp unit failed complete oxi test. Replaced lamp unit. | Defect | HW | Oxi module | Lamp unit
ABL825 FLEX | Fehlersuche durchgeführt. Neues Oxymodul eingebaut. Keine Fehler mehr vorhanden. | Fehlersuche durchgeführt. Neues Oxymodul eingebaut. Keine Fehler mehr vorhanden. | SWO-189993 | No | Troubleshooting. New oxymodule built-in. No more errors. | Troubleshooting. New oxymodule built-in. No more errors. | Defect | HW | Oxi module | Oxi module
ABL825 FLEX |  | B | D27D90D2DC5EED6B | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-190099 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Chip coding
ABL800 FLEX | Try to replace Power Supply unit- but not switching ON. Try to replace Analyser control PCB but Same problem found. CPU unit cleaned and try to reconnect but not switching ON. A burn of IC found upon CPU unit. Need to replace CPU unit. Sending TO for this will tack 1 or 2 week. | Try to replace Power Supply unit- but not switching ON. Try to replace Analyser control PCB but Same problem found. CPU unit cleaned and try to reconnect but not switching ON. A burn of IC found upon CPU unit. Need to replace CPU unit. Sending TO for this will tack 1 or 2 week. | SWO-190161 | No | Try to replace Power Supply unit- but not switching ON. Try to replace Analyser control PCB but Same problem found. CPU unit cleaned and try to reconnect but not switching ON. A burn of IC found upon CPU unit. Need to replace CPU unit. Sending TO for this will tack 1 or 2 week. | Try to replace Power Supply unit- but not switching ON. Try to replace Analyser control PCB but Same problem found. CPU unit cleaned and try to reconnect but not switching ON. A burn of IC found upon CPU unit. Need to replace CPU unit. Sending TO for this will tack 1 or 2 week. | Revisit required | HW | DMS_PC | CPU | Return module to EJNSC
AQT90 FLEX | 代替機を設置し、機器引上げ。 | B | SWO-193231 | No | An alternative machine is installed and the equipment is raised. |  | Defect | HW | Process Module | Optical Unit | 129
ABL90 FLEX |  | B | SWO-190164 | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-190205 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Lac
AQT90 FLEX | 更换903-236后修复(By Ran) | 更换903-236后修复(By Ran) | SWO-189837 | No | Repair after replacement 903-236 (By Ran) | Repair after replacement 903-236 (By Ran) | Defect | HW | Process Module | Optical Unit | 129
AQT90 FLEX | 更换636-541及903-329后修复 | 追加借用备件 903-329 | SWO-190117 | No | Repair after replacement 636-541 and 903-329 | Additional borrowed spare parts 903-329 | Revisit required | HW | Aspiration Module | Needle Wash Unit | 1999
AQT90 FLEX | 追加借用备件 903-329 | 追加借用备件 903-329 | SWO-190121 | No | Additional borrowed spare parts 903-329 | Additional borrowed spare parts 903-329 | Defect | HW | Aspiration Module | Needle Wash Unit | 1999
ABL80 FLEX BASIC |  | B | SWO-190368 | No
ABL80 FLEX CO-OX | Korrekturen die im Bereich Date und Time gemacht und gespeichert wurden, werden beibehalten. Auch eine kurzzeitige Ausschaltung verändert das Datum nicht. die Sommerzeit deaktiviert. Falls weiterhin Zeitveränderung vorhanden, erneute Anfahrt. | Fr. Ütmü hat wie gebeten SC und LP erneuert, zuvor das Datum+Uhrzeit richtig gestellt. Die Uhrzeit verstellt sich weiterhin nach Neustart um 1 Std. Vermutlich Sommer Winterzeit Problem. Das Datum ändert sich wohl im Betrieb auf Morgen. Fr. Ütmü bittet um einen Service Einsatz. | SWO-190538 | No | Corrections made and saved in the Date and Time area are retained. Even a short-term shutdown does not change the date. daylight saving time is deactivated. If there is still time change, re-enter. | Ms. Ütmü has renewed SC and LP as requested, previously set the date+time correctly. The time still adjusts by 1 hour after reboot. Probably summer winter time problem. The date probably changes to tomorrow. Ms. Ütmü asks for a service.
ABL90 | Non linear pH, pCO2 and p02. Level 3 reading high and Level 2 readings low. | Replaced sensor cassette and Ran Verification Kit, VK-R5, lot number 478. | D28148F73E097C0D | No | Non linear pH, pCO2 and p02. Level 3 reading high and Level 2 readings low. | Replaced sensor cassette and Ran Verification Kit, VK-R5, lot number 478. | Assisted Replacement | SC | SC | BG | VRK
ABL90 FLEX | Inbetriebnahme für Problauf durchgeführt. Neues LP (Lot: YI-61 defekt. LP erneuert (Lot: YI-52). | Inbetriebnahme für Problauf durchgeführt. Neues LP (Lot: YI-61 defekt. LP erneuert (Lot: YI-52). | SWO-190466 | No | commissioning for Problauf. New LP (Lot: YI-61 defective. LP renewed (lot: YI-52). | commissioning for Problauf. New LP (Lot: YI-61 defective. LP renewed (lot: YI-52). | Assisted Replacement | SP | SP | FT
ABL90 FLEX PLUS |  | B | D281889B3E0D16AD | Yes
ABL90 FLEX PLUS |  | B | D29301A4B261F127 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Zeitumstellung nach RMED Anleitung durchgeführt. Gerät zum Testen 2 x neu gestartet. Alles i.O | Zeitumstellung nach RMED Anleitung durchgeführt. Gerät zum Test 2 x neu gestartet. Alles ok. | SWO-191013 | No | time change according to RMED instructions. Device to test 2 x restarted. Everything i.O. | time change according to RMED instructions. Device to test 2 x restarted. All ok. | Setup | Time setting | Time setting | Time setting
ABL90 FLEX PLUS |  | B | D28E043ADA3A5025 | Yes
AQT90 FLEX |  | B | SWO-190260 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换备件905-953及636-551后修复 | B | SWO-190179 | No | Replacement of spare parts 905-953 and 636-551 after repair |  | Defect | HW | Aspiration Module | Barcode reader, Inlet
ABL815 FLEX | 自行更换备用件902-458，902-844两个，902-700后修复(By Ran) | 自行更换备用件902-458，902-844两个，902-700后修复(By Ran) | SWO-190232 | No | Replace spare parts 902-458, 902-844 two, 902-700 post-fix (By Ran) | Replace spare parts 902-458, 902-844 two, 902-700 post-fix (By Ran) | Defect | HW | Oxi module | Lamp unit | USED Analyzer
ABL837 FLEX | ELモジュール部、フラットメンブランバルブ用メンブラン破断の為、交換。 | B | SWO-190574 | No | Replacement for EL module part, membrane breakage for flat membrane valve. |  | Defect | HW | El/Met Module | El/Met Module | Sample transport
ABL800 BASIC | Replaced The Membranes But Still Calibration Error In Lactate, Glucose & Sodium. Further Testing Will be Done After Replacing Reference Membrane | Replaced The Membranes But Still Calibration Error In Lactate, Glucose & Sodium. Further Testing Will be Done After Replacing Reference Membrane | SWO-193688 | No | Replaced The Membranes But Still Calibration Error In Lactate, Glucose & Sodium. Further Testing Will be Done After Replacing Reference Membrane | Replaced The Membranes But Still Calibration Error In Lactate, Glucose & Sodium. Further Testing Will be Done After Replacing Reference Membrane | Unknown | HW | El/Met Module | Tubings | Email
ABL800 BASIC | Cleaned the inlet gasket & waste tube performed protien removal 2 point calibration done sample processed found ok | Cleaned the inlet gasket & waste tube performed protien removal 2 point calibration done sample processed found ok | SWO-190897 | No | Cleaned the inlet gasket & waste tube performed protien removal 2 point calibration done sample processed found ok | Cleaned the inlet gasket & waste tube performed protien removal 2 point calibration done sample processed found ok | Cleaned | HW | Inlet module | Inlet Gasket
ABL825 FLEX | Erneuerung der pCO2 Elektrode, Reinigung der Messkammer und der Elektrodenkontakte. Reinigung auch der pH Messkammer. Werte nach mehrmaligen Kalibration wieder in Ordnung. | Kunde hat pCO2-Elektrode quer getauscht, der Kalibrierfehler geht mit der Elektrode mit - Kunde biitet um Service, da er meint, die pCO2-Elektrode ist defekt | SWO-190537 | No | Renewal of the pCO2 electrode, cleaning of the measuring chamber and the electrode contacts. Cleaning also of the pH measuring chamber. Values are fine again after several calibrations. | Customer has swapped pCO2 electrode across, the calibration error goes with the electrode - customer biitet for service, because he thinks the pCO2 electrode is defective | Assisted Replacement | Electrode | Electrode | pCO2
ABL800 BASIC | Found head roller and lock problem need to replace with new, suggested to customer we can’t give new printer head, took old ABL800 printer and replaced the same restarted instrument found printer was working satisfactorily | Found head roller and lock problem need to replace with new, suggested to customer we can’t give new printer head, took old ABL800 printer and replaced the same restarted instrument found printer was working satisfactorily | SWO-191787 | No | Found head roller and lock problem need to replace with new, suggested to customer we can’t give new printer head, took old ABL800 printer and replaced the same restarted instrument found printer was working satisfactorily | Found head roller and lock problem need to replace with new, suggested to customer we can’t give new printer head, took old ABL800 printer and replaced the same restarted instrument found printer was working satisfactorily | Defect | HW | Printer Module | Printer head
ABL825 FLEX | CONFIRMED ISSUE WHEN ONSITE. 14-POLE FLEX CABLE WAS NOT INSTALLED PROPERLY. | REPLACED FLEX CABLE. | D2810F943E3D4A55 | No | CONFIRMED ISSUE WHEN ONSITE. 14-POLE FLEX CABLE WAS NOT INSTALLED PROPERLY. | REPLACED FLEX CABLE. | Defect | HW | Autocheck module | FPC Cable
AQT90 FLEX | Kein kontinuierlicher Strahl des Scanners, flackert. Kompletten Scanner ausgetauscht. Analysator eingeschaltet. Initialisierungsphase. Barcodes der QC Chargen eingelesen und korrigiert. Volle Funktion des Scanners. QC wird durch Frau Vogelsberger gemessen. | Kein kontinuierlicher Strahl des Scanners, flackert. Kompletten Scanner ausgetauscht. Analysator eingeschaltet. Initialisierungsphase. Barcodes der QC Chargen eingelesen und korrigiert. Volle Funktion des Scanners. QC wird durch Frau Vogelsberger gemessen. | SWO-190495 | No | No continuous beam of the scanner, flickers. Complete scanner replaced. analyzer switched on. Initialization phase. barcodes of the QC batches are read and corrected. Full function of the scanner. QC is measured by Ms Vogelsberger. | No continuous beam of the scanner, flickers. Complete scanner replaced. analyzer switched on. Initialization phase. barcodes of the QC batches are read and corrected. Full function of the scanner. QC is measured by Ms Vogelsberger. | Defect | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX | Startphase | QC Gemessen. | SWO-190646 | No | Start | QC Measured. | Defect | HW | Analyzer Control Module | Cable, Display
AQT90 FLEX | Fehlersuche, div. Anwendungen im Serviceprogr. durchgef., Gerät aus, Blower + Filter ausgebaut, gereinigt u. eingebaut, Gerät ein, Startphase abgewartet, QC i.O. | Fehlersuche; div. Anwendungen im Serviceprogr. durchgef.; Gerät aus; Blower + Filter ausgebaut, gereinigt u. eingebaut; Gerät ein; Startphase abgewartet; QC i.O. | SWO-190443 | No | Troubleshooting, various applications in the service progr. through, device off, Blower + filter removed, cleaned and installed, device on, start-up phase awaited, QC i.O. | Troubleshooting; various applications in the service progr. through it; device off; Blower + filter removed, cleaned and installed; device; start-up phase awaited; QC in the first | Cleaned | HW | Process Module | Filter, Cup dryer | 168
ABL80 FLEX Basic version | Comprobado conexionado. Sustituida sensord boar (placa grande frontal) por la del equipo de laboratorio con el cable incluido para comprobación. | B | D2826CEEB433B7A0 | No | Proven connection. Replaced sensord boar (large front plate) with that of laboratory equipment with the included cable for testing.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-190878 | No | The replacement test card is OK. | The replacement test card is OK.
ABL90 FLEX | Touchscreen Kalibrierung durchgeführt. EWFMGR C:- commit durchgeführt. Einlasssonde ist verbogen, Sonde und Einlassdichtung aus Kundenbestand erneuert. | Touchscreen Kalibrierung durchgeführt. EWFMGR C:- commit durchgeführt. Einlasssonde ist verbogen, Sonde und Einlassdichtung aus Kundenbestand erneuert. | SWO-191256 | No | Touchscreen calibration performed. EWFMGR C:- commit performed. Inlet probe is bent, probe and inlet seal is renewed from customer stock. | Touchscreen calibration performed. EWFMGR C:- commit performed. Inlet probe is bent, probe and inlet seal is renewed from customer stock. | Adjustment | HW | DMS_PC | Touch screen
ABL90 FLEX | Gerät manuell gespült LP erneuert Gerät überprüft | Gerät manuell gespült LP erneuert Gerät überprüft | SWO-190805 | Yes | Device manually rinsed LP renewed Device checked | Device manually rinsed LP renewed Device checked | Re classified RMED | DEMO Setup
ABL90 | Printer won't print. Error code 1298, Printer lid open. Opening and closing lid doesn't work. Rebooting does not resolve issue. | Go to FedEx hold for pickup for part. Replaced printer. Also worked with Kim to get Sunquest LIS working. Needed ETXETB change and also changed low level protocol to (TCP)(ASTM). Results no crossing. | D282545DB400D1AE | No | Printer won't print. Error code 1298, Printer lid open. Opening and closing lid doesn't work. Rebooting does not resolve issue. | Go to FedEx hold for pickup for part. Replaced printer. Also worked with Kim to get Sunquest LIS working. Needed ETXETB change and also changed low level protocol to (TCP)(ASTM). Results no crossing. | Defect | HW | Printer Module | Printer Module
ABL90 FLEX | Die Einlassnadel ist verbogen - Einlassnadel erneuert. Testmessungen durchgeführt - Kapillar Proben werden wieder gemessen. Messverfahren mit BGA Probenehmer nochmal erklärt. Analysator Status grün. | Die Einlassnadel ist verbogen - Einlassnadel erneuert. Testmessungen durchgeführt - Kapillar Proben werden wieder gemessen. Messverfahren mit BGA Probenehmer nochmal erklärt. Analysator Status grün. | SWO-191022 | No | The inlet needle is bent - inlet needle renewed. Test measurements carried out - capillary samples are measured again. Measurement method with BGA sampler explained again. Analyzer status green. | The inlet needle is bent - inlet needle renewed. Test measurements carried out - capillary samples are measured again. Measurement method with BGA sampler explained again. Analyzer status green. | Replaced | HW | Inlet module | Probe bend
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-191198 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | Glu/Lac
TCM5 BASIC |  | B | D2825A3BB407F4D6 | Yes |  |  | Re classified RMED | Double entry
TCM5 BASIC |  | B | D28257D6B438F4D2 | No
ABL825 FLEX | rebooting brings you back to same code 2 error | removed roma.db, rebuilt setup, qc, LIs | D2826A17B400AA20 | Yes | rebooting brings you back to same code 2 error | removed roma.db, rebuilt setup, qc, LIs
ABL827 FLEX |  | B | D2825E0CB423AFAD | Yes |  |  | Re classified RMED | SWU
ABL827 FLEX |  | B | D2825D89B43CAFAB | Yes |  |  | Re classified RMED | Upgrade
AQT90 FLEX | Desmontagem do equipamento, limpeza interna e externa, limpeza dos sensores, da agulha, verificação do estado da bombas, verificação do funcionamento, tudo ok. | B | D282AFCAB45DAC04 | Yes | Disassembly of the equipment, internal and external cleaning, cleaning of the sensors, the needle, checking the state of the pumps, checking the operation, everything OK. |  | Re classified RMED | PM done | is scheduled in acc w prev.
AQT90 FLEX | Optical unit vervangen. | B | SWO-190753 | No | Optical Unit. |  | Defect | HW | Process Module | Optical Unit | 129
AQT90 FLEX | QC Gemessen. | QC Gemessen. | SWO-190666 | No | QC Measured. | QC Measured. | Defect | HW | Analyzer Control Module | Display
AQT90 FLEX |  | B | D282AD86B40D19BE | No |  |  | Defect | Consumables | SP | Cup Stack | 167
ABL90 | Printer roller broken on one end. | Replaced printer | D283EDC82A286D07 | No | Printer roller broken on one end. | Replaced printer | Defect | HW | Printer Module | Printer Module
ABL90 FLEX PLUS | Zeitumstellung nach RMED Anleitung durchgeführt. Gerät zum Test 2 x neu gestartet. Alles ok. | Zeitumstellung nach RMED Anleitung durchgeführt. Gerät zum Test 2 x neu gestartet. Alles ok. | SWO-190967 | No | time change according to RMED instructions. Device to test 2 x restarted. All ok. | time change according to RMED instructions. Device to test 2 x restarted. All ok. | Setup | Time setting | Time setting | Time setting
TCM CombiM | 更换945-737一个正常。(By Ran) | 更换945-737一个正常。(By Ran) | SWO-190909 | No | Replace the 945-737 one normal. (By Ran) | Replace the 945-737 one normal. (By Ran)
ABL90 FLEX | 更换903-581后修复(By Ran) | 更换903-581后修复(By Ran) | SWO-190640 | No | Repair after replacement 903-581 (By Ran) | Repair after replacement 903-581 (By Ran) | Defect | HW | Basic Unit | Sensor Interface PCB
ABL827 FLEX |  | B | D28397EF2A116DEC | Yes
ABL825 FLEX | Referenz Elektrode ersetzt und weiteren Datenextrakt gezogen. Neustart, Kalibrationen und QC Messung OK | Referenz Elektrode ersetzt und weiteren Datenextrakt gezogen. Neustart, Kalibrationen und QC Messung OK | SWO-190947 | No | Reference electrode replaced and further data extract pulled. Restart, calibrations and QC measurement OK | Reference electrode replaced and further data extract pulled. Restart, calibrations and QC measurement OK | Assisted Replacement | Electrode | Membrane | Reference
ABL825 FLEX | 无更换备件，检查到是液路到试剂下的灯，中间连接的排线破损导致短路，故处理了下排线后修复(By Ran) no spare part changed. Just the cable was broken.Fix the cable, then fix the problem. | 无更换备件，检查到是液路到试剂下的灯，中间连接的排线破损导致短路，故处理了下排线后修复(By Ran) no spare part changed. Just the cable was broken.Fix the cable, then fix the problem. | SWO-190168 | No | No replacement spare parts, check is the liquid road to the reagent under the lamp, the middle connection of the line break caused by a short circuit, so the lower row after repair (By Ran) no spare part part changed. Just the cable was broken. Fix the cable, then fix the problem. | No replacement spare parts, check is the liquid road to the reagent under the lamp, the middle connection of the line break caused by a short circuit, so the lower row after repair (By Ran) no spare part part changed. Just the cable was broken. Fix the cable, then fix the problem. | Defect | HW | Basic Unit | Cable
ABL825 FLEX |  | B | D28AE766EE131DBF | No |  |  | Cleaned | HW | pH/BG Module | Measuring chamber
ABL825 FLEX | Checked logs for errors, found 0316 Neon number too low and 0258 Neon intensity too low . Ran complete oximetry diagnostics and had failures on lamp unit. | Replaced Lamp Unit. Ran all tests with pass results. Ran tHB calibration with pass result. | D283B6EE2A54973B | No | Checked logs for errors, found 0316 Neon number too low and 0258 Neon intensity too low . Ran complete oximetry diagnostics and had failures on lamp unit. | Replaced Lamp Unit. Ran all tests with pass results. Ran tHB calibration with pass result. | Defect | HW | Oxi module | Lamp unit
AQT90 FLEX | Kalkassette geliefert | Kalkassette geliefert | SWO-191058 | Yes | Calf cassette supplied | Calf cassette supplied | Re classified RMED | Logistical upSWO
AQT90 FLEX | to replace Optical unit | to replace Optical unit | SWO-190991 | No | to replace Optical unit | to replace Optical unit | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | Gerät neu eingestellt, Schacht der Reag.cassette von alten Cups gereinigt, neue Reag.cassette eingesetzt, Gerät neu gestartet, | Gerät neu eingestellt, Schacht der Reag.cassette von alten Cups gereinigt, neue Reag.cassette eingesetzt, Gerät neu gestartet, | SWO-191153 | Yes | Device re-adjusted, shaft of Reag.cassette cleaned of old cups, new Reag.cassette inserted, device restarted, | Device re-adjusted, shaft of Reag.cassette cleaned of old cups, new Reag.cassette inserted, device restarted,
ABL80 FLEX BASIC | 更换902-972一个正常。(By Ran) | 更换902-972一个正常。(By Ran) | SWO-190915 | No | Replace 902-972 one normal. (By Ran) | Replace 902-972 one normal. (By Ran)
ABL80 FLEX BASIC | 更换902-970一个正常。(By Ran) | 更换902-970一个正常。(By Ran) | SWO-190687 | No | Replace the 902-970 one normal. (By Ran) | Replace the 902-970 one normal. (By Ran)
ABL80 FLEX BASIC | 更换902-972后修复(By Ran) | 更换902-972后修复(By Ran) | SWO-190667 | No | Repair after replacement 902-972 (By Ran) | Repair after replacement 902-972 (By Ran)
ABL90 FLEX | Sensorkassette ersetzt, Startphase abgewartet, weitere Korrekte Kal abgewartet, QK ok, weitere Ansprachefehler des pCO2-Sensors -> Info an Kunden diesen Parameter zu beobachten und ggf. Rückruf an mich | Sensorkassette ersetzt, Startphase abgewartet, weitere Korrekte Kal abgewartet, QK ok, weitere Ansprachefehler des pCO2-Sensors -> Info an Kunden diesen Parameter zu beobachten und ggf. Rückruf an mich | SWO-191308 | No | Sensor cassette replaced, start phase awaited, further correct cal waited, QK ok, further speech errors of the pCO2 sensor -> Info to customers to observe this parameter and, if necessary, call back to me | Sensor cassette replaced, start phase awaited, further correct cal waited, QK ok, further speech errors of the pCO2 sensor -> Info to customers to observe this parameter and, if necessary, call back to me | Assisted Replacement | SC | SC | pCO2
ABL90 | System will not power on. | Replaced power supply. | D284F421A027715A | No | System will not power on. | Replaced power supply. | Defect | HW | Power Supply | Power Supply
TCM CombiM | 更换903-365后修复(By Ran) | 更换903-365后修复(By Ran) | SWO-190732 | No | Repair after replacement 903-365 (By Ran) | Repair after replacement 903-365 (By Ran)
AQT90 FLEX | 重新定标正常。(By Ran) Recal, OK. | 重新定标正常。(By Ran) Recal, OK. | SWO-190619 | No | Recalibration is normal. (By Ran) Recal, OK. | Recalibration is normal. (By Ran) Recal, OK. | Assisted Replacement | Consumables | Test Kit | Cartridge, Myo
AQT90 FLEX |  | B | SWO-191248 | No |  |  | Assisted Replacement | Consumables | Test Kit | Cartridge, TNI
ABL826 FLEX |  | B | D2929DC1B21EDE5F | No |  |  | Defect | HW | Printer Module | Printer head
ABL800 BASIC | Stripped and cleaned the entire AQC module scanner and mirror | Stripped and cleaned the entire AQC module scanner and mirror | SWO-193851 | No | Stripped and cleaned the entire AQC module scanner and mirror | Stripped and cleaned the entire AQC module scanner and mirror | Cleaned | HW | Autocheck module | Barcode Scanner
ABL825 FLEX |  | B | D284E5FAA00C257C | Yes
ABL827 FLEX | 装置引き上げ | B | SWO-191438 | No | Equipment pulling up |  | Unknown | Unknown | Unknown | Unknown
ABL815 FLEX | Proteinentfernung durchgeführt, Lufteinlass gespült, Datei Archiv Kal. und QC überprüft - Kal. und QC Fehler auch bei Glu, Lac und Na angezeigt, pCO2 nochmal neu membraniert - ohne Erfolg - Status zu gering, pCO2 Elektrode E788 getauscht, im Zuge des Einsatzes Glu, Lac und Ref. neu membraniert, bei Glu und Lac fehlt der untere Dichtring - Dichtring ersetzt, alle Messkammern gereinigt, Elektrodenkontakte und Connectoren mit Glasfaserpinsel gereinigt, Einlassnadel und Einlassdichtung aus Kundenbestand getauscht - Nadel war verbogen, LS Justage und Pumpen Kalibration durchgeführt, Nullstrom Kal. für pO2, Glu und Lac durchgeführt, System- und Funktionskontrolle in Ordnung | Proteinentfernung durchgeführt, Lufteinlass gespült, Datei Archiv Kal. und QC überprüft - Kal. und QC Fehler auch bei Glu, Lac und Na angezeigt, pCO2 nochmal neu membraniert - ohne Erfolg - Status zu gering, pCO2 Elektrode E788 getauscht, im Zuge des Einsatzes Glu, Lac und Ref. neu membraniert, bei Glu und Lac fehlt der untere Dichtring - Dichtring ersetzt, alle Messkammern gereinigt, Elektrodenkontakte und Connectoren mit Glasfaserpinsel gereinigt, Einlassnadel und Einlassdichtung aus Kundenbestand getauscht - Nadel war verbogen, LS Justage und Pumpen Kalibration durchgeführt, Nullstrom Kal. für pO2, Glu und Lac durchgeführt, System- und Funktionskontrolle in Ordnung | SWO-191570 | No | Protein removal performed, air inlet flushed, file archive Kal. and QC checked - Kal. and QC error also indicated at Glu, Lac and Na, pCO2 re-membraneed - without success - status too low, pCO2 electrode E788 exchanged, in the course of using Glu, Lac and Ref. newly membraneed, at Glu and Lac the lower sealing ring is missing - sealing ring replaced, all measuring chambers cleaned, electrode contacts and connectors cleaned with glass fiber brush, inlet needle and inlet seal replaced from customer stock - needle was bent, LS adjustment and Pumps calibration carried out, zero current cal. performed for pO2, Glu and Lac, system and function control in order | Protein removal performed, air inlet flushed, file archive Kal. and QC checked - Kal. and QC error also indicated at Glu, Lac and Na, pCO2 re-membraneed - without success - status too low, pCO2 electrode E788 exchanged, in the course of using Glu, Lac and Ref. newly membraneed, at Glu and Lac the lower sealing ring is missing - sealing ring replaced, all measuring chambers cleaned, electrode contacts and connectors cleaned with glass fiber brush, inlet needle and inlet seal replaced from customer stock - needle was bent, LS adjustment and Pumps calibration carried out, zero current cal. performed for pO2, Glu and Lac, system and function control in order | User error | Fluidic pathway | Fluidic pathway | Maintenance Parts
ABL800 BASIC | 装置返却作業 | B | SWO-191503 | Yes | Equipment return work |  | Re classified RMED | OOB, although 27 days
ABL825 FLEX | Customer replaced a new PO2 electrode and membrane. PO2 fails calibrations frequently. | Replaced PH measuring chamber, air valve membrane, electrode pump tubing and upper liquid sensor. All tests are successful. | D284E91FA0417FFF | No | Customer replaced a new PO2 electrode and membrane. PO2 fails calibrations frequently. | Replaced PH measuring chamber, air valve membrane, electrode pump tubing and upper liquid sensor. All tests are successful. | Defect | HW | pH/BG Module | Measuring chamber
ABL825 FLEX | CA is out of customer range for Level3 QC. Replaced E Met electrode pump tubing, cleaned CA and Reference electrodes chambers and contacts. | All tests are successful. | D28501F9A00CBFC5 | No | CA is out of customer range for Level3 QC. Replaced E Met electrode pump tubing, cleaned CA and Reference electrodes chambers and contacts. | All tests are successful. | Cleaned | HW | El/Met Module | Measuring chamber
ABL827 FLEX | Abfallstrecke, Abfallschläuche, Anschlussstücke der Einlassdichtung, Verschlauchung von Inlet zu Abfallpumpe und Einlassdichtung erneuert. Inlet- und LS Justage. Pumpen Kal ohne Beanstandung Kalibrationen durchgeführt. QC Messung und Testmessung ohne Beanstandung Verwendete Probennehmer sind Pico50 | Abfallstrecke, Abfallschläuche, Anschlussstücke der Einlassdichtung, Verschlauchung von Inlet zu Abfallpumpe und Einlassdichtung erneuert. Inlet- und LS Justage. Pumpen Kal ohne Beanstandung Kalibrationen durchgeführt. QC Messung und Testmessung ohne Beanstandung Verwendete Probennehmer sind Pico50 | SWO-191412 | No | Waste section, waste hoses, connecting pieces of the inlet seal, hoseing from Inlet to waste pump and inlet seal renewed. Inlet and LS adjustment. Pumps Kal without complaint calibrations. QC Measurement and test measurement without complaint Samplers used are Pico50 | Waste section, waste hoses, connecting pieces of the inlet seal, hoseing from Inlet to waste pump and inlet seal renewed. Inlet and LS adjustment. Pumps Kal without complaint calibrations. QC Measurement and test measurement without complaint Samplers used are Pico50 | Defect | HW | Basic Unit | Waste tubes
ABL827 FLEX | Alle Lösungen erneuert, Membraniert, offene Fragen geklärt, Schulung durchgeführt ( Geräteverantwortlicher ) Crea in Empf. 3 mit Fehlermeldung --> ABL war eine Woche in Ruhestellung--> erneuter Besuch am 12.05.2017 Display ohne Funktion --> nach Neustart wieder i.O --> Ersatzteil wird bestellt | Elektroden über Kreuz getauscht --> Fehler 959 erscheint nicht --> beide Elektroden erneuert und neu membraniert --> Kal und QC durchgeführt i.O | SWO-191415 | No | All solutions renewed, membraned, open questions clarified, training carried out ( device manager ) Crea in Empf. 3 with error message --> ABL was a week in retirement--> another visit on 12.05.2017 Display without function --> after restart i.O --> Spare part is ordered | Electrodes exchanged over cross --> Error 959 does not appear --> both electrodes renewed and re-membraneed --> Kal and QC performed i.O | Assisted Replacement | Consumables | Solution bottle | Crea
AQT90 FLEX | need to replace vision control PCB and Cable. | need to replace vision control PCB and Cable. | SWO-192555 | No | need to replace vision control PCB and Cable. | need to replace vision control PCB and Cable. | Defect | HW | Aspiration Module | PCB, Vacu vision
AQT90 FLEX | Optical Unit en Processwheel vervangen. | B | SWO-191352 | No | Optical Unit and Processwheel replacement. |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 機器返却実施。 | OPTユニット交換 プロセスホィール交換 システムチェック確認 | SWO-194774 | No | Equipment return ed. | OPT unit replacement Process wheel exchange System check check | Defect | HW | Process Module | Optical Unit | 129
ABL80 FLEX BASIC | Sensor Cassette flushed and solution pack reset. Cleaned both. Reinstall SP and SC. Analyser serviced and calibrated properly. Handed over in a fine working condition. | Sensor Cassette flushed and solution pack reset. Cleaned both. Reinstall SP and SC. Analyser serviced and calibrated properly. Handed over in a fine working condition. | SWO-191801 | No | Sensor Cassette flushed and solution pack reset. Cleaned both. Reinstall SP and SC. Analyser serviced and calibrated properly. Handed over in a fine working condition. | Sensor Cassette flushed and solution pack reset. Cleaned both. Reinstall SP and SC. Analyser serviced and calibrated properly. Handed over in a fine working condition.
ABL90 FLEX | Datenbank des ABl geprüft --> i.O Ausdruck CC | Datenbank des ABl geprüft --> i.O Ausdruck CC | SWO-191642 | No | Database of OJ checked --> i.O Expression CC | Database of OJ checked --> i.O Expression CC | Assisted Replacement | SP | SP | pO2
ABL90 FLEX | Kal., THB kal, QCs durchgeführt, OK | Kal., THB kal, QCs durchgeführt, OK | SWO-191802 | No | Kal., THB kal, QCs durchgeführt, OK | Kal., THB kal, QCs durchgeführt, OK | Adjustment | HW | Oxi module | tHb Cal
ABL90 |  | B | D2861D1A1642831C | Yes |  |  | Re classified RMED | No fact
ABL90 FLEX | Testscan ohne Erfolg..Scanner funktioniert liest aber keine Daten. SW neu installieren erfolgreich. Scanner List jetzt auch Pat . Etiketten. Initialisierung abgewartet.QC durchgeführt und in Ordnung. | Service Einsatz | SWO-191719 | No | Test scan without success.. Scanner works but does not read data. Reinstall SW successfully. Scanner List now also Pat . Labels. initialization awaited. QC performed and in order. | Service use | Defect | SW | RMED SW | Barcode Reader
ABL90 FLEX | LP erneuert und Bügelventil gangbar gemacht Kalibration und QC durchgeführt | LP erneuert und Bügelventil gangbar gemacht Kalibration und QC durchgeführt | SWO-191590 | No | LP renewed and ironing valve made workable calibration and QC performed | LP renewed and ironing valve made workable calibration and QC performed | Defect | HW | Oxi module | Pinch Valve
ABL90 FLEX PLUS | Remplacé le kit de maintenance annuel e.o Nettoyage système fluidique e.o Calibration e.o contrôle de qualité e.o | Remplacé le kit de maintenance annuel e.o Nettoyage système fluidique e.o Calibration e.o contrôle de qualité e.o | SWO-191730 | No | Replaced the annual maintenance kit e.o Cleaning fluid system e.o Calibration e.o quality control e.o | Replaced the annual maintenance kit e.o Cleaning fluid system e.o Calibration e.o quality control e.o | Defect | HW | Basic Unit | PM Kit
ABL90 FLEX PLUS |  | B | D28A19A2782F785B | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL835 FLEX | CHEQUEO DE ANALYZER CONTROL PCB, TENSIONES DE LA POWER SUPPLY Y VENTILADOR INLET SIENDO TODOO OK Y PERSISTIENDO EL PROBLEMA. SE DESMONTA CADA UNO DE LOS 4 MICRO FUSIBLES Y SE LIMPIAN LOS CONTACTOS DE ÉSTOS. REINICIO DE EQUIPO Y SOLUCIÓN DEL TEMA. | LIMPIEZA DE MICROFUSIBLES | D28632481661D62A | No | ANALYZER CONTROL PCB CHECK, POWER SUPPLY TENSIONS AND INLET FAN BEING ALL OK AND PERSISTING THE PROBLEM. EACH OF THE 4 MICRO FUSIBLES IS DISMOUNTED AND THE CONTACTS OF THESE ARE CLEANED. RESTARTING OF EQUIPMENT AND SOLUTION OF THE THEME. | CLEANING MICROFUSIBLE | Defect | HW | Basic Unit | Analyzer control PCB
ABL800 BASIC | Clメンブラン取り付け直し ｐHシーリングリング取り付け直し タンパク除去 エアライン・測定チャンバー清掃 LS調整・ポンプCAL確認 CAL・QC測定データ確認 | Clメンブラン取り付け直し ｐHシーリングリング取り付け直し タンパク除去 エアライン・測定チャンバー清掃 LS調整・ポンプCAL確認 CAL・QC測定データ確認 | SWO-191877 | No | Cl Membrane Reinstallation Reseating the pH sealing ring Protein removal Airline measurement chamber cleaning LS adjustment and pump CAL confirmation Cal/QC measurement data confirmation | Cl Membrane Reinstallation Reseating the pH sealing ring Protein removal Airline measurement chamber cleaning LS adjustment and pump CAL confirmation Cal/QC measurement data confirmation | Cleaned | HW | pH/BG Module | Measuring chamber
ABL800 BASIC | pH電極交換 エアライン・測定チャンバー清掃 LS調整・ポンプCAL確認 タンパク除去、Glu・Lac・Refメンブラン交換 CAL・QC測定データ確認 | B | SWO-191876 | No | pH electrode replacement Airline measurement chamber cleaning LS adjustment and pump CAL confirmation Protein removal, Glu Lac, Ref membrane replacement Cal/QC measurement data confirmation |  | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Neon lámpa csere után megszűnt a hibaüzenet. Oxi tesztek, kalibrálás, QC mérés. | B | SWO-191717 | No | After a Neon lamp replacement, the error message is resolved. Oxy tests, calibration, QC measurement. |  | Defect | HW | Oxi module | Lamp unit | Neon
ABL827 FLEX | pCO2 und pO2 neu membraniert, Nullstrom Kal. für pO2, Glu, Lac und Crea durchgeführt, Messkammern gereinigt, Lufteinlass gespült, Einlassnadel überprüft - ist gerade, System- und Funktionskontrolle in Ordnung | pCO2 und pO2 neu membraniert, Nullstrom Kal. für pO2, Glu, Lac und Crea durchgeführt, Messkammern gereinigt, Lufteinlass gespült, Einlassnadel überprüft - ist gerade, System- und Funktionskontrolle in Ordnung | SWO-191649 | No | pCO2 and pO2 re-membraneed, zero current Cal. for pO2, Glu, Lac and Crea performed, measuring chambers cleaned, air inlet flushed, inlet needle checked - is straight, system and function control in order | pCO2 and pO2 re-membraneed, zero current Cal. for pO2, Glu, Lac and Crea performed, measuring chambers cleaned, air inlet flushed, inlet needle checked - is straight, system and function control in order | Assisted Replacement | Electrode | Membrane | pCO2
ABL827 FLEX | Elektroden über Kreuz getauscht --> Fehler 959 erscheint nicht --> beide Elektroden erneuert und neu membraniert --> Kal und QC durchgeführt i.O | Elektroden über Kreuz getauscht --> Fehler 959 erscheint nicht --> beide Elektroden erneuert und neu membraniert --> Kal und QC durchgeführt i.O | SWO-191641 | No | Electrodes exchanged over cross --> Error 959 does not appear --> both electrodes renewed and re-membraneed --> Kal and QC performed i.O | Electrodes exchanged over cross --> Error 959 does not appear --> both electrodes renewed and re-membraneed --> Kal and QC performed i.O | User error | Electrode | Electrode | Crea mixed | Email
ABL835 FLEX | Had Keith Clark on site and looked through the resulst for qc. Found too great of varience in results. Ordered new spectrometer with fiber optic cable and replaced. Powered up and had neon voltage too low errors on spectrometer. DOA. Replaced oxi complete with old analyzer from 693N004 with customers permission and started up fine. Matched lamp to new unite. tHb at 555 pA. Other analyzer at 560 pA. Ran all levels of controls. Level 2 FCOHb not passing at 3.2. Analyzer range set to 2.0-2.8. Radometer range 1-5 with mean at 3. Lot 147. All others pass. Spoke with operator and no manager on duty. Rouella said not to adjust ranges and to make sure the qc passes. Not able to pass level 2. Ordered oxi complete for monday delivery. Analyzer oxi down. | Ordered oxi complete for Monday morning. will come in and install. Asked operator to disable oxi tests until return. | D286596216363552 | No | Had Keith Clark on site and looked through the resulst for qc. Found too great of varience in results. Ordered new spectrometer with fiber optic cable and replaced. Powered up and had neon voltage too low errors on spectrometer. DOA. Replaced oxi complete with old analyzer from 693N004 with customers permission and started up fine. Matched lamp to new unite. tHb at 555 pA. Other analyzer at 560 pA. Ran all levels of controls. Level 2 FCOHb not passing at 3.2. Analyzer range set to 2.0-2.8. Radometer range 1-5 with mean at 3. Lot 147. All others pass. Spoke with operator and no manager on duty. Rouella said not to adjust ranges and to make sure the qc passes. Not able to pass level 2. Ordered oxi complete for monday delivery. Analyzer oxi down. | Ordered Oxi complete for Monday morning. Will come in and install. Asked operator to disable OXI tests until return. | Revisit required | HW | Oxi module | Oxi module
ABL817 FLEX | Proteinentfernung durchgeführt, neues AutoCheck Modul (R2709N001) gegen Leihmodul ausgetauscht, AutoCheck Modul im Service überprüft, mehrere Kal. durchgeführt, tHb Kal. durchgeführt, QC Messungen gestartet, AutoCheck Modul konditioniert Ampullen wie gewohnt, System- und Funktionskontrolle in Ordnung | Proteinentfernung durchgeführt, neues AutoCheck Modul (R2709N001) gegen Leihmodul ausgetauscht, AutoCheck Modul im Service überprüft, mehrere Kal. durchgeführt, tHb Kal. durchgeführt, QC Messungen gestartet, AutoCheck Modul konditioniert Ampullen wie gewohnt, System- und Funktionskontrolle in Ordnung | SWO-191691 | No | Protein removal performed, new AutoCheck module (R2709N001) replaced with rental module, AutoCheck module checked in service, several cal. tHb Kal. performed, QC measurements started, AutoCheck module conditioned ampoules as usual, system and function control in order | Protein removal performed, new AutoCheck module (R2709N001) replaced with rental module, AutoCheck module checked in service, several cal. tHb Kal. performed, QC measurements started, AutoCheck module conditioned ampoules as usual, system and function control in order | Defect | HW | Autocheck module | Autocheck module
AQT90 FLEX | İĞNE DEĞİŞTİRİLEREK SORUN GİDERİLDİ. | İĞNE DEĞİŞTİRİLEREK SORUN GİDERİLDİ. | SWO-191710 | Yes | THE PROBLEM WAS ADDRESSED BY CHANGING THE NEEDLE. | THE PROBLEM WAS ADDRESSED BY CHANGING THE NEEDLE.
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-191816 | Yes | Replace the test card | Replace the test card
ABL90 FLEX |  | B | SWO-191790 | No |  |  | Defect | HW | Basic Unit | Actuator Unit
ABL90 FLEX | 更换910-342后修复 by andrew | 更换910-342后修复 by andrew | SWO-191191 | No | Repair by andrew after replacing 910-342 | Repair by andrew after replacing 910-342 | Defect | HW | DMS_PC | USER INTERFACE MODULE
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-192939 | No | The replacement test card is OK. | The replacement test card is OK. | Assisted Replacement | SC | SC | pCO2
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-192675 | No | Replace the test card | Replace the test card
ABL80 FLEX BASIC | Ausbau der Flüsdigkeits-Verschlauchung und Erneuerung durch einen Jahreswartungs-Kit. Abfalltrichter gereinigt. Flüssigkeit läuft gut ab, wird gut abgesaugt. | Ausbau der Flüsdigkeits-Verschlauchung und Erneuerung durch einen Jahreswartungs-Kit. Abfalltrichter gereinigt. Flüssigkeit läuft gut ab, wird gut abgesaugt. | SWO-192153 | No | Removal of fluid hose and renewal by an annual maintenance kit. Waste hopper cleaned. Liquid runs well, is well extracted. | Removal of fluid hose and renewal by an annual maintenance kit. Waste hopper cleaned. Liquid runs well, is well extracted.
ABL80 FLEX BASIC | Found that the Sensor connector board is damaged, tried installing new SC, same error was observed, the machine is displaying SC temp not stable even if no SC is connected. Required Sensor Connector Board as well as Signal Data Board for testing purpose. | Found that the Sensor connector board is damaged, tried installing new SC, same error was observed, the machine is displaying SC temp not stable even if no SC is connected. Required Sensor Connector Board as well as Signal Data Board for testing purpose. | SWO-192040 | No | Found that the Sensor connector board is damaged, tried installing new SC, same error was observed, the machine is displaying SC temp not stable even if no SC is connected. Required Sensor Connector Board as well as Signal Data Board for testing purpose. | Found that the Sensor connector board is damaged, tried installing new SC, same error was observed, the machine is displaying SC temp not stable even if no SC is connected. Required Sensor Connector Board as well as Signal Data Board for testing purpose.
ABL90 FLEX | Kal wiederholt i.O. fehlerhafte QC Level A wiederholt Kal Fehler bei pCO2 während QC Messung SC Fach im Serviceprog. geöffnet, es scheint es seine die Pins nicht alle auf einer höhe, Sensor Interface PCB bestellt ->info an Kunden rep. am Mittwoch | B | SWO-192057 | No | Kal repeats I.O. faulty QC Level A repeats Kal error in pCO2 during QC measurement SC Fach in the service prog. open, it seems its the pins not all ordered on a height, sensor interface PCB ->info to customers rep. on Wednesday |  | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX | ABL90 durchgespült und Teile im Umfang eines Jahreskits erneuert. Einlasssonde zur Bevorratung übergeben. ( Einlasssonde berechnen ) | ABL90 durchgespült und Teile im Umfang eines Jahreskits erneuert. Einlasssonde zur Bevorratung übergeben. ( Einlasssonde berechnen ) | SWO-192234 | No | ABL90 rinsed and parts in the scope of an annual kit renewed. inlet probe for stockpiling. ( Calculate inlet probe ) | ABL90 rinsed and parts in the scope of an annual kit renewed. inlet probe for stockpiling. ( Calculate inlet probe ) | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D289E54A78621B8E | No |  |  | Defect | HW | Oxi module | Spectrometer
ABL90 FLEX |  | B | SWO-191962 | No |  |  | Defect | HW | Basic Unit | Solution Pack ID PCB | 1134
AQT90 FLEX |  | B | SWO-191854 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
ABL827 FLEX |  | B | D28A10EE781E3946 | Yes
ABL827 FLEX | Display-unit defekt ( keine Bedienung möglich Touchscreen funktioniert nicht ) --> Display unit erneuert | Display-unit defekt ( keine Bedienung möglich Touchscreen funktioniert nicht ) --> Display unit erneuert | SWO-191984 | No | Display-unit defective ( no operation possible Touchscreen does not work ) --> Display unit renewed | Display-unit defective ( no operation possible Touchscreen does not work ) --> Display unit renewed | Defect | HW | DMS_PC | Touch screen
ABL835 FLEX |  | Replaced the oximetry. Lamp adjustment set to 3250. No need to adjust. Ran complete oxi test and all passed. Ran calibration and all passed. Range for FCOHb level 2 qc 2-2.8. Peer data at 2.5-3.3. Customer will adjust range and establish a mean. Ran level 2 qc as a precision at 3, 3, 3.1 for FCOHb. Will return tomorrow after analyzer settles at run VK-R7. | D28A5D767848BF27 | No |  | Replaced the oximetry. Lamp adjustment set to 3250. No need to adjust. Ran complete oxi test and all passed. Ran calibration and all passed. Range for FCOHb level 2 qc 2-2.8. Peer data at 2.5-3.3. Customer will adjust range and establish a mean. Ran level 2 qc as a precision at 3, 3, 3.1 for FCOHb. Will return tomorrow after analyzer settles at run VK-R7. | Defect | HW | Oxi module | Oxi module | FCOHb
AQT90 FLEX | Scanner neu justiert Fax 05861838099 | Scanner neu justiert Fax 05861838099 | SWO-192168 | No | Scanner re-adjusted Fax 05861838099 | Scanner re-adjusted Fax 05861838099 | Adjustment | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | Fehlersuche, div. Sicherheitseinst. angepaßt, Anwender klärt Zugriffsrechte mit Labor Limbach | Fehlersuche; div. Sicherheitseinst. angepaßt; Anwender klärt Zugriffsrechte mit Labor Limbach | SWO-192002 | No | Troubleshooting, various security one. adapted, user clarifies access rights with Laboratory Limbach | Troubleshooting; div. Security one. adapted; User clarifies access rights with Limbach laboratory | User error | SW | User access | User access
ABL80 FLEX BASIC | Kept in fridge for 30 minutes and reinstalled sensor cassette. But still it is showing Lac sensitivity low. Need to change SC. | Kept in fridge for 30 minutes and reinstalled sensor cassette. But still it is showing Lac sensitivity low. Need to change SC. | SWO-193205 | No | Kept in fridge for 30 minutes and reinstalled sensor cassette. But still it is showing Lac sensitivity low. Need to change SC. | Kept in fridge for 30 minutes and reinstalled sensor cassette. But still it is showing Lac sensitivity low. Need to change SC.
ABL80 CO-OX | pink fluid all over counter. Paged FSE | B | D28B0498EE3DD0A9 | Yes | pink fluid all over counter. Paged FSE
ABL90 |  | B | D28B67D7EE15A2FF | Yes
ABL90 |  | B | D28B67D8EE31DBE2 | Yes
ABL90 | checked out the report and found that they failed 3 out of 5 for THB | Open techline case, ran THB cal, preformed complete oxi test, Preformed Linearity for THB, and grabbed Service Dump and data logs for Techline case. | D28B4B86EE470983 | No | checked out the report and found that they failed 3 out of 5 for THB | Open techline case, ran THB cal, preformed complete oxi test, Preformed Linearity for THB, and grabbed Service Dump and data logs for Techline case. | Adjustment | HW | Oxi module | tHb Cal
ABL90 | Pinch valve not opening closing. | Replaced Pinch valve and tubing for valve with ring. | D28B37C0EE54E22C | No | Pinch valve not opening closing. | Replaced Pinch valve and tubing for valve with ring. | Defect | HW | Oxi module | Pinch valve | 1184
ABL90 FLEX |  | B | SWO-192188 | No |  |  | Defect | HW | Power Supply | Power Supply
AQT90 FLEX |  | B | SWO-192282 | No |  |  | Defect | Consumables | Test Kit | Cartridge | 5784
ABL835 FLEX | nach Neustart Display wieder in Funktion, vermutlich CPU Unit defekt, ET bestellt | Der Kunde wurde informiert, den Bildschirmrand zu reinigen. Der Kunde lehnt ab. Der Kunde bittet um einen Service. | SWO-192451 | No | after restart display back in function, presumably CPU unit defective, ET ordered | The customer has been informed to clean the edge of the screen. The customer refuses. The customer asks for a service. | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX | Touch screen is not out of calibration but is not responding correctly. The number 6 and letter G are populating when they should not be and scroll arrows are not working. PO2 zero errors found when checking the logs for performance. | Power down and restarted the analyzer and tested with Mike. All functions that had issues are now responding as they should. Performed the zero procedure in service and ran calibrations to confirm. | D28B3F7DEE5374B8 | No | Touch screen is not out of calibration but is not responding correctly. The number 6 and letter G are populating when they should not be and scroll arrows are not working. PO2 zero errors found when checking the logs for performance. | Power down and restarted the analyzer and tested with Mike. All functions that had issues are now responding as they should. Performed the zero procedure in service and ran calibrations to confirm. | Adjustment | HW | DMS_PC | Touch screen
ABL825 FLEX |  | B | D28D9053DA427BDC | No |  |  | Revisit required | HW | Oxi module | tHb Cal | Air in Fluidic Pathway
ABL820 FLEX | Fails PCO2 on level 1&2 QC on the lower range of packet insert. PCO2 electrode was remembraned and sill the same issue. | Remembraned with a different lot#R0795. Replaced PH electrode seal ring. Upgraded main software to version 6.17 Will order a new PCO2 electrode under warranty. | D28B751CEE45CE1C | No | Fails PCO2 on level 1&2 QC on the lower range of packet insert. PCO2 electrode was remembraned and sill the same issue. | Remembraned with a different lot#R0795. Replaced PH electrode seal ring. Upgraded main software to version 6.17 Will order a new PCO2 electrode under warranty. | Assisted Replacement | Electrode | Membrane | pCO2 | pH
ABL825 FLEX | Glu installiert. Mehrfavh kalibriert. | Glu installiert. Mehrfavh kalibriert. | SWO-192449 | Yes | Glu installed. Multifavh calibrated. | Glu installed. Multifavh calibrated. | Re classified RMED | Glu missing
AQT90 FLEX | Tausch der Optischen Einheit und des Process Wheels. Überprüfung ob Zentral gemessen wird. Werte besser wie vorher. Neukalibration der beiden Parameter TnT und D- Dimer. Testkits aus der Geschäftsstelle vorher geholt. Qualitätskontrolle außer Grenzen und höher gemessen als vorher. Einbau der alten Optischen Einheit und Ausbau der neuen. Erneute Kalibration der beiden Parameter. Qualitätskontroll Werte wieder in Ordnung. | Tausch der Optischen Einheit und des Process Wheels. Überprüfung ob Zentral gemessen wird. Werte besser wie vorher. Neukalibration der beiden Parameter TnT und D- Dimer. Testkits aus der Geschäftsstelle vorher geholt. Qualitätskontrolle außer Grenzen und höher gemessen als vorher. Einbau der alten Optischen Einheit und Ausbau der neuen. Erneute Kalibration der beiden Parameter. Qualitätskontroll Werte wieder in Ordnung. | SWO-192506 | No | Replacement of the Optical Unit and the Process Wheel. Check whether central is measured. Values better than before. Recalibration of the two parameters TnT and D- Dimer. Test kits from the office beforehand. Quality control out of bounds and measured higher than before. Installation of the old optical unit and extension of the new one. Re-calibration of the two parameters. Quality control values back in order. | Replacement of the Optical Unit and the Process Wheel. Check whether central is measured. Values better than before. Recalibration of the two parameters TnT and D- Dimer. Test kits from the office beforehand. Quality control out of bounds and measured higher than before. Installation of the old optical unit and extension of the new one. Re-calibration of the two parameters. Quality control values back in order. | Adjustment | HW | Process Module | Process wheel
AQT90 FLEX | Techniker Kollege hat noch Testmessungen durchgeführt um den Zustand der Optischen Einheit zu prüfen. Tausch von Hardware heute aus Zeitgründen nicht mehr möglich. | Techniker Kollege hat noch Testmessungen durchgeführt um den Zustand der Optischen Einheit zu prüfen. Tausch von Hardware heute aus Zeitgründen nicht mehr möglich. | SWO-192507 | No | Technician colleague has still carried out test measurements to check the condition of the optical unit. Replacement of hardware is no longer possible due to time constraints. | Technician colleague has still carried out test measurements to check the condition of the optical unit. Replacement of hardware is no longer possible due to time constraints. | Revisit required | HW | Process Module | Optical Unit | 211
AQT90 FLEX | Blutbilder gereinigt, Schlauchsystem mit Hypochlorid gespült, Sensoren gereinigt, Piercingtool und Wash-Unit gereinigt, Processwheel gereinigt 5 Blankmessungen - ok Photocounts ok 3 mal Level 1 LQC TnI - ok Level 2 und Level 3 LQC - TnI ok | Blutbilder gereinigt, Schlauchsystem mit Hypochlorid gespült, Sensoren gereinigt, Piercingtool und Wash-Unit gereinigt, Processwheel gereinigt 5 Blankmessungen - ok Photocounts ok 3 mal Level 1 LQC TnI - ok Level 2 und Level 3 LQC - TnI ok | SWO-192533 | No | Blood images cleaned, hose system flushed with hypochloride, sensors cleaned, piercing tool and wash unit cleaned, processwheel cleaned 5 blank measurements - ok Photocounts ok 3 times Level 1 LQC TnI - ok Level 2 and Level 3 LQC - TnI ok | Blood images cleaned, hose system flushed with hypochloride, sensors cleaned, piercing tool and wash unit cleaned, processwheel cleaned 5 blank measurements - ok Photocounts ok 3 times Level 1 LQC TnI - ok Level 2 and Level 3 LQC - TnI ok | Cleaned | HW | Process Module | Process wheel
ABL80 FLEX BASIC | 连续更换2张测试卡后修复 | 连续更换2张测试卡后修复 | SWO-193278 | Yes | Repair after 2 consecutive test cards | Repair after 2 consecutive test cards
ABL80 CO-OX | Ran zero calibration test and passed, checked sensor connector board for damage and none was found. | Replaced the analog board sensor connector board. | D28C78566460B0F7 | No | Ran zero calibration test and passed, checked sensor connector board for damage and none was found. | Replaced the analog board sensor connector board.
ABL80 FLEX BASIC | 更换测试卡 | 更换测试卡 | SWO-192599 | No | Replace the test card | Replace the test card
ABL90 FLEX | Ein und ausschalten des Anlysators | Ein und ausschalten des Anlysators | SWO-192835 | No | Switching the analyzer on and off | Switching the analyzer on and off | Assisted Replacement | SC | SC | Thermistor
ABL90 FLEX | Replace Oxi pinch valve tube. | Replace Oxi pinch valve tube. | SWO-192829 | No | Replace Oxygen Pinch valve tube. | Replace Oxygen Pinch valve tube. | Defect | HW | Oxi module | Pinch valve tube | 1342
ABL90 FLEX | Bei eintreffen ist Radianceverbindung nicht aktiv-> Haken für Kommunikationsverbindung entfernt gesetzt -> keine Verbessung -> Haus EDV nicht erreichbar bemängelte Barcodes eingescannt-> werden nicht erkannt, Druck des Barcodes auf dem Etikettenrand, externe Scanner angeschlossen -> Barcodes werden nicht erkannt -> Info an Station das die Qualität der Barcoddruck zu prüfen ist-> Rücksprache mit Hr. Gebhard in EvB Potsdam Gerät neu gestartet, Netzwerkabel geprüft-> Verbindung in Netzwerkdose lose -> befestigt ( Wastebehälter steht davor und drückt dagegen)-> Info an Stat.leitung das Kabel durch Haus EDV ersetzt werden soll, Pat ID Code per Hand eingegeben -> Daten kommen im Soaris an Labor: Telefonat mit Fr. Roland: Problem der Qualität der Barcodeetiketten ist bekannt und soll behoben werden | Bei eintreffen ist Radianceverbindung nicht aktiv-> Haken für Kommunikationsverbindung entfernt gesetzt -> keine Verbessung -> Haus EDV nicht erreichbar bemängelte Barcodes eingescannt-> werden nicht erkannt, Druck des Barcodes auf dem Etikettenrand, externe Scanner angeschlossen -> Barcodes werden nicht erkannt -> Info an Station das die Qualität der Barcoddruck zu prüfen ist-> Rücksprache mit Hr. Gebhard in EvB Potsdam Gerät neu gestartet, Netzwerkabel geprüft-> Verbindung in Netzwerkdose lose -> befestigt ( Wastebehälter steht davor und drückt dagegen)-> Info an Stat.leitung das Kabel durch Haus EDV ersetzt werden soll, Pat ID Code per Hand eingegeben -> Daten kommen im Soaris an Labor: Telefonat mit Fr. Roland: Problem der Qualität der Barcodeetiketten ist bekannt und soll behoben werden | SWO-192731 | No | When ireads arrive, the radiance connection is not active-> hook for communication connection removed -> no improvement -> house computer not reachable criticized barcodes scanned-> are not recognized, printing of the barcode on the label edge, external scanners are connected -> barcodes are not recognized -> Info to station that the quality of barcod printing is to be checked-> Consultation with Mr. Gebhard in EvB Potsdam Device restarted, network-abel checked-> connection in network socket loose -> fixed (waste container stands in front of it and presses against it)-> Info to Stat.leitung the cable should be replaced by house COMPUTER, Pat ID code entered by hand -> Data arrive in the Soaris Laboratory: Phone call with Ms. Roland: Problem of the quality of the barcode labels is known and should be fixed | When ireads arrive, the radiance connection is not active-> hook for communication connection removed -> no improvement -> house computer not reachable criticized barcodes scanned-> are not recognized, printing of the barcode on the label edge, external scanners are connected -> barcodes are not recognized -> Info to station that the quality of barcod printing is to be checked-> Consultation with Mr. Gebhard in EvB Potsdam Device restarted, network-abel checked-> connection in network socket loose -> fixed (waste container stands in front of it and presses against it)-> Info to Stat.leitung the cable should be replaced by house COMPUTER, Pat ID code entered by hand -> Data arrive in the Soaris Laboratory: Phone call with Ms. Roland: Problem of the quality of the barcode labels is known and should be fixed | User error | HW | Cable | Cable
ABL90 FLEX | Date and Time Einstellungen auf Windows-Ebene angepasst, Speicherbefehl ausgeführt, 2 Neustart des Analysators durchgeführt --> keine Zeitverstellung mehr feststellbar, System und Funktion i.O. | Date and Time Einstellungen auf Windows-Ebene angepasst, Speicherbefehl ausgeführt, 2 Neustart des Analysators durchgeführt --> keine Zeitverstellung mehr feststellbar, System und Funktion i.O. | SWO-192664 | No | Date and Time settings adapted at the Windows level, memory command executed, 2 restart of the analyzer performed --> no time adjustment can be determined anymore, system and function i.O. | Date and Time settings adapted at the Windows level, memory command executed, 2 restart of the analyzer performed --> no time adjustment can be determined anymore, system and function i.O. | Setup | SW | Time setting | Time setting
ABL90 FLEX | Date and Time Einstellungen auf Windows-Ebene angepasst, Speicherbefehl ausgeführt, 2 Neustart des Analysators durchgeführt --> keine Zeitverstellung mehr feststellbar, System und Funktion i.O. | Date and Time Einstellungen auf Windows-Ebene angepasst, Speicherbefehl ausgeführt, 2 Neustart des Analysators durchgeführt --> keine Zeitverstellung mehr feststellbar, System und Funktion i.O. | SWO-192663 | No | Date and Time settings adapted at the Windows level, memory command executed, 2 restart of the analyzer performed --> no time adjustment can be determined anymore, system and function i.O. | Date and Time settings adapted at the Windows level, memory command executed, 2 restart of the analyzer performed --> no time adjustment can be determined anymore, system and function i.O. | Setup | SW | Time setting | Time setting
ABL90 | Power supply LEDs not all on. Battery voltage higher than warning threshold. | Replaced power supply and back-up battery pack. | D28C7D8B6401C76F | No | Power supply LEDs not all on. Battery voltage higher than warning threshold. | Replaced power supply and back-up battery pack. | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | Netzwerkeinstellungen, Speicherbefehl und die Startphase durchgeführt. | Netzwerkeinstellungen, Speicherbefehl und die Startphase durchgeführt. | SWO-192822 | No | network settings, memory command, and the boot phase. | network settings, memory command, and the boot phase. | Setup | SW | His/lis | His/lis
ABL90 FLEX | Fließwege gespült, Startphase durchlaufen. Ventilfunktion im Service überprüft, Ventil leichtgängig. Mehrere Testmessungen durchgeführt, Funktion OK. | Fließwege gespült, Startphase durchlaufen. Ventilfunktion im Service überprüft, Ventil leichtgängig. Mehrere Testmessungen durchgeführt, Funktion OK. | SWO-192920 | No | Flow paths rinsed, go through the start phase. Valve function checked in service, valve smooth. Several test measurements performed, function OK. | Flow paths rinsed, go through the start phase. Valve function checked in service, valve smooth. Several test measurements performed, function OK. | Cleaned | HW | Oxi module | Fluidic pathway
ABL90 FLEX PLUS | Mixer Tray erneuert. | Mixer Tray erneuert. | SWO-192787 | No | Mixer Tray renewed. | Mixer Tray renewed. | Defect | HW | Basic Unit | Mixer Cover
ABL90 FLEX | 重新校准触摸屏后修复 by andrew | 重新校准触摸屏后修复 by andrew | SWO-192556 | No | Fix by andrew after recalibrating the touch screen | Fix by andrew after recalibrating the touch screen | Adjustment | HW | DMS_PC | Touch screen
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-195590 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Lac
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-195591 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Thermistor | 165
AQT90 FLEX |  | B | SWO-192607 | No |  |  | Revisit required | HW | Analyzer Control Module | Barcode reader, display | 1218
AQT90 FLEX |  | B | SWO-192547 | No |  |  | Defect | HW | Process Module | Optical Unit | 124
ABL835 FLEX | TRAS RECIBIR MATERIAL SE PROCEDE A CHECHEO DE MÓDULO PC, MEDICIÓN DE TENSIONES DE LA POWER SUPPLY, CABLES DE CONEXIÓN PC-AU, ANALYZER CONTROL PCB, FUSIBLES, Y CONEXIONES Y CABLES EN GENERAL DESCONECTANADO Y CONECTANDO TODOS DE NUEVO, NO SE OBSERVA NADA ANÓMALO. SE SUSTITUYEN LOS 4 FUSIBLES Y SE VERIFICAN Y REFUERZAN LOS CONTACTOS EN LA ANALYZER CONTROL PCB, SE VERIFICA CONTINUIDAD EN CONTACTOS, SE SUSTITUYE CABLE DE COMUNICACIÓN DE PC UNIT-ANALIZER CONTROL.. | SUSTITUCIÓN DE MICRO FUSIBLES Y CABLE DE COMUNICACIÓN MÓDULO PC - ANALYZER COONTROL REFUERZO DE CONTACTOS PORTAFUSIBLES | D28C768E6405D63F | No | AFTER RECEIVING MATERIAL, PC MODULE REGULATION, POWER SUPPLY TENSION MEASURE, PC-AU CONNECTION CABLES, ANALYZER CONTROL PCB, FUSIBLE, AND GENERAL CONNECTIONS AND CABLES ARE AVAILABLE AND CONNECTED ALL NEW, NOTHING IS NOT OBSERVED. THE 4 FUSIBLES ARE SUBSTITUTED AND THE CONTACTS ARE CHECKED AND REFURBED IN THE ANALYZER CONTROL PCB, CONTINUITY IS VERIFIED IN CONTACTS, PC UNIT-ANALIZER CONTROL COMMUNICATION CABLE IS SUBSTITUTE. | SUBSTITUTE OF MICRO FUSIBLES AND COMMUNICATION CABLE MODULE PC - ANALYZER COONTROL REFUND OF PORTAFUSIBLE CONTACTS | Revisit required | HW | Basic Unit | Analyzer control PCB | Email
ABL825 FLEX | felszívó tű cseréje, bemeneti egység tisztítása, LS és pumpa kalibrációk, QC mérés | B | SWO-192657 | No | Suction needle replacement, input unit cleaning, LS and pump calibrations, QC measurement |  | Defect | HW | Inlet module | Probe bend
ABL825 FLEX | Betriebssystem neu installiert, Analysatorsoftware installiert, Backup zurückgespielt, keine Sound hörbar, Lautsprecher ausgebaut - Kabel an Stecker angelötet, wieder zusammengesetzt, Neustart abgewartet, mehrere Kal, Sound eingestellt, QC | Betriebssystem neu installiert, Analysatorsoftware installiert, Backup zurückgespielt, keine Sound hörbar, Lautsprecher ausgebaut - Kabel an Stecker angelötet, wieder zusammengesetzt, Neustart abgewartet, mehrere Kal, Sound eingestellt, QC | SWO-192700 | No | Operating system reinstalled, analyzer software installed, backup backplayed, no sound audible, speakers removed - cable soldered to plug, reassembled, restart awaited, multiple kal, sound set, QC | Operating system reinstalled, analyzer software installed, backup backplayed, no sound audible, speakers removed - cable soldered to plug, reassembled, restart awaited, multiple kal, sound set, QC | Defect | HW | DMS_PC | Loudspeaker
ABL800 BASIC | pH電極交換 タンパク除去 エアライン・測定チャンバー清掃 LS調整・ポンプCAL確認 CAL・QC測定データ確認 | pH電極交換 タンパク除去 エアライン・測定チャンバー清掃 LS調整・ポンプCAL確認 CAL・QC測定データ確認 | SWO-193961 | No | pH electrode replacement Protein removal Airline measurement chamber cleaning LS adjustment and pump CAL confirmation Cal/QC measurement data confirmation | pH electrode replacement Protein removal Airline measurement chamber cleaning LS adjustment and pump CAL confirmation Cal/QC measurement data confirmation | Assisted Replacement | Electrode | Electrode | pH | email
ABL800 BASIC | pH BGベースプレート交換 タンパク除去 エアライン・測定チャンバー清掃 LS調整・ポンプCAL確認 CAL・QC測定データ確認 | pH BGベースプレート交換 タンパク除去 エアライン・測定チャンバー清掃 LS調整・ポンプCAL確認 CAL・QC測定データ確認 | SWO-193952 | No | pH BG base plate replacement Protein removal Airline measurement chamber cleaning LS adjustment and pump CAL confirmation Cal/QC measurement data confirmation | pH BG base plate replacement Protein removal Airline measurement chamber cleaning LS adjustment and pump CAL confirmation Cal/QC measurement data confirmation | Assisted Replacement | Electrode | Electrode | pH | Email
ABL835 FLEX | Lac nach Service Tipp gesäubert und mit Proteinentferner behandelt, Elektrode gespült und neu bezogen (LOT 3488) mehrfach 1 Punkt-Kalibrierung durchgeführt, Empfindlichkeit deutlich höher als zuvor aber Empf. springt z.Z. noch, Kunde beobachtet weiteres Verhalten | Lac nach Service Tipp gesäubert und mit Proteinentferner behandelt, Elektrode gespült und neu bezogen (LOT 3488) mehrfach 1 Punkt-Kalibrierung durchgeführt, Empfindlichkeit deutlich höher als zuvor aber Empf. springt z.Z. noch, Kunde beobachtet weiteres Verhalten | SWO-192722 | No | Lac cleaned according to service tip and treated with protein remover, electrode rinsed and re-sourced (LOT 3488) performed several times 1 point calibration, sensitivity significantly higher than before but empf. is still jumping at the time, customer is watching further behavior | Lac cleaned according to service tip and treated with protein remover, electrode rinsed and re-sourced (LOT 3488) performed several times 1 point calibration, sensitivity significantly higher than before but empf. is still jumping at the time, customer is watching further behavior | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
AQT90 FLEX | Ausbau der Optischen Einheit und des Process Wheels und Einbau einer neuen Optischen Einheit und eines neuen Procees Wheels. Alte Kalibrationen gelöscht und neu kalibriert. Kalibration Werte in Ordnung, Qualitätskontrolle wieder zielwertig. | Ausbau der Optischen Einheit und des Process Wheels und Einbau einer neuen Optischen Einheit und eines neuen Procees Wheels. Alte Kalibrationen gelöscht und neu kalibriert. Kalibration Werte in Ordnung, Qualitätskontrolle wieder zielwertig. | SWO-192761 | No | Expansion of the optical unit and the process wheel and installation of a new optical unit and a new Procees Wheel. Old calibrations deleted and recalibrated. Calibration values in order, quality control again targeted. | Expansion of the optical unit and the process wheel and installation of a new optical unit and a new Procees Wheel. Old calibrations deleted and recalibrated. Calibration values in order, quality control again targeted. | Revisit required | HW | Process Module | Optical Unit | 211
ABL80 FLEX CO-OX OSM | Daten auf einen USB Stick gesichert und am Gerät gelöscht, | Daten auf einen USB Stick gesichert und am Gerät gelöscht, | SWO-193026 | No | data backed up to a USB stick and deleted from the device, | data backed up to a USB stick and deleted from the device,
ABL80 FLEX BASIC | 更换903-334后修复 by andrew | 更换903-334后修复 by andrew | SWO-192680 | No | Repair by andrew after replacing 903-334 | Repair by andrew after replacing 903-334
ABL80 CO-OX | Paula ran level 1 external QC on May 17th at 8:35 am. The Glu result was 25. The range on the analyzer is (26-38). The analyzer stayed green and did not flat the result as being out low. They did not catch it at the time because it was green. She was looking at some results and found the issue. | Gathered data logs and opend CQC case. | D28DB816DA16F8A0 | No | Paula ran level 1 external QC on May 17th at 8:35 am. The Glu result was 25. The range on the analyzer is (26-38). The analyzer stayed green and did not flat the result as being out low. They did not catch it at the time because it was green. She was looking at some results and found the issue. | Gathered data logs and opend CQC case.
ABL90 | Verified Lac, level slot 3 Quality Control Plot. Carbon dust in dual valve housing. | Replaced all El Met tubbing, cleaned El Met measuring chanber with protein remover, cleaned dual valve housing, replaced flat membrane valve, | D28DE34CDA4D28D7 | Yes | Verified Lac, level slot 3 Quality Control Plot. Carbon dust in dual valve housing. | Replaced all El Met tubbing, cleaned El Met measuring chanber with protein remover, cleaned dual valve housing, replaced flat membrane valve, | Re classified RMED | ABL8XX entry
ABL90 FLEX | heutige Überprüfung ergibt, Sperrung Lac nicht mehr vorhanden ( nach Neustart?) QC Werte Lac nach wie vor alle außer Meßbereich Kal, Reinitialisierung LP Anderes LP eingesetzt, anschl. wieder Lac gesperrt, QC Werte nach wie vor außerhalb Meßbereich | komplette Radiometer SW neu installiert, Haken bei Abfrage komplett rausgenommen, ABL neu konfiguriert Service Dump gezogen- Weiterleitung an TS Ursache Sperrung Lac ist , wenn K, Na oder Ca deaktiviert werden, wird Lac automatisch gesperrt, beim Kundengerät jetzt Ca zugeschaltet | SWO-192911 | Yes | Today's review reveals blocking Lac no longer exists (after reboot?) QC Values Lac still all except measurement range Kal, reinitialization LP Other LP used, followed by again Lac locked, QC values still outside measuring range | complete radiometer SW reinstalled, hook completely removed when query, ABL reconfigured Service Dump Pulled- Forwarding to TS Cause blocking Lac is when K, Na or Ca are deactivated, Lac is automatically locked, with customer device now Ca switched on | Re classified RMED | SWU | FSE not following Service Tip | Disabled electrolyte | Lac
ABL90 | Inspection of the system revealed the power supply is not providing continuous power to the system. | Replaced the power supply. | D28D276BDA47879E | No | Inspection of the system revealed the power supply is not providing continuous power to the system. | Replaced the power supply. | Defect | HW | Basic Unit | Power Supply
ABL90 FLEX | Barcode Einstellungen für ITS 2 aus 5 hinzugefügt. Änderungen im System gespeichert. Startphase durchgeführt und mit dem Anwender eine Testmessung durchgeführt. | Barcode Einstellungen für ITS 2 aus 5 hinzugefügt. Änderungen im System gespeichert. Startphase durchgeführt und mit dem Anwender eine Testmessung durchgeführt. | SWO-193057 | No | Added barcode settings for ITS 2 out of 5. changes stored in the system. phase and a test measurement carried out with the user. | Added barcode settings for ITS 2 out of 5. changes stored in the system. phase and a test measurement carried out with the user. | Setup | SW | Roma.ini | Barcode
ABL90 FLEX PLUS |  | B | D28EFB70505F56BF | No |  |  | Revisit required | HW | DMS_PC | CPU
ABL90 FLEX PLUS | Bayunit, contactelemenren heb gereinigd. | Bayunit, contactelemenren heb gereinigd. | SWO-193201 | No | Bayunit, Contactelemenren have cleaned. | Bayunit, Contactelemenren have cleaned. | Cleaned | HW | Basic Unit | Sensor Bay Unit
AQT90 FLEX |  | B | SWO-192924 | Yes |  |  | Re classified RMED | Logistical upSWO |  |  | 1218
AQT90 FLEX | 自行更换备用光单903-364，重新定标后修复 by andrew | 自行更换备用光单903-364，重新定标后修复 by andrew | SWO-192309 | Yes | Replace the spare light sheet 903-364 yourself and fix by Andrew after re-calibration | Replace the spare light sheet 903-364 yourself and fix by Andrew after re-calibration | Re classified RMED | Stock recovery
ABL80 CO-OX | Kit de tubos. Revisada neumática, calibración y QCS. Todo ok. | B | D28DAE08DA0512F2 | No | Tube kit. Pneumatic review, calibration and QCS. It's all right.
ABL800 FLEX | Replace the crea reagent & calibrate the system. | Replace the crea reagent & calibrate the system. | SWO-193592 | No | Replace the crea reagent & calibrate the system. | Replace the crea reagent & calibrate the system. | Assisted Replacement | Consumables | Solution bottle | Crea
ABL815 FLEX | 更换902-833一个正常。 by andrew | 更换902-833一个正常。 by andrew | SWO-192630 | No | Replace 902-833 one normal. by andrew | Replace 902-833 one normal. by andrew | Defect | HW | Oxi module | Lamp unit | Neon
ABL810 FLEX | Pumpenkalibration bringt auf dauer keine Verbesserung. Pumpenschläuche der Elektroden Module aus dem Kundenbestand ausgetauscht. | Pumpenkalibration bringt auf dauer keine Verbesserung. Pumpenschläuche der Elektroden Module aus dem Kundenbestand ausgetauscht. | SWO-193161 | No | Pump calibration does not bring any improvement in the long run. Pump hoses of the electrode modules are replaced from the customer's inventory. | Pump calibration does not bring any improvement in the long run. Pump hoses of the electrode modules are replaced from the customer's inventory. | Defect | HW | pH/BG Module | Pump tubes
ABL825 FLEX | Lüfter im Modul ersetzt, Gerät kalibriert und QC durchgeführt, | Lüfter im Modul ersetzt, Gerät kalibriert und QC durchgeführt, | SWO-193066 | No | replaced in the module, device calibrated and QC performed, | replaced in the module, device calibrated and QC performed, | Defect | HW | Autocheck module | Fan | 763
ABL825 FLEX | Customer had done several rinses and the problem resolved itself before FSE onsite. | Performed Zero Current calibration- pO2 changed by 10pA to 18pA | D28DC367DA234A79 | No | Customer had done several rinses and the problem resolved itself before FSE onsite. | Performed Zero Current calibration- pO2 changed by 10pA to 18pA | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL810 FLEX | Reparatur im Rahmen der Halbjahreswartung durchgeführt. | Reparatur im Rahmen der Halbjahreswartung durchgeführt. | SWO-193327 | No | repair carried out as part of half-year maintenance. | repair carried out as part of half-year maintenance. | Adjustment | HW | pH/BG Module | Zero current
ABL835 FLEX | Found lamp intensity on the edge of out. Sometimes in, sometime out. | Adjusted lamp. Performed all necessary calibrations. | D28DAC92DA068FE8 | No | Found lamp intensity on the edge of out. Sometimes in, sometime out. | Adjusted lamp. Performed all necessary calibrations. | Adjustment | HW | Oxi module | Lamp unit
ABL80 FLEX BASIC | 更换903-334后修复 by andrew | 更换903-334后修复 by andrew | SWO-192595 | No | Repair by andrew after replacing 903-334 | Repair by andrew after replacing 903-334
ABL80 FLEX BASIC | 更换902-972后修复 by andrew | 更换902-972后修复 by andrew | SWO-192940 | No | Repair by andrew after replacement 902-972 | Repair by andrew after replacement 902-972
ABL90 FLEX | komplette Radiometer SW neu installiert, Haken bei Abfrage komplett rausgenommen, ABL neu konfiguriert Service Dump gezogen- Weiterleitung an TS Ursache Sperrung Lac ist , wenn K, Na oder Ca deaktiviert werden, wird Lac automatisch gesperrt, beim Kundengerät jetzt Ca zugeschaltet | komplette Radiometer SW neu installiert, Haken bei Abfrage komplett rausgenommen, ABL neu konfiguriert Service Dump gezogen- Weiterleitung an TS Ursache Sperrung Lac ist , wenn K, Na oder Ca deaktiviert werden, wird Lac automatisch gesperrt, beim Kundengerät jetzt Ca zugeschaltet | SWO-193392 | Yes | complete radiometer SW reinstalled, hook completely removed when query, ABL reconfigured Service Dump Pulled- Forwarding to TS Cause blocking Lac is when K, Na or Ca are deactivated, Lac is automatically locked, with customer device now Ca switched on | complete radiometer SW reinstalled, hook completely removed when query, ABL reconfigured Service Dump Pulled- Forwarding to TS Cause blocking Lac is when K, Na or Ca are deactivated, Lac is automatically locked, with customer device now Ca switched on | Re classified RMED | SWU | FSE not following Service Tip | Disabled electrolyte | Lac
ABL90 FLEX | muß Änderung in der roma.ini vorgenommen werden ABL erkennt Tastatur nicht, andere Tastatur nicht vorhanden Kunde informiert Problem muß noch behoben werden | muß Änderung in der roma.ini vorgenommen werden ABL erkennt Tastatur nicht, andere Tastatur nicht vorhanden Kunde informiert Problem muß noch behoben werden | SWO-193563 | No | must be changed in the roma.ini ABL does not recognize keyboard, other keyboard does not exist Customer informed Problem still needs to be fixed | must be changed in the roma.ini ABL does not recognize keyboard, other keyboard does not exist Customer informed Problem still needs to be fixed | Revisit required | SW | Roma.ini | Roma.ini
ABL90 FLEX | Gasket hold clean. Gasket module clean. tHb calibration work. Sensor cassette clean. Calibration check QC check. Sample test. Activity log check. | Gasket hold clean. Gasket module clean. tHb calibration work. Sensor cassette clean. Calibration check QC check. Sample test. Activity log check. | SWO-195224 | No | Gasket hold clean. Gasket module clean. tHb calibration work. Sensor cassette clean. Calibration check QC check. Sample test. Activity log check. | Gasket hold clean. Gasket module clean. tHb calibration work. Sensor cassette clean. Calibration check QC check. Sample test. Activity log check. | Adjustment | HW | Oxi module | tHb Cal
ABL90 FLEX PLUS |  | B | D28EFC97506156C0 | No |  |  | Defect | HW | DMS_PC | CPU
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-193517 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Unknown
ABL827 FLEX | Met2ﾓｼﾞｭｰﾙｴｱﾊﾞﾙﾌﾞがｴﾗｰ出していた。また、PHﾓｼﾞｭｰﾙﾘﾝｽ側ﾊﾞﾙﾌﾞの動きもいまいちのため交換。 | B | SWO-194751 | No | The Met2 module air valve was out of error. In addition, the movement of the PH module rinse side valve is also bad, so we replaced it. |  | Defect | HW | Met2 Module | Valve
ABL800 BASIC | 回路洗浄実施（次亜塩素酸溶液） 全メンブラン交換 Na電極交換 Cal2 QC OK | B | SWO-193265 | No | Circuit cleaning (hypochlorite solution) All Membrane Replacement Na electrode replacement Cal2 QC OK |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | pCO2 Elektrode (E788 JT04) geprüft und ersetzt, Funktionskontrolle ok | FSE-Einsatz erforderlich | SWO-193533 | No | pCO2 electrode (E788 JT04) tested and replaced, function control ok | FSE deployment required | Assisted Replacement | Electrode | Electrode | pCO2 | Revisit?
ABL825 FLEX | pCO2 Elektrode geprüft und ausgetauscht | FSE-Einsatz erforderlich | SWO-193534 | Yes | pCO2 electrode tested and replaced | FSE deployment required | Re classified RMED | Double entry
AQT90 FLEX |  | B | D28F3470501F9D19 | No |  |  | Defect | Hw | Aspiration Module | Needle Wash Unit | 3412
AQT90 FLEX | 重新定标后修复 by andrew | 重新定标后修复 by andrew | SWO-193339 | No | Fix by andrew after re-calibration | Fix by andrew after re-calibration | Assisted Replacement | Consumables | Test Kit | Cartridge | TnI
ABL80 FLEX BASIC |  | B | SWO-193666 | No
AQT90 FLEX | 更换库存机903-364后修复 by andrew | 更换库存机903-364后修复 by andrew | SWO-193678 | No | Repair by andrew after replacing stock machine 903-364 | Repair by andrew after replacing stock machine 903-364 | Defect | HW | Process Module | Optical Unit | 135
AQT90 FLEX | 更换903-364，903-236各一个正常。 by jesson | 更换903-364，903-236各一个正常。 by jesson | SWO-193602 | No | Replace 903-364, 903-236 each normal. by jesson | Replace 903-364, 903-236 each normal. by jesson | Defect | HW | Process Module | Optical Unit | 124
ABL800 BASIC | Performed Lamp unit adjustment, 2 point calibration performed successfully. | Performed Lamp unit adjustment, 2 point calibration performed successfully. | SWO-193698 | No | Performed Lamp unit adjustment, 2 point calibration performed successfully. | Performed Lamp unit adjustment, 2 point calibration performed successfully. | Adjustment | HW | Oxi module | Lamp unit
ABL80 FLEX BASIC |  | B | SWO-193864 | No
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-194177 | Yes | The replacement test card is OK. | The replacement test card is OK.
ABL90 |  | B | D292DF1AB222C2B4 | Yes
ABL90 FLEX | Änderung in der roma.ini vorgenommen Neustart und Test (i.O.) | Änderung in der roma.ini vorgenommen Neustart und Test (i.O.) | SWO-193993 | No | Change made in the roma.ini Restart and Test (i.O.) | Change made in the roma.ini Restart and Test (i.O.) | Setup | SW | Roma.ini | Barcode reader
ABL90 FLEX | None required Analyser working within specification | None required Analyser working within specification | SWO-194081 | No | None required Analyser working within specification | None required Analyser working within specification | No fault found | Performance | Oxi module | ctBil
ABL90 FLEX PLUS |  | B | D29324A6B22E80E1 | Yes
ABL90 FLEX PLUS |  | B | D2931E66B217DAFE | No |  |  | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX |  | B | SWO-193880 | No |  |  | Defect | HW | Oxi module | Hemolyzer | 581
ABL800 FLEX | CPU Unit replaced. Analyzer serviced and calibrated properly. Handed over in fine working condition. Analyzer is on observation till next day. | Analyser checked and found working fine. Analyzer calibrated and found all parameters well with in range. Solutions replaced. | SWO-195452 | No | CPU Unit replaced. Analyzer serviced and calibrated properly. Handed over in fine working condition. Analyzer is on observation till next day. | Analyser checked and found working fine. Analyzer calibrated and found all parameters well with in range. Solutions replaced. | Defect | HW | DMS_PC | CPU
ABL800 BASIC | モジュール引き上げ | B | SWO-193796 | No | Module pull up |  | Defect | HW | Oxi module | Oxi module | Neon
ABL800 BASIC | Reassembled the gears, 2 point calibration performed successfully. | Reassembled the gears, 2 point calibration performed successfully. | SWO-194848 | No | Reassembled the gears, 2 point calibration performed successfully. | Reassembled the gears, 2 point calibration performed successfully. | Defect | HW | Printer Module | Printer head
ABL825 FLEX |  | B | D293363DB24480E7 | Yes
ABL835 FLEX | OXI PCB udskiftet, service tHb og alm. tHb kalibreret, slutafprøvet. | OXI PCB udskiftet, service tHb og alm. tHb kalibreret, slutafprøvet. | SWO-194087 | No | OXI PCB udskiftet, service tHb og alm. tHb kalibreret, slutafprøvet. | OXI PCB udskiftet, service tHb og alm. tHb kalibreret, slutafprøvet. | Defect | HW | Oxi module | PCB, Oxi module
ABL90 | C USTOMER REQUESTS R ADIANCE CHANGED TO AQURE, T IME CHANGED TO 24H F ORMAT AND I.T C ONNECTION ESTABLISHED | R ADIANCE CHANGED TO AQURE; T IME CHANGED TO 24H F ORMAT AND I.T C ONNECTION E STABLISHED . C USTOMER REQUESTS SOFTWARE UPGRADE BE PERFORMED AT A LATER DATE . | D29423BB28396F12 | Yes | C USTOMER REQUESTS R ADIANCE CHANGED TO AQURE, T IME CHANGED TO 24H F ORMAT AND I.T C ONNECTION ESTABLISHED | R ADIANCE CHANGED TO AQURE; T IME CHANGED TO 24H F ORMAT AND I.T C ONNECTION E STABLISHED . C USTOMER REQUESTS SOFTWARE UPGRADE BE PERFORMED AT A LATER DATE . | Re classified RMED | Planned by Customer request
ABL90 FLEX PLUS |  | B | D2953ED39E36F284 | Yes |  |  | Re classified RMED | Installation
ABL90 FLEX PLUS | Upgraded and rebooted software. | B | SWO-194579 | No | Upgraded and rebooted software. |  | Defect | SW | RMED SW | Data base | 145
ABL90 FLEX PLUS | Replaced bent inlet probe. | user has replaced IG SP and retested, finished stat up, states replace SP, advised user to check analyser call back in 10. got user into service found high clot value for SC flow. advised to replace SC and re test, Customer spoke to on-call engineer, Still faulty after replacing SC, customer requested visit from engineer. | SWO-194581 | No | Replaced bent inlet probe. | user has replaced IG SP and retested, finished stat up, states replace SP, advised user to check analyser call back in 10. got user into service found high clot value for SC flow. advised to replace SC and re test, Customer spoke to on-call engineer, Still faulty after replacing SC, customer requested visit from engineer. | Replaced | HW | Auto Inlet module | Probe bend
AQT90 FLEX | 更换905-908一个正常。by jesson | 更换905-908一个正常。by jesson | SWO-194292 | No | Replace 905-908 one normal. by jesson | Replace 905-908 one normal. by jesson | Defect | HW | Process Module | Sensor, BNP
AQT90 FLEX |  | B | SWO-194367 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 测试卡复位后修复 | B | SWO-194807 | No | Repair after test card reset |  | Defect | Consumables | Test Kit | Cartridge | Broken wings
ABL800 FLEX | Analyser checked and found working fine. Analyzer calibrated and found all parameters well with in range. Solutions replaced. | Analyser checked and found working fine. Analyzer calibrated and found all parameters well with in range. Solutions replaced. | SWO-195453 | Yes | Analyser checked and found working fine. Analyzer calibrated and found all parameters well with in range. Solutions replaced. | Analyser checked and found working fine. Analyzer calibrated and found all parameters well with in range. Solutions replaced. | Re classified RMED | Follow Up
ABL827 FLEX |  | B | D29523CF9E3CB519 | No |  |  | Revisit required | HW | DMS_PC | DMS_PC
ABL835 FLEX | Lösungen abgezogen und eingespannt Kein Fehler feststellbar, bei QC Überprüfung O2 instabil pO2 neu membraniert Gerät kalibriert und überprüft | Lösungen abgezogen und eingespannt Kein Fehler feststellbar, bei QC Überprüfung O2 instabil pO2 neu membraniert Gerät kalibriert und überprüft | SWO-194424 | No | Solutions removed and clamped No error detectable, in QC review O2 unstable pO2 re-membraneed Device calibrated and checked | Solutions removed and clamped No error detectable, in QC review O2 unstable pO2 re-membraneed Device calibrated and checked | Cleaned | HW | Basic Unit | Barcode reader, Basic Unit
ABL837 FLEX | Checked logs. Checked all valves. Checked pump rotors. Checked waste lines. | Flushed waste lines. Replaced Inlet gasket and Probe. Flushed waste connector as it was clogged. 6 Month PM done. Replaced Cleaning solution and Crea membranes. | D29407D9286284B5 | No | Checked logs. Checked all valves. Checked pump rotors. Checked waste lines. | Flushed waste lines. Replaced Inlet gasket and Probe. Flushed waste connector as it was clogged. 6 Month PM done. Replaced Cleaning solution and Crea membranes. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX | 0251, halogen lamp intensity too high and 0579, not possible to measure errors are present. Halogen lamp intensity is observed to be higher than the 3100-3900 range at >4000. | Adjusted the halogen lamp intensity and performed a complete oximetry test. Performed all necessary calibrations. Analyzer is within acceptable operating specifications. | D2940285285F5925 | No | 0251, halogen lamp intensity too high and 0579, not possible to measure errors are present. Halogen lamp intensity is observed to be higher than the 3100-3900 range at >4000. | Adjusted the halogen lamp intensity and performed a complete oximetry test. Performed all necessary calibrations. Analyzer is within acceptable operating specifications. | Adjustment | HW | Oxi module | Lamp unit | Halogen
TCM/Tosca | The touch screen is out of calibration. Also, there is a bubble under the screen and unable to perform calibration. | Replaced the touch screen. | D293DA48281F0407 | No | The touch screen is out of calibration. Also, there is a bubble under the screen and unable to perform calibration. | Replaced the touch screen.
ABL80 FLEX BASIC |  | B | SWO-194607 | No
ABL80 FLEX BASIC |  | B | SWO-194769 | No
ABL80 FLEX CO-OX | Gerät überprüft, Vergleichsmessungen durchgeführt, tHb Kal und Zero Abgleich durchgeführt, Konstanzmessungen o.k. | Gerät überprüft, Vergleichsmessungen durchgeführt, tHb Kal und Zero Abgleich durchgeführt, Konstanzmessungen o.k. | SWO-195233 | No | device checked, comparative measurements carried out, tHb Kal and Zero matching performed, Constance measurements ok | device checked, comparative measurements carried out, tHb Kal and Zero matching performed, Constance measurements ok
ABL80 FLEX BASIC |  | B | SWO-194749 | No
ABL90 FLEX | Lösungspack vom Analysator auswerfen lassen, selben Lösungspack nochmals installiert, Überprüfung des Analysators im Service (Adjust Manifold, Pressuretest, LS-Justage, Pumpenkal., SP Check, Clot detect remove), Analysator gespült, System und Funktion i.O. --> aktuelle SC (435-112) bei vorheriger SC(425-189 --> installiert am 19.05.17) ebenfalls QC-Fehler bei pCO2, welcher sich erst am nächsten Tag (20.05.17) bemerkbar machte --> Kal. und QC-Fehler bei pCO2 --> keine SC mehr vor Ort --> FSE (TobEh) organisiert sich eine SC und wird diese am Freitag einbauen | SC (R0539) eingebaut und (R435) ausgebaut, System und Funktion i.O. | SWO-194827 | No | Let the analyzer eject the solution pack, reinstall the same solution pack, check the analyzer in the service (Adjust Manifold, Pressuretest, LS-Adjustment, Pump Calca, SP Check, Clot detect detect remove), analyzer flushed, system and function i.O. --> current SC (435-112) with previous SC(425-189 --> installed on 19.05.17) also QC error with pCO2, which only became noticeable the next day (20.05.17) --> Kal. and QC error at pCO2 --> no more SC on site --> FSE (TobEh) organizes a SC and will install it on Friday | SC (R0539) built-in and (R435) removed, system and function in the | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX PLUS |  | B | D2953F189E05F285 | Yes
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-194606 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Thermistor
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-194630 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Reference | 1361
ABL90 FLEX | 更换两个试剂包正常。 | 更换两个试剂包正常。 | SWO-194987 | No | Replacing both reagent kits is normal. | Replacing both reagent kits is normal. | Assisted Replacement | SP | SP | FT | 1531
ABL90 FLEX | 更换910-342一个正常。by jesson | 更换910-342一个正常。by jesson | SWO-194291 | No | Replace 910-342 one normal. by jesson | Replace 910-342 one normal. by jesson | Defect | HW | DMS_PC | CPU
AQT90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-195084 | No | Repair after replacing test card | Repair after replacing test card | Defect | Consumables | Test Kit | Cartridge | 5784
ABL825 FLEX |  | B | D29594BD9E287ED9 | Yes
ABL825 FLEX |  | B | D2955D639E297825 | Yes
ABL827 FLEX |  | B | D2952B959E63B51D | No |  |  | Defect | HW | DMS_PC | CPU
ABL815 FLEX | pCO2 neu Membraniert, Messkammer gereinigt, LS Justage und Pumpenkal durchgeführt. Mehrfach kalibriert. | pCO2 neu Membraniert, Messkammer gereinigt, LS Justage und Pumpenkal durchgeführt. Mehrfach kalibriert. | SWO-195287 | No | pCO2 newly membraneed, measuring chamber cleaned, LS adjustment and pump scale performed. Multiple calibrated. | pCO2 newly membraneed, measuring chamber cleaned, LS adjustment and pump scale performed. Multiple calibrated. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Friction Detection Unit bestellt | Friction Detection Unit bestellt | SWO-195302 | No | Friction Detection Unit ordered | Friction Detection Unit ordered | Revisit required | HW | Inlet module | Friction detector unit
ABL825 FLEX | 更换902-833后修复 by jesson | 更换902-833后修复 by jesson | SWO-194256 | No | Repair by jesson after replacement 902-833 | Repair by jesson after replacement 902-833 | Defect | HW | Oxi module | Oxi module | Hemolyzer
ABL825 FLEX | QC soution is observed leaking in and around the analyzer. The pH sealing ring is also observed to be missing from the pH electrode. | Replaced the pH sealing ring, verified that the leak is no longer present and verified successful QC's and calibrations. Also upgraded system software to current version, v6.17. Analyzer is within acceptable operating specifications. | D29553CD9E219A1A | No | QC soution is observed leaking in and around the analyzer. The pH sealing ring is also observed to be missing from the pH electrode. | Replaced the pH sealing ring, verified that the leak is no longer present and verified successful QC's and calibrations. Also upgraded system software to current version, v6.17. Analyzer is within acceptable operating specifications. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
AQT90 FLEX | De barcodescanner van de Inletmodule heb gereinigd. | De barcodescanner van de Inletmodule heb gereinigd. | SWO-194961 | No | The barcode scanner of the Inletmodule has been cleaned. | The barcode scanner of the Inletmodule has been cleaned. | Cleaned | HW | Aspiration Module | Sensor, Vacutainer
AQT90 FLEX | Eine Kunststoff-Nippel an der Testkassette isst abgebrochen, somit konnte der Analysator die Testkassette nicht mehr ins Storage Wheel zurückziehen. Testkassette herausgenommen und Bereich des Storage Wheels überprüft, ob Nippel vorhanden und nicht Rotation blockiert. Kalibration gestartet. Kalibration on Ordnung. | Eine Kunststoff-Nippel an der Testkassette isst abgebrochen, somit konnte der Analysator die Testkassette nicht mehr ins Storage Wheel zurückziehen. Testkassette herausgenommen und Bereich des Storage Wheels überprüft, ob Nippel vorhanden und nicht Rotation blockiert. Kalibration gestartet. Kalibration on Ordnung. | SWO-194962 | No | A plastic nipple on the test cartridge eats aborted, so the analyzer could no longer retract the test cartridge into the storage wheel. Test cartridge removed and area of the storage wheel checks whether nipples exist and does not block rotation. calibration started. Calibration on order. | A plastic nipple on the test cartridge eats aborted, so the analyzer could no longer retract the test cartridge into the storage wheel. Test cartridge removed and area of the storage wheel checks whether nipples exist and does not block rotation. calibration started. Calibration on order. | Defect | Consumables | Test Kit | Cartridge | Broken tooth
ABL90 FLEX | Sensor Kassette in Bottrop geholt und zum Kunden geliefert. Kunde hatte noch eine SC dieser er in der Zwischenzeit eingesetzt hatte. Gerät kalibriert und QC durchgeführt. Problematik ist entstanden, weil der Kunde den ABL90 in Kombination mit nicht heparinisierten Spritzen verwendet. Mit dem Kunden die Problematik besprochen. Kunde wird in Zukunft Pico50 Pn verwenden. | Sensor Kassette in Bottrop geholt und zum Kunden geliefert. Kunde hatte noch eine SC dieser er in der Zwischenzeit eingesetzt hatte. Gerät kalibriert und QC durchgeführt. Problematik ist entstanden, weil der Kunde den ABL90 in Kombination mit nicht heparinisierten Spritzen verwendet. Mit dem Kunden die Problematik besprochen. Kunde wird in Zukunft Pico50 Pn verwenden. | SWO-195378 | Yes | Sensor cassette in Bottrop and delivered to the customer. Customer still had a SC which he had used in the meantime. calibrated and QC performed. Problems arose because the customer uses the ABL90 in combination with non-heparinized syringes. Discussed the issue with the customer. Customer will use Pico50 Pn in the future. | Sensor cassette in Bottrop and delivered to the customer. Customer still had a SC which he had used in the meantime. calibrated and QC performed. Problems arose because the customer uses the ABL90 in combination with non-heparinized syringes. Discussed the issue with the customer. Customer will use Pico50 Pn in the future. | Re classified RMED | Logistical upSWO
ABL90 FLEX | Anhang zum Notdienst. Lieferung von 2 Sensor Kasetten und 1 LP. | Sensor Kassette in Bottrop geholt und zum Kunden geliefert. Kunde hatte noch eine SC dieser er in der Zwischenzeit eingesetzt hatte. Gerät kalibriert und QC durchgeführt. Problematik ist entstanden, weil der Kunde den ABL90 in Kombination mit nicht heparinisierten Spritzen verwendet. Mit dem Kunden die Problematik besprochen. Kunde wird in Zukunft Pico50 Pn verwenden. | SWO-195379 | Yes | Appendix to the emergency service. Delivery of 2 sensor sets and 1 LP. | Sensor cassette in Bottrop and delivered to the customer. Customer still had a SC which he had used in the meantime. calibrated and QC performed. Problems arose because the customer uses the ABL90 in combination with non-heparinized syringes. Discussed the issue with the customer. Customer will use Pico50 Pn in the future. | Re classified RMED | See SWO-195378
AQT90 FLEX | 更换903-364一个正常. by jesson | 更换903-364一个正常. by jesson | SWO-195012 | No | Replace 903-364 one normal. by jesson | Replace 903-364 one normal. by jesson | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364,903-584各一个正常。by jesson | 更换903-364,903-584各一个正常。by jesson | SWO-195029 | No | Replace 903-364, 903-584 each normal. by jesson | Replace 903-364, 903-584 each normal. by jesson | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-195066 | No |  |  | Defect | HW | Incubator module | Cable
ABL800 BASIC | Lacメンブラン交換。QC測定。 | B | SWO-195169 | No | Lac Membrane replacement. QC measurement. |  | Assisted Replacement | Electrode | Membrane | Lac | Sensitivity
ABL800 BASIC | Clean all the chambers ,clean PO2 electrode and done 2-point calibration which is OK but near the high limit.Run QC which is also OK. | Again clean all chambers, again scan both Gases, run zero current calibration and done 2-point calibration but problem remains same.Replace PO2 electrode but not effected much more.Again check and found PCo2 membrane is leaking. Replace po2 & pco2 membrane and done calibration but sensitivity is more than 40.Replace electrode with new membrane and done calibration and problem has been resolved. | SWO-195588 | No | Clean all the chambers ,clean PO2 electrode and done 2-point calibration which is OK but near the high limit. Run QC which is also OK. | Again clean all chambers, again scan both Gases, run zero current calibration and done 2-point calibration but problem remains same. Replace PO2 electrode but not effected much more. Again check and found PCo2 membrane is leaking. Replace po2 & pco2 membrane and done calibration but sensitivity is more than 40.Replace electrode with new membrane and done calibration and problem has been resolved. | Assisted Replacement | Electrode | Membrane | pCO2
ABL825 FLEX | Error happening on patient samples. Minor calibration drift recently. | Rechecked valves. Replaced pH seal ring again. | D296AABA14309E2F | No | Error happening on patient samples. Minor calibration drift recently. | Rechecked valves. Replaced pH seal ring again. | Assisted Replacement | Electrode | Electrode | pH | Sealing ring
ABL825 FLEX |  | B | D296E8A114637A09 | No |  |  | Defect | HW | Oxi module | Oxi module | Neon
ABL825 FLEX |  | B | D296EB2014587A0E | No |  |  | Defect | HW | Oxi module | Spectrometer | LEH
ABL80 FLEX BASIC | 更换902-972一个正常。by jesson | 更换902-972一个正常。by jesson | SWO-195049 | No | Replace 902-972 one normal. by jesson | Replace 902-972 one normal. by jesson
ABL90 FLEX | SC (R0539) eingebaut und (R435) ausgebaut, System und Funktion i.O. | SC (R0539) eingebaut und (R435) ausgebaut, System und Funktion i.O. | SWO-195310 | No | SC (R0539) built-in and (R435) removed, system and function in the | SC (R0539) built-in and (R435) removed, system and function in the | Assisted Replacement | SC | SC | pCO2
ABL90 FLEX PLUS | Performed Manual flush Replaced gasket and ran fluid transport checks. Analyser is now passing. Initialised, calibrated and has passed all QC's. | Performed Manual flush Replaced gasket and ran fluid transport checks. Analyser is now passing. Initialised, calibrated and has passed all QC's. | SWO-195466 | No | Performed Manual flush Replaced gasket and ran fluid transport checks. Analyser is now passing. Initialised, calibrated and has passed all QC's. | Performed Manual flush Replaced gasket and ran fluid transport checks. Analyser is now passing. Initialised, calibrated and has passed all QC's. | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Einlass entfernt und Stellrad bewegt - lässt sich ohne Probleme bewegen und man spürt keine Blockade oder unrunde Drehung, ABL beginnt ohne Einlassmodul die Startphase von 10 Min., nach 10 Min. ABL mittels Reset-Knopf ausgeschaltet, im ausgeschalteten Zustand Einlassmodul wieder angesetzt und ABL neu gestartet, Startphase von 50 Min. hat begonnen, nach Startphase mehrere Testmessungen in Spritzen- und Kapillarmodus durchgeführt - Einlass öffnet und schließt ganz normal und ohne krachenden Geräusche, System- und Funktionskontrolle in Ordnung | Einlass entfernt und Stellrad bewegt - lässt sich ohne Probleme bewegen und man spürt keine Blockade oder unrunde Drehung, ABL beginnt ohne Einlassmodul die Startphase von 10 Min., nach 10 Min. ABL mittels Reset-Knopf ausgeschaltet, im ausgeschalteten Zustand Einlassmodul wieder angesetzt und ABL neu gestartet, Startphase von 50 Min. hat begonnen, nach Startphase mehrere Testmessungen in Spritzen- und Kapillarmodus durchgeführt - Einlass öffnet und schließt ganz normal und ohne krachenden Geräusche, System- und Funktionskontrolle in Ordnung | SWO-195534 | No | Inlet removed and adjusting wheel moved - can be moved without problems and you can not feel any blockage or uncircular rotation, ABL starts the start phase of 10 min. without inlet module, after 10 min. ABL is switched off by means of a reset button, when the inlet module is switched off again started and ABL restarted, start phase of 50 min. has begun, after start phase several test measurements in syringe and capillary mode - inlet opens and closes normally and without crashing noise, system and function control in order | Inlet removed and adjusting wheel moved - can be moved without problems and you can not feel any blockage or uncircular rotation, ABL starts the start phase of 10 min. without inlet module, after 10 min. ABL is switched off by means of a reset button, when the inlet module is switched off again started and ABL restarted, start phase of 50 min. has begun, after start phase several test measurements in syringe and capillary mode - inlet opens and closes normally and without crashing noise, system and function control in order | Adjustment | SW | Auto Inlet module | Off/On | Inlet Open
ABL90 FLEX | 更换910-324一个正常。by jesson | 更换910-324一个正常。by jesson | SWO-194997 | No | Replace 910-324 one normal. by jesson | Replace 910-324 one normal. by jesson | Defect | HW | DMS_PC | Touch screen
AQT90 FLEX | 更换903-364一个，重新定标正常。by jesson | 更换903-364一个，重新定标正常。by jesson | SWO-195041 | No | Replace the 903-364 one and re-label normally. by jesson | Replace the 903-364 one and re-label normally. by jesson | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 清洁处理盘，并系统保养 | 清洁处理盘，并系统保养 | SWO-195501 | No | Clean the processing disc and maintain the system | Clean the processing disc and maintain the system | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | Lak. Membran gewechselt (alte Membran noch auf der Elektrodenspitze) , Probelauf ok | Lak. Membran gewechselt (alte Membran noch auf der Elektrodenspitze) , Probelauf ok | SWO-195438 | No | Lak. Membrane changed (old membrane still on the electrode tip) , test run ok | Lak. Membrane changed (old membrane still on the electrode tip) , test run ok | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | pH Messkammer gereinigt, Gerät kalibriert und QC durchgeführt. | pH Messkammer gereinigt, Gerät kalibriert und QC durchgeführt. | SWO-195388 | No | pH measuring chamber cleaned, device calibrated and QC performed. | pH measuring chamber cleaned, device calibrated and QC performed. | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL80 FLEX BASIC | 更换636-520后修复 by jesson | 更换636-520后修复 by jesson | SWO-195199 | Yes | Repair by jesson after replacement 636-520 | Repair by jesson after replacement 636-520
ABL90 FLEX PLUS | 回路洗浄、ガスケット・溶液パック排液ドレイン掃除 | B | SWO-195625 | No | Circuit cleaning, gasket and solution pack drain drain cleaning |  | Cleaned | HW | Auto Inlet module | Gasket - Clot
AQT90 FLEX | 更换光单，重新定标 | 更换光单，重新定标 | SWO-195500 | No | Replace the light sheet and re-label | Replace the light sheet and re-label | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | 更换903-364一个正常. by jesson | SWO-195541 | No |  | Replace 903-364 one normal. by jesson | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-198后修复(by tim) | 更换903-198后修复(by tim) | SWO-195074 | No | Repair after replacement 903-198 (by tim) | Repair after replacement 903-198 (by tim) | Defect | HW | Analyzer Control Module | PCB, Main Board | 1
AQT90 FLEX | 更换903-364后修复 by jesson | 更换903-364后修复 by jesson | SWO-195262 | No | Repair by jesson after replacing 903-364 | Repair by jesson after replacing 903-364 | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换备件903-329及924-504后修复(by tim) | 更换备件903-329及924-504后修复(by tim) | SWO-195079 | No | Replacement of spare parts 903-329 and 924-504 after repair (by tim) | Replacement of spare parts 903-329 and 924-504 after repair (by tim) | Defect | HW | Aspiration Module | Needle Wash Unit | 3412
ABL825 FLEX | 更换910-275后修复 by jesson | 更换910-275后修复 by jesson | SWO-195263 | Yes | Repair by jesson after replacement 910-275 | Repair by jesson after replacement 910-275
ABL800 BASIC | Again clean all chambers, again scan both Gases, run zero current calibration and done 2-point calibration but problem remains same.Replace PO2 electrode but not effected much more.Again check and found PCo2 membrane is leaking. Replace po2 & pco2 membrane and done calibration but sensitivity is more than 40.Replace electrode with new membrane and done calibration and problem has been resolved. | Again clean all chambers, again scan both Gases, run zero current calibration and done 2-point calibration but problem remains same.Replace PO2 electrode but not effected much more.Again check and found PCo2 membrane is leaking. Replace po2 & pco2 membrane and done calibration but sensitivity is more than 40.Replace electrode with new membrane and done calibration and problem has been resolved. | SWO-195592 | No | Again clean all chambers, again scan both Gases, run zero current calibration and done 2-point calibration but problem remains same. Replace PO2 electrode but not effected much more. Again check and found PCo2 membrane is leaking. Replace po2 & pco2 membrane and done calibration but sensitivity is more than 40.Replace electrode with new membrane and done calibration and problem has been resolved. | Again clean all chambers, again scan both Gases, run zero current calibration and done 2-point calibration but problem remains same. Replace PO2 electrode but not effected much more. Again check and found PCo2 membrane is leaking. Replace po2 & pco2 membrane and done calibration but sensitivity is more than 40.Replace electrode with new membrane and done calibration and problem has been resolved. | Assisted Replacement | Electrode | Membrane | pO2 | Sensitivity
ABL810 FLEX | Customer states Inlet clip broken. Probe not secure. | Replaced inlet clip assembly. | D298F0BF003001FE | No | Customer states Inlet clip broken. Probe not secure. | Replaced inlet clip assembly. | Defect | HW | Inlet module | Inlet clip
ABL825 FLEX | Intermittently occurring on patient samples and QC. | Replaced pump tubes and interconnecting tubings. Rechecked valves. | D298F070003C1B01 | No | Intermittently occurring on patient samples and QC. | Replaced pump tubes and interconnecting tubings. Rechecked valves. | Defect | HW | pH/BG Module | Pump tubes | 653
ABL80 FLEX BASIC | 备件没有使用，重装软件后修复 by jesson | 备件没有使用，重装软件后修复 by jesson | SWO-195503 | No | Spare parts not used, reload software to repair by jesson | Spare parts not used, reload software to repair by jesson
ABL80 FLEX Basic version |  | B | D29AD86676072D41 | No
ABL90 FLEX | 更换903-199后修复 by jesson | 更换903-199后修复 by jesson | SWO-195277 | No | Repair by jesson after replacement 903-199 | Repair by jesson after replacement 903-199 | Defect | HW | Basic Unit | Cassette retainer
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-201252 | No | Repair after replacing test card | Repair after replacing test card
ABL80 FLEX BASIC | SC数回交換。 | B | SWO-195920 | No | SC several times replaced.
ABL90 FLEX PLUS | Remplacé le rotor péristaltique e.o (blocage) e.o Calibration e.o contrôle de qualité e.o | Remplacé le rotor péristaltique e.o (blocage) e.o Calibration e.o contrôle de qualité e.o | SWO-195882 | No | Replaced the peristaltic rotor e.o (blocking) e.o Calibration e.o quality control e.o | Replaced the peristaltic rotor e.o (blocking) e.o Calibration e.o quality control e.o | Defect | HW | Basic Unit | Pump
ABL90 FLEX PLUS | Kundin hatte LP bereits getauscht. ABL war in 55 minütiger Startphase. Kalibrationen und QC's vorgenommen. | Kundin hatte LP bereits getauscht. ABL war in 55 minütiger Startphase. Kalibrationen und QC's vorgenommen. | SWO-195720 | No | Customer had already exchanged LP. ABL was in 55 minutes of launch. calibrations and QC's. | Customer had already exchanged LP. ABL was in 55 minutes of launch. calibrations and QC's. | Assisted Replacement | SP | SP | FT
ABL800 BASIC | 比較電極交換 Cal2 QC OK | 回路洗浄実施（次亜塩素酸溶液） 全メンブラン交換 Na電極交換 Cal2 QC OK | SWO-195660 | No | Comparative electrode replacement Cal2 QC OK | Circuit cleaning (hypochlorite solution) All Membrane Replacement Na electrode replacement Cal2 QC OK | Assisted Replacement | Electrode | Electrode | Na
ABL835 FLEX |  | B | SWO-196060 | No |  |  | Revisit required | HW | Autocheck module | Autocheck module
AQT90 FLEX |  | B | D29BB5C9EC395029 | Yes
AQT90 FLEX |  | B | D29B9943EC05D2CE | No |  |  | Replaced | Analyzer | Analyzer replaced | Analyzer replaced
ABL80 FLEX BASIC | Preseason study done on QC and two different patient samples and found all acceptable. Analyser Serviced and calibrated properly. Handed over in a fine working condition. Staff Satisfied with results. | Preseason study done on QC and two different patient samples and found all acceptable. Analyser Serviced and calibrated properly. Handed over in a fine working condition. Staff Satisfied with results. | SWO-196587 | No | Preseason study done on QC and two different patient samples and found all acceptable. Analyser Serviced and calibrated properly. Handed over in a fine working condition. Staff Satisfied with results. | Preseason study done on QC and two different patient samples and found all acceptable. Analyser Serviced and calibrated properly. Handed over in a fine working condition. Staff Satisfied with results.
ABL80 FLEX BASIC | Found calibration values Lac sensitivity was low, tri to calibrate but problem was remining same, need to replace with new Sensor Cassette and confirm. | B | SWO-196832 | Yes | Found calibration values Lac sensitivity was low, tri to calibrate but problem was remining same, need to replace with new Sensor Cassette and confirm.
ABL9 | Level 1 gelöscht und als Level 3 angelegt . Dann QC gemessen , System hat sich entsperrt, QC wieder richtig konfiguriert, Kalibration und QC durchgeführt | Level 1 gelöscht und als Level 3 angelegt . Dann QC gemessen , System hat sich entsperrt, QC wieder richtig konfiguriert, Kalibration und QC durchgeführt | SWO-196026 | No | Level 1 deleted and created as Level 3 . Then QC measured, system has been unlocked, QC reconfigured correctly, calibration and QC performed | Level 1 deleted and created as Level 3 . Then QC measured, system has been unlocked, QC reconfigured correctly, calibration and QC performed
ABL90 FLEX | Abfallschlauch und Ventilschlauch überprüft - kein Flüssigkeitsaustritt. Abfallschlauch am Waste Connector nachgesetzt. ABL manuell gespült. Pinch Valve im Service überprüft - schaltet einwandfrei. Oxi Complete Test ok. pO2 Complete und Pressure Test ok. tHb Kalibration ausgeführt (Fcuv = 101,2%). Einlassdichtung erneuert. LP aus Kundenbestand erneuert. | Abfallschlauch und Ventilschlauch überprüft - kein Flüssigkeitsaustritt. Abfallschlauch am Waste Connector nachgesetzt. ABL manuell gespült. Pinch Valve im Service überprüft - schaltet einwandfrei. Oxi Complete Test ok. pO2 Complete und Pressure Test ok. tHb Kalibration ausgeführt (Fcuv = 101,2%). Einlassdichtung erneuert. LP aus Kundenbestand erneuert. | SWO-196017 | No | Waste hose and valve hose checked - no liquid leakage. waste hose on the Waste Connector. ABL manually rinsed. Pinch Valve checked in service - switches flawlessly. Oxi Complete Test ok. pO2 Complete and Pressure Test ok. tHb calibration (Fcuv = 101.2%). Inlet seal renewed. LP from customer base. | Waste hose and valve hose checked - no liquid leakage. waste hose on the Waste Connector. ABL manually rinsed. Pinch Valve checked in service - switches flawlessly. Oxi Complete Test ok. pO2 Complete and Pressure Test ok. tHb calibration (Fcuv = 101.2%). Inlet seal renewed. LP from customer base. | Cleaned | Fluidic pathway | Fluidic pathway | Wetability
ABL90 FLEX PLUS |  | B | D29D3A3D621B79BA | Yes |  |  | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Replaced battery and test ok. | Replaced battery and test ok. | SWO-196427 | No | Replaced battery and test ok. | Replaced battery and test ok. | Defect | HW | Battery | Battery
ABL800 BASIC | Replaced laser scan unit from a working ABL800, now error is rectified | Replaced laser scan unit from a working ABL800, now error is rectified | SWO-197040 | No | Replaced laser scan unit from a working ABL800, now error is rectified | Replaced laser scan unit from a working ABL800, now error is rectified | Defect | HW | Autocheck module | Barcode reader, Autocheck Module
ABL837 FLEX | Thoroughly dust out autocheck module, clean mirror on laser scanning unit, fit new Autocheck fan. On power on and calibration had leak currents detected, carried out extensive tests to pinpoint cause of leak currents, found nothing conclusive but still had leak currents. Stripped and cleaned pH chamber and seals. Perform 2 x 1 point calibration, and run all 4 levels of QC to clear errors. | Thoroughly dust out autocheck module, clean mirror on laser scanning unit, fit new Autocheck fan. On power on and calibration had leak currents detected, carried out extensive tests to pinpoint cause of leak currents, found nothing conclusive but still had leak currents. Stripped and cleaned pH chamber and seals. Perform 2 x 1 point calibration, and run all 4 levels of QC to clear errors. | SWO-196808 | No | Thoroughly dust out autocheck module, clean mirror on laser scanning unit, fit new Autocheck fan. On power on and calibration had leak currents detected, carried out extensive tests to pinpoint cause of leak currents, found nothing conclusive but still had leak currents. Stripped and cleaned pH chamber and seals. Perform 2 x 1 point calibration, and run all 4 levels of QC to clear errors. | Thoroughly dust out autocheck module, clean mirror on laser scanning unit, fit new Autocheck fan. On power on and calibration had leak currents detected, carried out extensive tests to pinpoint cause of leak currents, found nothing conclusive but still had leak currents. Stripped and cleaned pH chamber and seals. Perform 2 x 1 point calibration, and run all 4 levels of QC to clear errors. | Defect | HW | Autocheck module | FAN | OOB
ABL825 FLEX | Fehlersuche: Nach Reinigung der Fließwege Oxi-Complete-Test absolviert, Neon Intensity too low, Neo Position failure => Modul ersetzt und Garantie beantragt. Abgleich und OCT, OK. Kalibrationen, tHb-Kal, QC Messungen OK | Fehlersuche: Nach Reinigung der Fließwege Oxi-Complete-Test absolviert, Neon Intensity too low, Neo Position failure => Modul ersetzt und Garantie beantragt. Abgleich und OCT, OK. Kalibrationen, tHb-Kal, QC Messungen OK | SWO-196232 | No | Troubleshooting: After cleaning the flow paths Completed Oxi-Complete-Test, Neon Intensity too low, Neo Position failure => Module replaced and warranty requested. Matching and OCT, OK. Calibrations, tHb-Kal, QC measurements OK | Troubleshooting: After cleaning the flow paths Completed Oxi-Complete-Test, Neon Intensity too low, Neo Position failure => Module replaced and warranty requested. Matching and OCT, OK. Calibrations, tHb-Kal, QC measurements OK | Defect | HW | Oxi module | Oxi module | Neon lamp
ABL810 FLEX | Oximetriás tesztek hibásak. Hemolizátor csere után a tesztek megfelelőek. Kalibrálás, tHb kalibrálás, QC mérés. 4-5 méréssel tesztelve a készülék. | Oximetriás tesztek hibásak. Hemolizátor csere után a tesztek megfelelőek. Kalibrálás, tHb kalibrálás, QC mérés. 4-5 méréssel tesztelve a készülék. | SWO-196595 | No | Oximetry tests are incorrect. After a haemolytic replacement, the tests are appropriate. Calibration, tHb calibration, QC measurement. 4-5 is tested by measuring the device. | Oximetry tests are incorrect. After a haemolytic replacement, the tests are appropriate. Calibration, tHb calibration, QC measurement. 4-5 is tested by measuring the device. | Defect | HW | Oxi module | Hemolyzer
ABL810 FLEX | Upon arrival the customers stated complaint was verified. Checked the fluidics module for leaks and any worn tubing with all tubing looking ok. Checked all tubing on the inlet module and discovered the tubing on the inlet heater was loose and leaking slightly, Also, the inlet gasket looked to have been past it's replacment date. | Replaced inlet heater tubing, and inlet gasket (provided by the customer) and flushed out the inlet probe. Performed all necessary inlet calibrationss in service mode and adjusted the unit. | D29CC88062227CEF | No | Upon arrival the customers stated complaint was verified. Checked the fluidics module for leaks and any worn tubing with all tubing looking ok. Checked all tubing on the inlet module and discovered the tubing on the inlet heater was loose and leaking slightly, Also, the inlet gasket looked to have been past it's replacment date. | Replaced inlet heater tubing, and inlet gasket (provided by the customer) and flushed out the inlet probe. Performed all necessary inlet calibrationss in service mode and adjusted the unit. | Defect | HW | Inlet module | Tubings
ABL825 FLEX | checked data logs for errors, checkec valves and liquid sensors, checked electrodes and electrode chambers, checked electrode membranes, | salt deposit at pump tubing, removed and repalced pump tube, removed modules and cleaned valves, cleaned electrode chambers and replaced all electrode membranes, | D29D54E9623F83FC | No | checked data logs for errors, checkec valves and liquid sensors, checked electrodes and electrode chambers, checked electrode membranes, | salt deposit at pump tubing, removed and repalced pump tube, removed modules and cleaned valves, cleaned electrode chambers and replaced all electrode membranes, | Defect | HW | pH/BG Module | Pump tubes | Email
ABL820 FLEX | Error code 476 & 293 on blood samples. All calibrations and quality controls are in. | Used Ref and pH electrode from stock to determine pH electrode is defective. Provided lot # of pH to customer Service to determine if warranty replacement. Replacement pH electrode is being shipped overnite. Will record new electrode on SWO and pick up failed electrode to be shipped back to Westlake. | D29D0FE662380F39 | No | Error code 476 & 293 on blood samples. All calibrations and quality controls are in. | Used Ref and pH electrode from stock to determine pH electrode is defective. Provided lot # of pH to customer Service to determine if warranty replacement. Replacement pH electrode is being shipped overnite. Will record new electrode on SWO and pick up failed electrode to be shipped back to Westlake. | Assisted Replacement | Electrode | Electrode | pH | 476
ABL825 FLEX | Viewed QC logs - QC intermittently getting sample errors. 14 pole cable in auto check crimped. | Replaced the cable. | D29CC3E2621E0421 | No | Viewed QC logs - QC intermittently getting sample errors. 14 pole cable in auto check crimped. | Replaced the cable. | Defect | HW | Autocheck module | FPC Cable
ABL825 FLEX | Checked data logs for errors. Zero current errors intermittently. All other checks ok. | Completed zero current calibration in service. Restart calibrations. | D29CEC1D622142DA | No | Checked data logs for errors. Zero current errors intermittently. All other checks ok. | Completed zero current calibration in service. Restart calibrations. | Adjustment | HW | pH/BG Module | Zero current
ABL825 FLEX |  | B | D29CDFFD62185F9D | No |  |  | Defect | HW | Inlet module | Nipple kit
ABL835 FLEX | Nach SW Update keine Auffälligkeiten bei Lac mehr feststellbar, Glu Ansprachefehler, Glu neu membraniert, Messkammer, Pumpenschläuche und Ventile überprüft, Glu einkalibriert (Lac weiterhin unauffällig), Kal. Und QC ohne Beanstandung | Nach SW Update keine Auffälligkeiten bei Lac mehr feststellbar, Glu Ansprachefehler, Glu neu membraniert, Messkammer, Pumpenschläuche und Ventile überprüft, Glu einkalibriert (Lac weiterhin unauffällig), Kal. Und QC ohne Beanstandung | SWO-196203 | No | After SW Update no more abnormalities can be detected at Lac, Glu speech error, Glu re-membraneed, measuring chamber, pump hoses and valves checked, Glu calibrated (Lac still inconspicuous), Kal. And QC without complaint | After SW Update no more abnormalities can be detected at Lac, Glu speech error, Glu re-membraneed, measuring chamber, pump hoses and valves checked, Glu calibrated (Lac still inconspicuous), Kal. And QC without complaint | Assisted Replacement | Electrode | Membrane | Glu
ABL80 FLEX CO-OX | Überprüfung Fließwege Austausch Schlauchsatz und Pumpen tHb Kal LP erneuert ( Fälligkeit lt. Plan) | Überprüfung Fließwege Austausch Schlauchsatz und Pumpen tHb Kal LP erneuert ( Fälligkeit lt. Plan) | SWO-196536 | No | Review flow paths Replacement hose set and pumps tHb Kal LP renewed (due date according to plan) | Review flow paths Replacement hose set and pumps tHb Kal LP renewed (due date according to plan)
ABL80 CO-OX | Zero check in service not stable. One pin on SC connector is black. | Replaced SC Connector board. Perform Zero check in service. Holds zero and is stable. | D29DF369D8237A54 | Yes | Zero check in service not stable. One pin on SC connector is black. | Replaced SC Connector board. Perform Zero check in service. Holds zero and is stable.
ABL90 FLEX | Done data log download as requested. Set the analyzers time to correct time as it was an hour ahead. Ran 5x samples and all ok. | Done data log download as requested. Set the analyzers time to correct time as it was an hour ahead. Ran 5x samples and all ok. | SWO-196693 | No | Done data log download as requested. Set the analyzers time to correct time as it was an hour ahead. Ran 5x samples and all ok. | Done data log download as requested. Set the analyzers time to correct time as it was an hour ahead. Ran 5x samples and all ok. | No fault found | No fault found | No fault found | No fault found
ABL90 FLEX | die Netzwerkkonfiguration auf Windowsebene war nicht mehr vorhanden, erneut eingegeben und auf CF Karte gesichert, Neustart abgewartet, Verbindung i.O. | die Netzwerkkonfiguration auf Windowsebene war nicht mehr vorhanden, erneut eingegeben und auf CF Karte gesichert, Neustart abgewartet, Verbindung i.O. | SWO-196795 | No | the network configuration at the Windows level was no longer present, re-entered and backed up to CF card, restart awaited, connection i.O. | the network configuration at the Windows level was no longer present, re-entered and backed up to CF card, restart awaited, connection i.O. | Setup | His/lis | His/lis | His/lis
ABL90 FLEX PLUS | Sampler Tray getauscht Gerät überprüft | Sampler Tray getauscht Gerät überprüft | SWO-196621 | No | Sampler Tray exchanged Device checked | Sampler Tray exchanged Device checked | Defect | HW | Sampler Mixer | Mixer Cover
AQT90 FLEX | 更换902-924一个正常。(by tim) | 更换902-924一个正常。(by tim) | SWO-195945 | No | Replace the 902-924 one normal. (by tim) | Replace the 902-924 one normal. (by tim) | Defect | HW | Analyzer Control Module | Barcode reader, display | BR2
AQT90 FLEX | 更换903-364和902-924后修复(By Ran) | 更换903-364和902-924后修复(By Ran) | SWO-196500 | No | Replacement 903-364 and 902-924 post-repair (By Ran) | Replacement 903-364 and 902-924 post-repair (By Ran) | Defect | HW | Process Module | Optical Unit | 124
ABL827 FLEX |  | B | D29DE963D807FBFE | Yes
ABL825 FLEX | Einlass gereinigt und Fließwege gespült, Messkammer gereinigt, pH Dichtring erneuert, zwei Kalibrierungen und QC's durchgeführt. | Einlass gereinigt und Fließwege gespült, Messkammer gereinigt, pH Dichtring erneuert, zwei Kalibrierungen und QC's durchgeführt. | SWO-196926 | No | Inlet cleaned and flow paths rinsed, measuring chamber cleaned, pH sealing ring renewed, two calibrations and QC's performed. | Inlet cleaned and flow paths rinsed, measuring chamber cleaned, pH sealing ring renewed, two calibrations and QC's performed. | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
AQT90 FLEX |  | B | D29E19B4D858889A | No |  |  | Defect | HW | Aspiration Module | Needle | Techline
AQT90 FLEX | Ersatzmonitor aus Demogerät testweise eingebaut. Kein Erfolg. Steckverbindungen geprüft, keine Auffälligkeiten. CPU Unit ausgebaut und von anderen Komponenten getrennt um Startverhalten der CPU zu testen. Keine Reaktion, kein Booten. CPU Unit voraussichtlich defekt. Monitor zurückgebaut. CPU Unit bestellt. Gerät mitgenommen um in GS17 zu reparieren wenn CPU Unit geliefert wird. Zubehör und ext. Drucker vom Kunden mitgenommen. | Der Kunde bittet um einen Service. | SWO-197335 | No | Replacement monitor from demo device installed on a test basis. No success. Connectors checked, no abnormalities. CPU Unit and disconnected from other components to test THE START behavior of the CPU. No reaction, no booting. CPU unit is expected to be defective. Monitor. CPU Unit. device to repair in GS17 when CPU Unit is supplied. Accessories and ext. printer taken by the customer. | The customer asks for a service. | Revisit required | HW | Analyzer Control Module | CPU
ABL80 FLEX BASIC | 更换902-972一个正常。(by tim ) | 更换902-972一个正常。(by tim ) | SWO-196800 | No | Replace 902-972 one normal. (by tim) | Replace 902-972 one normal. (by tim)
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-196840 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换测试卡正常。 | 更换测试卡正常。 | SWO-196859 | No | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC | 更换902-972后修复(by tim) | 更换902-972后修复(by tim) | SWO-196285 | No | Repair after replacement 902-972 (by tim) | Repair after replacement 902-972 (by tim)
ABL80 CO-OX | Examined sensor cassette and solution pack. Customer has replaced SC and SP. | Replaced 2 sensor cassettes and solution pack, for troubleshooting purposes. Performed zero check procedure. Verified operation of ABL80. | D2A0DA0DC453E1F1 | No | Examined sensor cassette and solution pack. Customer has replaced SC and SP. | Replaced 2 sensor cassettes and solution pack, for troubleshooting purposes. Performed zero check procedure. Verified operation of ABL80.
ABL90 FLEX | defekten Akku (910-311P0129N008) ausgebaut neuen Akku (910-311P0137N064) eingebaut Gerät kalibriert und überprüft | defekten Akku (910-311P0129N008) ausgebaut neuen Akku (910-311P0137N064) eingebaut Gerät kalibriert und überprüft | SWO-196785 | No | defective battery (910-311P0129N008) removed new battery (910-311P0137N064) built-in Device calibrated and checked | defective battery (910-311P0129N008) removed new battery (910-311P0137N064) built-in Device calibrated and checked | Defect | HW | Battery | Battery | 1559
ABL90 FLEX | SC (946-008 R0539) installiert und (946-008 R0545) deinstalliert, System und Funktion i.O. | SC (946-008 R0539) installiert und (946-008 R0545) deinstalliert, System und Funktion i.O. | SWO-196994 | No | SC (946-008 R0539) installed and uninstalled (946-008 R0545), system and function I.O. | SC (946-008 R0539) installed and uninstalled (946-008 R0545), system and function I.O. | Assisted Replacement | SC | SC | pH | Wetability
ABL90 FLEX PLUS | 2 mal SP reklamiert - Erläuterung etc. Analysator gespült und überprüft - Erneuerung Nadel (gekrümmt) und Einlassdichtung - Funktionsüberprüfung etc. | 2 mal SP reklamiert - Erläuterung etc. Analysator gespült und überprüft - Erneuerung Nadel (gekrümmt) und Einlassdichtung - Funktionsüberprüfung etc. | SWO-196876 | No | 2 times SP complained - explanation etc. Analyzer rinsed and checked - renewal needle (curved) and inlet seal - function check etc. | 2 times SP complained - explanation etc. Analyzer rinsed and checked - renewal needle (curved) and inlet seal - function check etc. | Replaced | HW | Auto Inlet module | Probe bend
ABL90 FLEX PLUS | De Inletmodule compleet heb vervangen vervolgens heb gecontroleerd, OK. | De Inletmodule compleet heb vervangen vervolgens heb gecontroleerd, OK. | SWO-196913 | No | The Inletmodule has completely replaced then checked, OK. | The Inletmodule has completely replaced then checked, OK. | User error | HW | Auto Inlet module | Broken Capillary
ABL90 FLEX PLUS |  | B | D29F27914E43728D | No |  |  | Assisted Replacement | SC | SC | Lac | Techline
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-198088 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | pH | 179
ABL90 | SUSTITUCIÓN DE FUENTE DE ALIMENTACIÓN | B | D29FC91D4E1B2584 | No | REPLACING THE POWER SUPPLY |  | Defect | HW | Power Supply | Power Supply | LEH Mail to Alex
AQT90 FLEX | 更换910-304后修复(by tim) | 更换910-304后修复(by tim) | SWO-196408 | No | Repair after replacement 910-304 (by tim) | Repair after replacement 910-304 (by tim) | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | 重新定标后修复(by tim) | 重新定标后修复(by tim) | SWO-196526 | No | Repair after re-calibration (by tim) | Repair after re-calibration (by tim) | Defect | HW | Process Module | Optical Unit | 124
ABL827 FLEX |  | B | D29F7EFD4E374E3A | No |  |  | Revisit required | HW | Oxi module | Unknown | COtHb & MetHb
ABL835 FLEX |  | B | SWO-196868 | No |  |  | Unknown | Unknown | Unknown | Unknown | LEH Mail
ABL815 FLEX | Datensicherung, Monitor überprüft (HWParam 4.10.007), Firmware installiert | Datensicherung, Monitor überprüft (HWParam 4.10.007), Firmware installiert | SWO-196953 | No | Data backup, monitor checked (HWParam 4.10.007), firmware installed | Data backup, monitor checked (HWParam 4.10.007), firmware installed | Defect | HW | DMS_PC | Touch screen
ABL825 FLEX |  | B | D29F284A4E07DAB4 | Yes |  |  | Re classified RMED | Logistical upSWO
ABL825 FLEX |  | B | D29F51F54E33A178 | No |  |  | Cleaned | Fluidic pathway | pH/BG Module | Measuring chamber
ABL800 FLEX | Drane and remove the Oxy Module. After Half hour again refix and calibrate, Cleaning, Protean removal, Oxy complete test and tHb Calibration done and pass. Analyser seriviced and calibrated properly. Handed over in a fine working condition. | Drane and remove the Oxy Module. After Half hour again refix and calibrate, Cleaning, Protean removal, Oxy complete test and tHb Calibration done and pass. Analyser seriviced and calibrated properly. Handed over in a fine working condition. | SWO-196881 | No | Drane and remove the Oxy Module. After Half hour again refix and calibrate, Cleaning, Protean removal, Oxy complete test and tHb Calibration done and pass. Analyser seriviced and calibrated properly. Handed over in a fine working condition. | Drane and remove the Oxy Module. After Half hour again refix and calibrate, Cleaning, Protean removal, Oxy complete test and tHb Calibration done and pass. Analyser seriviced and calibrated properly. Handed over in a fine working condition. | Cleaned | HW | Oxi module | Oxi module
ABL825 FLEX | PH Elektrode untere Dichtung fehlt komplett, Blut über Überlauf PH Modul ausgelaufen, Blutreste entfernt Dichtung PH Elektrode ersetzt, PH Fehler seit gestern, PCO2 Fehler seit membranieren heute morgen 7:00, pCO2 Membran ersetzt Kontakte der Elektrode gereinigt, Empf bleibt < 10%, PCo2 Elektrode ersetzt, Kal und QC durchgeführt | PH Elektrode untere Dichtung fehlt komplett, Blut über Überlauf PH Modul ausgelaufen, Blutreste entfernt Dichtung PH Elektrode ersetzt, PH Fehler seit gestern, PCO2 Fehler seit membranieren heute morgen 7:00, pCO2 Membran ersetzt Kontakte der Elektrode gereinigt, Empf bleibt < 10%, PCo2 Elektrode ersetzt, Kal und QC durchgeführt | SWO-197033 | No | PH electrode lower seal is missing completely, blood spilled over overflow PH module, blood residue removed seal PH electrode replaced, PH fault since yesterday, PCO2 fault since membraneing this morning 7:00, pCO2 membrane replaced contacts of the electrode cleaned, mpf remains 10%, PCo2 electrode replaced, Kal and QC performed | PH electrode lower seal is missing completely, blood spilled over overflow PH module, blood residue removed seal PH electrode replaced, PH fault since yesterday, PCO2 fault since membraneing this morning 7:00, pCO2 membrane replaced contacts of the electrode cleaned, mpf remains 10%, PCo2 electrode replaced, Kal and QC performed | Assisted Replacement | Electrode | Electrode | pCO2
AQT90 FLEX |  | B | D29F82994E635B15 | Yes |  |  | Re classified RMED | SWU
ABL90 FLEX | neue Einlaßdichtung eingesetzt, LP installiert, 1x LP hinterlegt, Kal. + QC i.O. | neue Einlaßdichtung eingesetzt; LP installiert; 1x LP hinterlegt; Kal. + QC i.O. | SWO-197367 | No | new inlet seal inserted, LP installed, 1x LP deposited, Kal. + QC in the first place | new inlet seal inserted; LP installed; 1x LP deposited; Kal. + QC in the first place | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX PLUS | Realigned sensor flag so that it did not obstruct closing of bay | advised to replace SP | SWO-197380 | No | Realigned sensor flag so that it did not obstruct closing of bay | advised to replace SP | Adjustment | HW | Basic Unit | Sensor door
ABL90 FLEX PLUS | Replaced inlet gasket and start up tests ok. | Replaced inlet gasket and start up tests ok. | SWO-197645 | No | Replaced inlet gasket and start up tests ok. | Replaced inlet gasket and start up tests ok. | Defect | HW | Auto Inlet module | Inlet Gasket
ABL90 FLEX PLUS | ABL90 manuell gespült Hemolyzer sowie Ventilschlauch erneuert, tHb kalibriert danach Glucose bei QC Level C zu tief, SCi + LP aus Kundenbestand erneuert nach SCi Tausch Glu im Level C um 2 mg dL zu hoch SCi reinstalliert Gerät kalibriert und überprüft | ABL90 manuell gespült Hemolyzer sowie Ventilschlauch erneuert, tHb kalibriert danach Glucose bei QC Level C zu tief, SCi + LP aus Kundenbestand erneuert nach SCi Tausch Glu im Level C um 2 mg dL zu hoch SCi reinstalliert Gerät kalibriert und überprüft | SWO-197194 | No | ABL90 manually rinsed Hemolyzer and valve hose renewed, tHb calibrated glucose at QC Level C too deep, SCi + LP renewed from customer base after SCi exchange glu at level C reinstalled by 2 mg dL too high SCi Device calibrated and checked | ABL90 manually rinsed Hemolyzer and valve hose renewed, tHb calibrated glucose at QC Level C too deep, SCi + LP renewed from customer base after SCi exchange glu at level C reinstalled by 2 mg dL too high SCi Device calibrated and checked | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Customer was properly running both test and validating the results, however they were under the impression that there was a script that was supposed to walk them through both parts. | Directed customer to the OM sect for running tHb cal verifying results. | D2A0DDA5C40B2B06 | Yes | Customer was properly running both test and validating the results, however they were under the impression that there was a script that was supposed to walk them through both parts. | Directed customer to the OM sect for running tHb cal verifying results. | Re classified RMED | Customer training
ABL835 FLEX | Elektrolyt-Ersatzmodul aus GS eingesetzt, SW angepasst und mehrfach einkalibriert, Kal. und QC i.O., Kunde wird weiteres Verhalten beobachten | Frau Wagener bittet um Rückruf Terminabsprache von Hr. Poisel | SWO-197799 | No | Electrolyte replacement module from GS used, SW adapted and multiple calibrated, Kal. and QC i.O., customer will observe further behavior | Ms. Wagener asks for recall appointment from Mr. Poisel | Defect | HW | El/Met Module | El/Met Module | Electrolyte and Met
AQT90 FLEX | Gerät überprüft, Cups entfernt, defekte Reagenzienkassette ersetzt, Messung o.k. | Gerät überprüft, Cups entfernt, defekte Reagenzienkassette ersetzt, Messung o.k. | SWO-197314 | No | Device checked, cups removed, defective reagent cartridge replaced, measurement ok | Device checked, cups removed, defective reagent cartridge replaced, measurement ok | Defect | Consumables | Solution Pack | Cup Stack
AQT90 FLEX | Einlaufverfahren Systemtest und QC durchgeführt. | Einlaufverfahren Systemtest und QC durchgeführt. | SWO-197771 | No | inlet procedure system test and QC. | inlet procedure system test and QC. | Defect | HW | Analyzer Control Module | CPU
AQT90 FLEX | Netzwerkeinstellungen an AQT eingegeben. Verbindung zu H+S wurde hergestellt. Med. Technik Herr Babel , Stationsleiter Herr Miehe und die IT wurden informiert und überzeugten sich vor Ort über die Funktionalität der Anbindung. | B | SWO-197773 | Yes | network settings to AQT. Connection to H+S has been established. Med. Technik Mr. Babel , station manager Mr. Miehe and the IT were informed and convinced themselves on site about the functionality of the connection. |  | Re classified RMED
AQT90 FLEX | Monovette aus Wet-Section entfernt Probelauf | Der Kunde bittet um einen Service und Rückruf. Der Kunde möchte informiert werden, wenn Kosten anfallen sollten. | SWO-197387 | No | Removed monovette from Wet-Section Rehearsal | The customer asks for a service and callback. The customer would like to be informed if costs should be incurred. | User error | User error | User error | User error | Vacutainer fallen from holder
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-197694 | No | Repair after replacing test card | Repair after replacing test card
ABL90 FLEX PLUS | Verbindungsschlauch getauscht. Einlassdichtung gereinigt. Alle Prozeduren in Service durchgeführt. Neustart 1 Std. | Verbindungsschlauch getauscht. Einlassdichtung gereinigt. Alle Prozeduren in Service durchgeführt. Neustart 1 Std. | SWO-197810 | No | Connection hose is replaced. Inlet seal cleaned. All procedures performed in service. Restart 1 hr. | Connection hose is replaced. Inlet seal cleaned. All procedures performed in service. Restart 1 hr. | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX | 更换902-924一个正常。(by tim) | 更换902-924一个正常。(by tim) | SWO-197122 | No | Replace the 902-924 one normal. (by tim) | Replace the 902-924 one normal. (by tim) | Defect | HW | Basic Unit | Barcode reader, display
ABL80 FLEX BASIC | 更换903-334，902-972各一个正常。(by tim) | 更换903-334，902-972各一个正常。(by tim) | SWO-197454 | No | Replace 903-334, 902-972 each normal. (by tim) | Replace 903-334, 902-972 each normal. (by tim)
ABL90 FLEX | 更换测试卡正常。 | 更换测试卡正常。 | SWO-197494 | Yes | The replacement test card is OK. | The replacement test card is OK.
ABL80 FLEX BASIC |  | B | SWO-197594 | No
ABL80 FLEX CO-OX OSM | 代替え器設置 | B | SWO-197646 | No | Installation of an alternative device
ABL80 FLEX BASIC | Sensor board replaced. Zeroing calibration done and found heater not functioning. Need to replace Signal Data board. Asked from Pune. will be coming till Wednesday. | Observation of analyser done. And found working fine. Analyser calibrated and serviced properly. Handed over in a fine working condition. | SWO-199081 | No | Sensor board replaced. Zeroing calibration done and found heater not functioning. Need to replace Signal Data board. Asked from Pune. will be coming till Wednesday. | Observation of analyser done. And found working fine. Analyser calibrated and serviced properly. Handed over in a fine working condition.
ABL90 | Customer said he received temperature errors 1065 and 1066. Prior to my arrival the errors cleared themselves and the analyzer was now working properly. Inspected SC pins and found all ok. Ran temperature calibration in the Service menu. Customer said they tried flushing the unit but no water came out of the inlet probe. | Customer changed the SC, SP, inlet gasket, and inlet probe. The unit could have had a clot. There were several options that they wanted to add to the drop down menus associated with each sample button.The changes were made to this analyzer. Copied the setup to all the other analyzers. Updated to software version 3.3. | D2A9DF2EFE5BE238 | No | Customer said he received temperature errors 1065 and 1066. Prior to my arrival the errors cleared themselves and the analyzer was now working properly. Inspected SC pins and found all ok. Ran temperature calibration in the Service menu. Customer said they tried flushing the unit but no water came out of the inlet probe. | Customer changed the SC, SP, inlet gasket, and inlet probe. The unit could have had a clot. There were several options that they wanted to add to the drop down menus associated with each sample button. The changes were made to this analyzer. Copied the setup to all the other analyzers. Updated to software version 3.3. | Cleaned | Fluidic pathway | Fluidic pathway | Clot | 166
ABL90 FLEX | 更换试剂包正常。 | 更换试剂包正常。 | SWO-202723 | Yes | The replacement kit is OK. | The replacement kit is OK.
AQT90 FLEX | 重新定标后修复(By Ran) recal, OK | 重新定标后修复(By Ran) recal, OK | SWO-197447 | No | Re-calibration repair (By Ran) recal, OK | Re-calibration repair (By Ran) recal, OK | Defect | HW | Process Module | Optical Unit | 131
ABL825 FLEX | Upon arrival, found unit had been running low on Peer 2. The past four days unit running well within Peer 2 ranges. Will install version 6.17 in unit. | No correction needed. Customer establishing own 2SD range at beginning of lot and not establishing new ranges monthly or when any PCO2 maintenance is done. This is causing False positives on unit. Explained to customer to use PEER2 ranges going forward. Will have CAS come in and review Protocol. | D2A44760260E676B | No | Upon arrival, found unit had been running low on Peer 2. The past four days unit running well within Peer 2 ranges. Will install version 6.17 in unit. | No correction needed. Customer establishing own 2SD range at beginning of lot and not establishing new ranges monthly or when any PCO2 maintenance is done. This is causing False positives on unit. Explained to customer to use PEER2 ranges going forward. Will have CAS come in and review Protocol. | Setup | Training | Training | Training
ABL835 FLEX | Tested oximetry system. Found halogen lamp above 3800. | Adjusted halogen lamp to 3150. Ran complete oximetry test with pass results. Restart calibrations. | D2A49606262B62E4 | No | Tested oximetry system. Found halogen lamp above 3800. | Adjusted halogen lamp to 3150. Ran complete oximetry test with pass results. Restart calibrations. | Adjustment | HW | Oxi module | Lamp unit | Halogen
ABL80 FLEX CO-OX | Gerät konfiguriert | Gerät konfiguriert | SWO-197929 | No | Device configured | Device configured
ABL90 FLEX | Touch regiert vor allem unten links nicht korrekt. Verschobene Touchfunktion trotz mehrmaliger Touch Kalibration. ( siehe vorherige SWOs ) Monitor erneuert. ewf commit durchgeführt. | Touch regiert vor allem unten links nicht korrekt. Verschobene Touchfunktion trotz mehrmaliger Touch Kalibration. ( siehe vorherige SWOs ) Monitor erneuert. ewf commit durchgeführt. | SWO-198005 | No | Touch does not govern correctly, especially at the bottom left. Shifted touch function despite several touch calibrations. ( see previous SWOs ) Monitor renewed. ewf commit. | Touch does not govern correctly, especially at the bottom left. Shifted touch function despite several touch calibrations. ( see previous SWOs ) Monitor renewed. ewf commit. | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX | Kaltstartphase abgewartet und LP installiert, Kal und QCs durchgeführt | Kaltstartphase abgewartet und LP installiert, Kal und QCs durchgeführt | SWO-198048 | Yes | Cold start phase waited and LP installed, Kal and QCs performed | Cold start phase waited and LP installed, Kal and QCs performed | Re classified RMED | Logistical upSWO
ABL90 FLEX | Einlassdichtung erneuert, System und Funktion i.O. | Einlassdichtung erneuert, System und Funktion i.O. | SWO-198116 | No | Inlet seal renewed, system and function in the | Inlet seal renewed, system and function in the | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Sample Mixer Unit erneuert Gerät kalibriert und überprüft | Sample Mixer Unit erneuert Gerät kalibriert und überprüft | SWO-197766 | No | Sample Mixer Unit renewed Device calibrated and checked | Sample Mixer Unit renewed Device calibrated and checked | Defect | HW | Sampler Mixer | Sampler Mixer
ABL90 FLEX | Konfiguration überprüft, in Ordnung. Trance mit astmnet erstellt und Herrn Handl, Synlab zugeschickt. Benutzer-ID auf Trance, aber nicht in Conworx sichtbar. Rückruf Herr Korte, Siemens. Laut Herrn Korte ist Benutzer-ID in der Datenbank sichtbar, wird aber nicht an Synlab weitergeleitet. Herr Korte übernimmt die weitere Fehlersuche. | Konfiguration überprüft, in Ordnung. Trance mit astmnet erstellt und Herrn Handl, Synlab zugeschickt. Benutzer-ID auf Trance, aber nicht in Conworx sichtbar. Rückruf Herr Korte, Siemens. Laut Herrn Korte ist Benutzer-ID in der Datenbank sichtbar, wird aber nicht an Synlab weitergeleitet. Herr Korte übernimmt die weitere Fehlersuche. | SWO-198122 | No | Configuration checked, all right. Trance created with astmnet and sent to Mr. Handl, Synlab. User ID on trance, but not visible in Conworx. Recall Mr. Korte, Siemens. According to Mr. Korte, user ID is visible in the database, but is not forwarded to Synlab. Mr Korte will continue to search for the bug. | Configuration checked, all right. Trance created with astmnet and sent to Mr. Handl, Synlab. User ID on trance, but not visible in Conworx. Recall Mr. Korte, Siemens. According to Mr. Korte, user ID is visible in the database, but is not forwarded to Synlab. Mr Korte will continue to search for the bug. | No fault found | His/lis | His/lis | His/lis
ABL90 FLEX PLUS |  | B | D2A5911B9C2B140F | No |  |  | Defect | HW | Basic Unit | Solution Pack ID PCB
ABL90 FLEX PLUS |  | B | D2A5AF569C546EEC | Yes
ABL90 | Sustitucion de la sensor interface pcb. Actualizacion soft 3.3 Cal, qc | B | D2A5E6FC9C577ED7 | No | Replacement of the pcb interface sensor. Soft update 3.3 Cal, qc |  | Defect | HW | Basic Unit | Sensor Interface PCB
AQT90 FLEX |  | B | SWO-197687 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 没有故障，工程师报错仪器序列号了。(By Ran) wrong SN. CASE closed. | 没有故障，工程师报错仪器序列号了。(By Ran) wrong SN. CASE closed. | SWO-197888 | No | There was no fault and the engineer reported the wrong instrument serial number. (By Ran) wrong SN. CASE closed. | There was no fault and the engineer reported the wrong instrument serial number. (By Ran) wrong SN. CASE closed. | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-119备件一个后修复 | B | SWO-197593 | No | Replace 903-119 spare parts with one post-fix |  | Defect | HW | Process Module | Cup wash unit
AQT90 FLEX |  | B | SWO-197899 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-197688 | No |  |  | Assisted Replacement | Consumables | Test Kit | Cartridge, Tni
ABL800 BASIC | Fehler wurde durch fällige Wartung behoben | Fehler wurde durch fällige Wartung behoben | SWO-198654 | No | Error fixed due maintenance | Error fixed due maintenance | Defect | HW | pH/BG Module | Pump tubes | Assymetri
ABL800 BASIC | 910-332有质量问题，重新开单借910-274、910-332和910-321各一个 | 更换910-332后修复(By Ran) | SWO-197588 | No | 910-332 has quality issues, re-billing borrowing 910-274, 910-332 and 910-321 each | Repair after replacement 910-332 (By Ran) | Revisit required | HW | DMS_PC | CPU | Used analyzer
ABL825 FLEX | 更换LAC电极膜正常。 | 更换LAC电极膜正常。 | SWO-197764 | No | Replacing the LAC electrode film is normal. | Replacing the LAC electrode film is normal. | Assisted Replacement | Electrode | Membrane | Lac
ABL825 FLEX | 更换电滤板备件一个 | 更换电滤板备件一个 | SWO-197695 | No | Replace one of the panel spare parts | Replace one of the panel spare parts | Defect | HW | El/Met Module | PCB, El/Met | 48, LEH contact
ABL825 FLEX | Szennyeselszívó cső eldugult, cső csere után a folyás megszűnt. | B | SWO-197991 | No | Dirty suction tube is blocked, after pipe replacement the discharge is discontinued. |  | Defect | HW | Inlet module | Tubings | Clot
ABL827 FLEX |  | B | D2A5CB9D9C006233 | No |  |  | Setup | SW | His/lis | His/lis
ABL825 FLEX | Halogenlamp justiert, Kal, QC | Halogenlamp justiert, Kal, QC | SWO-198059 | No | Halogenlamp justiert, Kal, QC | Halogenlamp justiert, Kal, QC | Adjustment | HW | Oxi module | Lamp unit | Halogen
ABL80 CO-OX | CONFIRMED ISSUE WHEN ONSITE. | PROCEED TO ADJUST HEMOLYZER SETTINGS AND CALIBRATED tHb AFTERWARDS. | D2A6E5F31232B253 | Yes | CONFIRMED ISSUE WHEN ONSITE. | PROCEED TO ADJUST HEMOLYZER SETTINGS AND CALIBRATED tHb AFTERWARDS.
ABL80 FLEX BASIC | Signal Data board replaced. Zeroing calibration done. SC and SP installed. Sample analysis done. Analyser serviced and calibrated properly. Handed over in a fine working condition. | Observation of analyser done. And found working fine. Analyser calibrated and serviced properly. Handed over in a fine working condition. | SWO-199087 | No | Signal Data board replaced. Zeroing calibration done. SC and SP installed. Sample analysis done. Analyser serviced and calibrated properly. Handed over in a fine working condition. | Observation of analyser done. And found working fine. Analyser calibrated and serviced properly. Handed over in a fine working condition.
ABL90 FLEX | Set both machines back to centralised user use. Site can now push operators. | Set both machines back to centralised user use. Site can now push operators. | SWO-198263 | Yes | Set both machines back to centralised user use. Site can now push operators. | Set both machines back to centralised user use. Site can now push operators. | Re classified RMED | His/Lis settup
ABL90 FLEX | Fehlersuche: Einlassdichtung ersetzt - weiterhin Fehler. System in beide Richtungen gespült, ohne Erfolg. SP ersetzt, ohne Erfolg. Alle Fließwege auf Dichtheit überprüft, Funktion und Leichtgängigkeit des Bügelventils geprüft, ok. Einlassnadel ersetzt, Fluidic Programme im Servicemenü absolviert, weiterhin fehlerhafter Flüssigkeitstransport, diverse Lösungen werden nicht gefördert. Jahreswartung mit Wartungskit vorgezogen (s. gesonderter case) und Hemolyzer auf Durchlässigkeit überprüft. Neustart, nach etwa 5 min wieder in der geführten Fehlerbehebung => SC ersetzt und erneut sämtliche bezogenen Serviceprogramme mehrfach wiederholt, nun (mit neuer SC) regelmäßig Vakuumfehler => SC zurück getauscht. Fluidic Programme nun ok, Neustart, Startphase, Kal und QC Messungen OK. | Ein Vorortservice wird gewünscht | SWO-198446 | No | Troubleshooting: Inlet seal replaced - still errors. System flushed in both directions, to no avail. SP replaced, to no avail. All flow paths checked for leak, function and smoothness of the ironing valve checked, ok. Replacement of the inlet needle, fluidic programs completed in the service menu, further faulty fluid transport, various solutions are not promoted. Annual maintenance with maintenance kit preferred (see separate case) and Hemolyzer checked for permeability. Restart, replaced after about 5 min in the guided troubleshooting => SC and again repeated all service programs several times, now (with new SC) regularly replaced vacuum errors => SC. Fluidic programs now ok, restart, start phase, Kal and QC measurements OK. | A suburban service is desired | Revisit required | SC | SC | FT | Vacum
ABL90 FLEX PLUS |  | B | D2A843C688363DBA | No |  |  | Revisit required | HW | Basic Unit | Sensor Interface PCB | Techline RCZE 78139653
ABL5 | Störung im Rahmen der Wartung beseitigt | Störung im Rahmen der Wartung beseitigt | SWO-198185 | No | Trouble-free maintenance | Trouble-free maintenance
AQT90 FLEX |  | B | SWO-198105 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-198081 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-198200 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
ABL835 FLEX |  | B | SWO-198228 | No |  |  | Defect | HW | Printer Module | Printer Module
ABL800 BASIC | 更换910-332后修复(By Ran) | 更换910-332后修复(By Ran) | SWO-197706 | No | Repair after replacement 910-332 (By Ran) | Repair after replacement 910-332 (By Ran) | Defect | HW | DMS_PC | HDD
ABL825 FLEX |  | B | D2A6E1DF1246AE3E | Yes |  |  | Re classified RMED | Preventive maintenance
ABL835 FLEX | Glu und Messkammer gereinigt, Glu neu membraniert und einkalibriert, Kal. Und QC i.O. | Frau Wagener bittet um Rückruf Terminabsprache von Hr. Poisel | SWO-198310 | No | Glu and measuring chamber cleaned, Glu re-membraneed and calibrated, Kal. And QC i.O. | Ms. Wagener asks for recall appointment from Mr. Poisel | Cleaned | HW | El/Met Module | Measuring chamber
ABL80 CO-OX | The solution pack door is broken again and not pushing the SP in all the way. The analyzer is failing Glu. Paula will let it run and then run the QC later. She said it needs to hydrate. | Replaced molded door hinge assembly. | D2A81FC98808EDAB | No | The solution pack door is broken again and not pushing the SP in all the way. The analyzer is failing Glu. Paula will let it run and then run the QC later. She said it needs to hydrate. | Replaced molded door hinge assembly.
ABL80 OSM | Paula had a possible fluid blockage. I suggested turning it off, remove solution pack and sensor cassette and then turn back on. When it came back up it showed corrupt system file and corrupt parameter file. It gavee all parametes accross the top and said no reportable parameters. Reloaded SW did not solve issue. Replaced SC and it solved the issue. | Replaced SC with customer stock. | D2A81919882BEDA8 | No | Paula had a possible fluid blockage. I suggested turning it off, remove solution pack and sensor cassette and then turn back on. When it came back up it showed corrupt system file and corrupt parameter file. It gavee all parametes accross the top and said no reportable parameters. Reloaded SW did not solve issue. Replaced SC and it solved the issue. | Replaced SC with customer stock.
ABL80 FLEX BASIC | Observation of analyser done. And found working fine. Analyser calibrated and serviced properly. Handed over in a fine working condition. | Observation of analyser done. And found working fine. Analyser calibrated and serviced properly. Handed over in a fine working condition. | SWO-199089 | No | Observation of analyser done. And found working fine. Analyser calibrated and serviced properly. Handed over in a fine working condition. | Observation of analyser done. And found working fine. Analyser calibrated and serviced properly. Handed over in a fine working condition.
ABL80 FLEX CO-OX | Gerät geprüft , Diverse Messungen mit Hypochlorid durchgeführt. Sensorcassette erneuert (Verfallsdatum erreicht) Gerät demontiert , Quetschventil und Fliesswege überprüft. Hemolyzer erneut gespült. Anschließend kal , thb kal und QC durchgeführt. Blutige Messung auch ok. | Service Einsatz | SWO-199148 | No | Device tested, Various measurements performed with hypochloride. Sensor cassette renewed (expiry date reached) Device dismantled, squeezing valve and flow paths checked. Hemolyzer rinsed again. Subsequently kal , thb kal and QC performed. Bloody measurement also ok. | Service use
ABL90 |  | B | D2A95B5EFE2AEBFA | No |  |  | Defect | HW | Basic Unit | SC Door
ABL90 FLEX | Erneuerung des Solutions pack gegen konditioniert Ware | Erneuerung des Solutions pack gegen konditioniert Ware | SWO-198657 | No | Renewal of the solution pack against conditioned goods | Renewal of the solution pack against conditioned goods | Assisted Replacement | SP | SP | Leak current | 1331
ABL90 FLEX | PC Unit getauscht, kein Erfolg. AU PCB Getauscht , kein Erfolg. Netzteil getauscht, kein Erfog. Monitor getauscht, Kein Erfolg. Alle getauschten Module in Demogerät eingebaut. alles o.k. Im Kundengerät Einlassmodul Einheit getauscht, kein Erfolg. Mischer Unit für sPico getauscht, alles o.k. . Gerät bootet wieder normal. Mischer Unit defekt. Alle Module wieder original zurückgebaut. MischerUnit von AU PCB getrennt. Gerät bootet normal . Mischer Unit muß bestellt werden. Einlaufverfahren durchgeführt. | PC Unit getauscht, kein Erfolg. AU PCB Getauscht , kein Erfolg. Netzteil getauscht, kein Erfog. Monitor getauscht, Kein Erfolg. Alle getauschten Module in Demogerät eingebaut. alles o.k. Im Kundengerät Einlassmodul Einheit getauscht, kein Erfolg. Mischer Unit für sPico getauscht, alles o.k. . Gerät bootet wieder normal. Mischer Unit defekt. Alle Module wieder original zurückgebaut. MischerUnit von AU PCB getrennt. Gerät bootet normal . Mischer Unit muß bestellt werden. Einlaufverfahren durchgeführt. | SWO-198687 | No | PC Unit exchanged, no success. AU PCB Exchanged, no success. Power supply replaced, no erfog. Monitor swapped, No success. All replaced modules built into demo device. everything ok. In the customer device inlet module unit exchanged, no success. Mixer unit for sPico exchanged, all OK. . Device boots back to normal. Mixer unit defective. All modules restored to their original. separating MixerUnit from AU PCB. Device boots normally. Mixer unit must be ordered. run-in procedures. | PC Unit exchanged, no success. AU PCB Exchanged, no success. Power supply replaced, no erfog. Monitor swapped, No success. All replaced modules built into demo device. everything ok. In the customer device inlet module unit exchanged, no success. Mixer unit for sPico exchanged, all OK. . Device boots back to normal. Mixer unit defective. All modules restored to their original. separating MixerUnit from AU PCB. Device boots normally. Mixer unit must be ordered. run-in procedures. | Revisit required | HW | Sampler Mixer | Sampler Mixer
ABL90 FLEX | ABL hat gestern Abend und heute morgen wieder den selben Fehler gezeigt, das SP wurde anwenderseitig 2 x neu initialisiert - beim Eintreffen grüne Ampel. Im Servicemenü Pressure Test aufgerufen, ok. Einlassdichtung erneuert => Vakuumfehler, permanent reproduzierbar, zuvor eingesetzte (gestern erneuerte) Einlassdichtung eingesetzt => Vakuumfehler, permanent reproduzierbar, weitere (3. neue) Einlassdichtung eingesetzt => keine weiteren Fehlermeldungen, mehrfach sämtliche Service-Testprogramme absolviert, alle fehlerfrei! Mehrere Blutmessungen Spritze Kapillare abgewartet, alle OK. ServiceDump gezogen, QC Messung OK | ABL hat gestern Abend und heute morgen wieder den selben Fehler gezeigt; das SP wurde anwenderseitig 2 x neu initialisiert - beim Eintreffen grüne Ampel. Im Servicemenü Pressure Test aufgerufen, ok. Einlassdichtung erneuert => Vakuumfehler, permanent reproduzierbar; zuvor eingesetzte (gestern erneuerte) Einlassdichtung eingesetzt => Vakuumfehler, permanent reproduzierbar; weitere (3. neue) Einlassdichtung eingesetzt => keine weiteren Fehlermeldungen, mehrfach sämtliche Service-Testprogramme absolviert, alle fehlerfrei! Mehrere Blutmessungen Spritze Kapillare abgewartet, alle OK. ServiceDump gezogen, QC Messung OK | SWO-198674 | No | ABL showed the same error last night and this morning, the SP was re-initialized 2 times on the user side - when green traffic lights arrived. Called in the pressure test service menu, ok. Inlet seal renewed => Vacuum fault, permanently reproducible, previously used (yesterday renewed) inlet seal used => Vacuum error, permanently reproducible, further (3rd new) inlet seal inserted => no further error messages, multiple completed all service test programs, all error-free! Several blood measurements syringe capillary waited, all OK. ServiceDump pulled, QC measurement OK | ABL showed the same mistake again last night and this morning; the SP has been reinitialized 2 times on the user side - when green traffic lights arrive. Called in the pressure test service menu, ok. Inlet seal renewed => Vacuum fault, permanently reproducible; previously used (yesterday renewed) inlet seal => vacuum fault, permanently reproducible; further (3rd new) inlet seal inserted => no further error messages, completed several times all service test programs, all error-free! Several blood measurements syringe capillary waited, all OK. ServiceDump pulled, QC measurement OK | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Fließwege mit Spülwerkzeug durchgespült. Fließwege wieder frei, verstopft durch Gerinnsel. Verwendung von Sarstedt Probenmehmern. Nach Neustart und erneuter Initialisierungsphase ist Analysator funktionsbereit. | Fließwege mit Spülwerkzeug durchgespült. Fließwege wieder frei, verstopft durch Gerinnsel. Verwendung von Sarstedt Probenmehmern. Nach Neustart und erneuter Initialisierungsphase ist Analysator funktionsbereit. | SWO-198841 | No | Flow paths rinsed with rinsing tool. Flow paths free again, clogged by clots. Use of Sarstedt sample companies. After rebooting and reinitialization, Analyser is ready for use. | Flow paths rinsed with rinsing tool. Flow paths free again, clogged by clots. Use of Sarstedt sample companies. After rebooting and reinitialization, Analyser is ready for use. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 1184
ABL90 FLEX | SC und LP zum Test ersetzt, aber Problem dass Empf. bei Lac = 0 ist, bleibt bestehen es ist bereits die 3. SC mit gleichem Fehlerbild vermutlich Sensor Interface Board defekt - muss bestellt werden Gerät wird zur Reparatur mitgenommen | Gerät wieder zum Kunden gebracht und funktionstüchtig übergeben | SWO-198847 | No | SC and LP replaced for testing, but problem that Empf. at Lac = 0, remains it is already the 3rd SC with the same error image presumably sensor interface board defective - must be ordered Device is taken for repair | Device brought back to the customer and handed over functionally | Revisit required | HW | Basic Unit | Sensor Interface PCB | Lac
ABL90 FLEX PLUS |  | B | D2A83EC6882642F5 | Yes
ABL90 FLEX PLUS | The disable patient record view box was unchecked. | The disable patient record view box was unchecked. | SWO-198836 | No | The disable patient record view box was unchecked. | The disable patient record view box was unchecked. | Setup | SW | His/lis | Remote support
ABL90 FLEX PLUS | De Inletmodule heb vervangen vervolgens heb gecontroleerd. | De Inletmodule heb vervangen vervolgens heb gecontroleerd. | SWO-198828 | No | The Inletmodule have replaced then checked. | The Inletmodule have replaced then checked. | Defect | HW | Auto Inlet module | Auto Inlet module
ABL90 FLEX PLUS | Reconfigured Analyser, checked and tested. Setup Analyser to appear on AQURE and allocated it to Newham - Rainbow Ward. | Reconfigured Analyser, checked and tested. Setup Analyser to appear on AQURE and allocated it to Newham - Rainbow Ward. | SWO-198849 | Yes | Reconfigured Analyser, checked and tested. Setup Analyser to appear on AQURE and allocated it to Newham - Rainbow Ward. | Reconfigured Analyser, checked and tested. Setup Analyser to appear on AQURE and allocated it to Newham - Rainbow Ward. | Re classified RMED | Installation | His/lis | Aqure
ABL90 FLEX PLUS |  | B | D2A8453C884C3DBD | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB | Techline RCZE 78139653
AQT90 FLEX | 更换903-364备件后修复（By Ran） | 更换903-364备件后修复（By Ran） | SWO-198212 | No | Repair after replacement of 903-364 spare parts (By Ran) | Repair after replacement of 903-364 spare parts (By Ran) | Defect | HW | Process Module | Optical Unit | 124
AQT90 FLEX |  | B | SWO-198467 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-198486 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-198490 | No |  |  | Assisted Replacement | Consumables | Test Kit | Cartridge, PCT
AQT90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-198672 | No | Repair after replacing test card | Repair after replacing test card | Defect | Consumables | Test Kit | Cartridge | 5784
AQT90 FLEX | 更换903-364备件后修复(By Ran) | 更换903-364备件后修复(By Ran) | SWO-198082 | Yes | Repair after replacement of 903-364 spare parts (By Ran) | Repair after replacement of 903-364 spare parts (By Ran) | Re classified RMED | OOB |  |  | 131
ABL827 FLEX |  | B | D2A85664881C0918 | Yes
ABL825 FLEX | Servicedump erstellt, Oximetrie-Modul ausgetauscht und Reklamatiert. Service tHb und tHb kal durchgeführt, Kalibriert und QC's durchgeführt. | Servicedump erstellt, Oximetrie-Modul ausgetauscht und Reklamatiert. Service tHb und tHb kal durchgeführt, Kalibriert und QC's durchgeführt. | SWO-198724 | No | Service dump created, oximetry module replaced and advertised. Service tHb and tHb kal performed, calibrated and QC's performed. | Service dump created, oximetry module replaced and advertised. Service tHb and tHb kal performed, calibrated and QC's performed. | Defect | HW | Oxi module | Oxi module
ABL80 FLEX CO-OX OSM | SC gespült, Fließwege gereinigt, LS Justage durch reboot ausgelöst, Funktion OK. | SC gespült, Fließwege gereinigt, LS Justage durch reboot ausgelöst, Funktion OK. | SWO-199107 | No | SC flushed, flow paths cleaned, LS adjustment triggered by reboot, function OK. | SC flushed, flow paths cleaned, LS adjustment triggered by reboot, function OK.
ABL80 FLEX BASIC | 更换902-972后修复(By Ran) | 更换902-972后修复(By Ran) | SWO-198579 | No | Repair after replacement 902-972 (By Ran) | Repair after replacement 902-972 (By Ran)
ABL90 | Customer attempted to replace the SC but the SC bay would not open. Although, the gears could be heard turning. Analyzer was just PM'd 6 8 2017. During the PM, the bay door was closed beyond the point that would allow the analyzer to normally open it. | Shut down the analyzer and manually opened the sensor bay. Verified that there were no obstructions. Opened and closed the sensor bay door in service mode with no errors. Analyzer is within acceptable operating specifications. | D2A9A3B5FE09D60E | Yes | Customer attempted to replace the SC but the SC bay would not open. Although, the gears could be heard turning. Analyzer was just PM'd 6 8 2017. During the PM, the bay door was closed beyond the point that would allow the analyzer to normally open it. | Shut down the analyzer and manually opened the sensor bay. Verified that there were no obstructions. Opened and closed the sensor bay door in service mode with no errors. Analyzer is within acceptable operating specifications.
ABL90 FLEX | 部品交換 | B | SWO-199318 | No | Parts replacement |  | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX | Gerät wieder zum Kunden gebracht und funktionstüchtig übergeben | Gerät wieder zum Kunden gebracht und funktionstüchtig übergeben | SWO-199155 | No | Device brought back to the customer and handed over functionally | Device brought back to the customer and handed over functionally | Defect | HW | Basic Unit | Sensor Interface PCB | Lac
ABL90 | Upon boot up analyzer would give system disk error, insert system disk message. | Replaced CF card. Installed Windows, 3.2 MR1, restored customers data & setup. Upgraded software to 3.3. | D2A96915FE06DEDD | No | Upon boot up analyzer would give system disk error, insert system disk message. | Replaced CF card. Installed Windows, 3.2 MR1, restored customers data & setup. Upgraded software to 3.3. | Defect | HW | DMS_PC | CF Card
ABL90 FLEX PLUS |  | B | D2A956B4FE0BF215 | Yes
ABL90 FLEX PLUS | Gerät und Scanner kontrolliert, Scanner erneuert | Gerät und Scanner kontrolliert, Scanner erneuert | SWO-199018 | No | Device and scanner checked, scanner renewed | Device and scanner checked, scanner renewed | Defect | HW | Basic Unit | Barcode reader, display
ABL90 FLEX | 连续更换2张测试卡后修复 | 连续更换2张测试卡后修复 | SWO-199201 | Yes | Repair after 2 consecutive test cards | Repair after 2 consecutive test cards
AQT90 FLEX | 更换903-364，重新定标正常。 by andrew | 更换903-364，重新定标正常。 by andrew | SWO-198968 | No | Replace the 903-364 and re-mark normally. by andrew | Replace the 903-364 and re-mark normally. by andrew | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-198，636-551后故障依旧，加借903-117，903-028各一个。 by andrew | 更换903-198，636-551后故障依旧，加借903-117，903-028各一个。 by andrew | SWO-199029 | No | After replacing 903-198, 636-551, the fault remains, plus 903-117, 903-028 each one. by andrew | After replacing 903-198, 636-551, the fault remains, plus 903-117, 903-028 each one. by andrew | Revisit required | HW | Analyzer Control Module | PCB, Main Board
ABL825 FLEX | 代替えモジュール設置でOxiモジュール引き上げ作業実施 | B | SWO-202120 | No | Oxi module pulling work implemented by installing an alternative module |  | Defect | HW | Oxi module | Oxi module
ABL835 FLEX | Software opdateret til ver. 6.17 Setup rettet | Software opdateret til ver. 6.17 Setup rettet | SWO-199381 | No | Software updated for Ver. 6.17 Setup fixed | Software updated for Ver. 6.17 Setup fixed | Defect | SW | SW | Blank Layout
ABL825 FLEX | Checked data logs. Found error code 0374: Inhomogeneous sample at pH BG module during QC runs. Some fluid leakage from front of analyzer. Customer already replaced inlet probe and inlet gasket. | Replaced tubing between modules, inlet gasket and pH seal ring. Cleaned pH measuring chamber and electrode. Ran zero current calibration in service. restart calibrations. | D2A989B9FE3F625D | No | Checked data logs. Found error code 0374: Inhomogeneous sample at pH BG module during QC runs. Some fluid leakage from front of analyzer. Customer already replaced inlet probe and inlet gasket. | Replaced tubing between modules, inlet gasket and pH seal ring. Cleaned pH measuring chamber and electrode. Ran zero current calibration in service. restart calibrations. | Adjustment | HW | Inlet module | Probe cal. | and Tubings
ABL90 FLEX | Einlasshalter gewechselt, dann Pressure Test i.O. Service Programme durchlaufen lassen i.O. Startup abgewartet - Gerät stand seit heute Nacht um 02:00Uhr Anmerkung: Bitte in Zukunft für ausreichend Material sorgen. Es müssen zu jeder Zeit mind. 2 LP, 2 SC und 2 Einlasshalter auf Vorrat gehalten werden. | B | SWO-199269 | No | Inlet holder changed, then Pressure Test i.O. Let service programs go through as a. Startup waited - device stood since tonight at 02:00o Note: Please ensure sufficient material in the future. At least 2 LP, 2 SC and 2 inlet holders must be kept in stock at all times. |  | Defect | HW | Inlet module | Inlet Gasket
AQT90 FLEX | 更换903-364后修复(By Ran) | 更换903-364后修复(By Ran) | SWO-198894 | No | Repair after replacement 903-364 (By Ran) | Repair after replacement 903-364 (By Ran) | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX | 更换903-364后修复(By Ran) | 更换903-364后修复(By Ran) | SWO-198895 | No | Repair after replacement 903-364 (By Ran) | Repair after replacement 903-364 (By Ran) | Defect | HW | Process Module | Optical Unit
AQT90 FLEX | Lichtschranken gereinigt, optische Einheit gereinigt, Neustart durchgeführt , Testmessungen und QC durchgeführt Wahrscheinlich technischer Defekt ,sporadisch immer diverse Fehlermeldungen Timeouts verschiedenster Module | B | SWO-199215 | No | Light barriers cleaned, optical unit cleaned, restart performed, test measurements and QC performed Probably technical defect ,sporadically always various error messages timeouts of various modules |  | Defect | HW | HW | Temperature
ABL90 FLEX |  | B | SWO-199236 | No |  |  | Defect | HW | Basic Unit | PCB, Wet section Control
AQT90 FLEX |  | 更换903-198，636-551后故障依旧，加借903-117，903-028各一个。 by andrew | SWO-199244 | No |  | After replacing 903-198, 636-551, the fault remains, plus 903-117, 903-028 each one. by andrew | Defect | HW | Analyzer Control Module | PCB, Main Board
ABL80 FLEX BASIC | センサーボード故障により現象発生。 センサーボード交換を実施しました。 | センサーボード故障により現象発生。 センサーボード交換を実施しました。 | SWO-199768 | No | A phenomenon occurs due to sensor board failure. The sensor board was replaced. | A phenomenon occurs due to sensor board failure. The sensor board was replaced.
ABL80 FLEX BASIC | done the zeroing and change the sensor board on standby purpose, need to replace the sensor board. | done the zeroing and change the sensor board on standby purpose, need to replace the sensor board. | SWO-201041 | No | done the zeroing and change the sensor board on standby purpose, need to replace the sensor board. | done the zeroing and change the sensor board on standby purpose, need to replace the sensor board.
ABL80 CO-OX | Oximetry failing on system cycle | Replaced Pinch valve and Hemolyzer assembly. Customer refused software upgrade at this time. | D2AD250E605D3542 | No | Oximetry failing on system cycle | Replaced Pinch valve and Hemolyzer assembly. Customer refused software upgrade at this time.
ABL90 FLEX | Reinigung Fließwege, Einlassdichtungshalter erneuert, Fluid Transport im Service durchgeführt, QC wiederholt - pCO2 zu hoch, pH in Level B zu tief, SC90 ersetzt, | Fehlerbehebung pH erfolgreich durch QC wiederholung, bez. pCO2 wird nochmals kalibriert mit anschl. QC sowie neue SC organisiert die dann ggf. ersetzt wird. Safe Pico vor Ort diese bzw Pico50 werden zukünftig bestellt, Bestelldauer vermutlich durch Hausinterne Vorgänge bedingt | SWO-199703 | No | Cleaning flow paths, inlet seal holder renewed, fluid transport carried out in service, QC repeated - pCO2 too high, pH in level B too deep, SC90 replaced, | Troubleshooting pH successfully by QC repetition, e.g. pCO2 is calibrated again with followed QC as well as new SC which will be replaced if necessary. Safe Pico on site this or Pico50 will be ordered in the future, order duration probably due to internal operations | Assisted Replacement | SC | SC | pCO2 | Rillibäk?
ABL90 FLEX | erneut kalibriert ohne Ergebnis, anschließend die SC herausgenommen, gespült und wieder eingesetzt,Kal i.O., QC i.O. (keine Messung die Ursache!) | erneut kalibriert ohne Ergebnis, anschließend die SC herausgenommen, gespült und wieder eingesetzt,Kal i.O., QC i.O. (keine Messung die Ursache!) | SWO-199744 | No | recalibrated without result, then the SC removed, rinsed and reinserted,Kal i.O., QC i.O. (no measurement the cause!) | recalibrated without result, then the SC removed, rinsed and reinserted,Kal i.O., QC i.O. (no measurement the cause!) | User error | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL825 FLEX |  | B | D2AD3D4F60199792 | No |  |  | Defect | HW | Inlet module | Inlet Gasket
ABL837 FLEX |  | B | D2B218C2384A9FB0 | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | Examine data logs. Check all inlet flow path. | Replaced parts from maintenance kit as needed for fault isolation. Replaced module and inlet tubings. Replaced inlet nipple at heater block input. Ran pump and liquid sensor calibrations with pass results. Observed operation. | D2AD55D56060C12B | No | Examine data logs. Check all inlet flow path. | Replaced parts from maintenance kit as needed for fault isolation. Replaced module and inlet tubings. Replaced inlet nipple at heater block input. Ran pump and liquid sensor calibrations with pass results. Observed operation. | Defect | HW | Inlet module | Nipple kit
ABL825 FLEX |  | B | D2AE337BD630AE50 | No |  |  | Defect | HW | Gas module | Gas Module
AQT90 FLEX |  | B | D2AD454D6006A5BA | Yes
ABL80 FLEX CO-OX OSM | 再現性なし。 ヘモライザーとピンチバルブ交換実施。 | 再現性なし。 ヘモライザーとピンチバルブ交換実施。 | SWO-199765 | No | No reproducibility. Hemolyzer and pinch valve replacement were implemented. | No reproducibility. Hemolyzer and pinch valve replacement were implemented.
ABL80 FLEX BASIC | Try to reinstall and reset SC and SP one by one. But not calibrated. SC replaced. PER returned for replacement. | Try to reinstall and reset SC and SP one by one. But not calibrated. SC replaced. PER returned for replacement. | SWO-200864 | No | Try to reinstall and reset SC and SP one by one. But not calibrated. SC replaced. PER returned for replacement. | Try to reinstall and reset SC and SP one by one. But not calibrated. SC replaced. PER returned for replacement.
ABL90 FLEX | Sample Mixer Unit erneuert. Einlaufverfahren durchgeführt. Alle Funktionen wieder o.k. | Sample Mixer Unit erneuert. Einlaufverfahren durchgeführt. Alle Funktionen wieder o.k. | SWO-200205 | No | Sample Mixer Unit renewed. run-in procedures. All functions again ok. | Sample Mixer Unit renewed. run-in procedures. All functions again ok. | Defect | HW | Basic Unit | Sampler Mixer
ABL90 | TECHLINE SUGGESTS POWER SUPPLY | REPLACED POWER SUPPLY | D2AE88CED613E5E2 | No | TECHLINE SUGGESTS POWER SUPPLY | REPLACED POWER SUPPLY | Defect | HW | Power Supply | Power Supply | Techline
ABL90 FLEX | Einlaßdichtung erneuert. Systemneustart. Kal, QC und Blutmessung iO. | Einlaßdichtung erneuert. Systemneustart. Kal, QC und Blutmessung iO. | SWO-200122 | No | Inlet seal renewed. System restart. Kal, QC and blood measurement iO. | Inlet seal renewed. System restart. Kal, QC and blood measurement iO. | Defect | HW | Inlet module | Inlet Gasket
ABL90 FLEX PLUS |  | B | D2AE8567D647CFD8 | Yes |  |  | Re classified RMED | Training
ABL90 FLEX PLUS | Na verwijderen inlethouder start instrument wel op met transportfoutemeldingen. Kan na plaatsen inlethouder geheel via servicemenu goed laten functioneren. Instrument opnieuw op laten starten, kalibreren en controles meten, OK. | Na verwijderen inlethouder start instrument wel op met transportfoutemeldingen. Kan na plaatsen inlethouder geheel via servicemenu goed laten functioneren. Instrument opnieuw op laten starten, kalibreren en controles meten, OK. | SWO-199960 | No | After removing the entry-holder start instrument with transport foulesthings. Can function properly after placing the holder in a service menu. Instrument to restart, calibrate and measure controls, OK. | After removing the entry-holder start instrument with transport foulesthings. Can function properly after placing the holder in a service menu. Instrument to restart, calibrate and measure controls, OK. | Adjustment | HW | Auto Inlet module | Probe cal. | LEH
ABL90 FLEX PLUS | SD-file heb gekopieerd voor RMED. | SD-file heb gekopieerd voor RMED. | SWO-200528 | Yes | SD-File copied for RMED. | SD-File copied for RMED. | Re classified RMED | To collect Auto inlet module
ABL90 FLEX PLUS | ホルダつきインレットガスケットの戻りが鈍い(キャピラリー測定時)。 ホルダ付インレットガスケット、インレットプローブ交換。 インレットモジュール分解点検実施。 別途ご依頼のレポートレイアウト変更。 | B | SWO-199933 | No | The return of the inlet gasket with the holder is dull (at capillary measurement). Inlet gasket with holder, inlet probe replacement. Inlet module disassembly inspection was carried out. The report layout change of the request separately. |  | Adjustment | HW | Auto Inlet module | Cap - Spring
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-201891 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | pCO2
AQT90 FLEX | 更换903-210、636-561和848-037各一个正常。 by andrew | 更换903-210、636-561和848-037各一个正常。 by andrew | SWO-199393 | No | Replace the 903-210, 636-561, and 848-037 each normal. by andrew | Replace the 903-210, 636-561, and 848-037 each normal. by andrew | Defect | HW | Process Module | PCB, Cup dryer | 17
ABL826 FLEX |  | B | D2AEEFF6D60C97A1 | Yes |  |  | Re classified RMED | Phone support
ABL825 FLEX | Reinigung, Kal.1, Servicekalibration, Gerät kalibriert und mit QCs überprüft, Ok | Reinigung, Kal.1, Servicekalibration, Gerät kalibriert und mit QCs überprüft, Ok | SWO-200149 | No | Cleaning, Kal.1, service calibration, device calibrated and checked with QCs, Ok | Cleaning, Kal.1, service calibration, device calibrated and checked with QCs, Ok | Adjustment | HW | Oxi module | Oxi module
ABL837 FLEX |  | B | D2AEEEBAD644979F | No |  |  | Revisit required | Electrode | Electrode | pH
AQT90 FLEX | Main Board erneuert , AQT neu kalibriert , keine QC durch geführt da noch diverse Parameter kalibriert werden müßen, Kunde mißt nach Durchführung die QC Frau Gansau informiert | Main Board erneuert , AQT neu kalibriert , keine QC durch geführt da noch diverse Parameter kalibriert werden müßen, Kunde mißt nach Durchführung die QC Frau Gansau informiert | SWO-200146 | No | Main board renewed, AQT recalibrated, no QC performed as various parameters still need to be calibrated, customer measures after carrying out the QC Informed Ms Gansau | Main board renewed, AQT recalibrated, no QC performed as various parameters still need to be calibrated, customer measures after carrying out the QC Informed Ms Gansau | Defect | HW | Analyzer Control Module | PCB, Main Board | 3264
TCM400 |  | B | D2AE8B5DD6181A28 | No
ABL80 FLEX BASIC | 更换902-972后修复 by andrew | 更换902-972后修复 by andrew | SWO-199954 | No | Repair by andrew after replacement 902-972 | Repair by andrew after replacement 902-972
ABL90 | Confirm when on site. | Replaced the sensor interface PCB. | D2AFE38E4C3F4A90 | No | Confirm when on site. | Replaced the sensor interface PCB. | Defect | HW | Basic Unit | Sensor Interface PCB
ABL90 FLEX | Einlassdichtung aus JW Satz eingebaut. Im Servicemenü Programme mehrfach wiederholt bis alle Fehler beseitigt waren. | Einlassdichtung aus JW Satz eingebaut. Im Servicemenü Programme mehrfach wiederholt bis alle Fehler beseitigt waren. | SWO-200584 | No | Inlet seal built from JW set. Repeated several times in the service menu programs until all errors have been fixed. | Inlet seal built from JW set. Repeated several times in the service menu programs until all errors have been fixed. | Defect | HW | Inlet module | Inlet Gasket
ABL90 | the analyzer showed no life | I checked the wall power and it was good, i disconnected the battery and still now power. I need a power supply | D2AF78374C3995E4 | No | the analyzer showed no life | I checked the wall power and it was good, i disconnected the battery and still now power. I need a power supply | Revisit required | HW | Power Supply | Power Supply
ABL90 FLEX PLUS | Dateien eingespielt, Gerät neu gestartet | Dateien eingespielt, Gerät neu gestartet | SWO-200482 | No | Files played in, device restarted | Files played in, device restarted | Setup | SW | RMED | Parameter setup
ABL90 FLEX | 更换903-183备件后修复 by andrew | 更换903-183备件后修复 by andrew | SWO-199814 | No | Repair by andrew after replacement of 903-183 spare parts | Repair by andrew after replacement of 903-183 spare parts | Defect | HW | Basic Unit | Sensor Interface PCB | 1326
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-200190 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Reference
ABL827 FLEX |  | B | D2AFD67A4C23E4F5 | Yes
ABL827 FLEX |  | B | D2AFD8D34C02E4FA | Yes
ABL825 FLEX | Undichte Messkammer gereinigt , Airkanal und Wastekanal. gespült , Ref.Elektrode erneuert. | Undichte Messkammer gereinigt , Airkanal und Wastekanal. gespült , Ref.Elektrode erneuert. | SWO-200526 | No | Leaky measuring chamber cleaned, air channel and waste channel. rinsed, Ref.electrode renewed. | Leaky measuring chamber cleaned, air channel and waste channel. rinsed, Ref.electrode renewed. | Assisted Replacement | Electrode | Electrode | Reference
ABL827 FLEX | Sensitivity high above upper limit. Problem occurred while unit was being validated for method comparisons. | Checked fluid flow and valves. Seem OK. Replaced El Med module pump tubing. | D2AFAADB4C3C51E0 | No | Sensitivity high above upper limit. Problem occurred while unit was being validated for method comparisons. | Checked fluid flow and valves. Seem OK. Replaced El Med module pump tubing. | Defect | HW | El/Met Module | Pump tubes
TCM400-2 |  | B | D2AF84DE4C048EED | No
ABL80 FLEX BASIC | 更换903-334及测试卡后修复 by andrew | 更换903-334及测试卡后修复 by andrew | SWO-199815 | No | Repair by andrew after replacing 903-334 and test card | Repair by andrew after replacing 903-334 and test card
ABL80 FLEX BASIC | クーリングファン交換。 | B | SWO-204314 | No | Cooling fan replacement.
ABL80 FLEX BASIC | 更换903-334后修复 by andrew | 更换903-334后修复 by andrew | SWO-200328 | No | Repair by andrew after replacing 903-334 | Repair by andrew after replacing 903-334
ABL90 |  | B | D2B14070C218C217 | No |  |  | Adjustment | HW | Basic Unit | SC Door
ABL90 FLEX | Austausch div Verschleißteile Ventil überprüft Überprüfung Fließwege | Austausch div Verschleißteile Ventil überprüft Überprüfung Fließwege | SWO-200822 | No | Replacement div consumables Valve checked Review flow paths | Replacement div consumables Valve checked Review flow paths | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | New analyzer
ABL90 FLEX | 更换924-816后修复 by jesson | 更换924-816后修复 by jesson | SWO-200998 | No | Repair by after replacement 924-816 jesson | Repair by after replacement 924-816 jesson | Defect | HW | Basic Unit | Inlet Gasket
ABL90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-200810 | No | Repair after replacing test card | Repair after replacing test card | Assisted Replacement | SC | SC | Chip coding
ABL800 BASIC | 更换903-125备件后修复 by andrew | 更换903-125备件后修复 by andrew | SWO-200327 | No | Repair by andrew after replacement of 903-125 spare parts | Repair by andrew after replacement of 903-125 spare parts | Defect | HW | Oxi module | PCB, Oxi module | 537
ABL825 FLEX |  | B | D2B2236F38509FB8 | No |  |  | Defect | SW | Windows | Windows | Error code 2
ABL800 BASIC | Clean the inlet gasket properly,clean the waste pump tube,flush the inlet probe,run the inlet cali,liquid sensor cal & pump cal,after successful calibration run few sample,got the satisfactory result. | Clean the inlet gasket properly,clean the waste pump tube,flush the inlet probe,run the inlet cali,liquid sensor cal & pump cal,after successful calibration run few sample,got the satisfactory result. | SWO-201264 | No | Clean the inlet gasket properly,clean the waste pump tube,flush the inlet probe,run the inlet cali,liquid sensor cal & pump cal,after successful calibration run few sample,got the satisfactory result. | Clean the inlet gasket properly,clean the waste pump tube,flush the inlet probe,run the inlet cali,liquid sensor cal & pump cal,after successful calibration run few sample,got the satisfactory result. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL827 FLEX | Checked all data logs for QC issues and found CA falls below the customer set ranges but are within the Radiometer packet insert ranges. Cleaned Reference and Ca electrode connector. | Upgraded main software to version 6.17 as per FAN. Run 3 l3v3ls of QC. All tests are successful. | D2B0F225C20395D9 | Yes | Checked all data logs for QC issues and found CA falls below the customer set ranges but are within the Radiometer packet insert ranges. Cleaned Reference and Ca electrode connector. | Upgraded main software to version 6.17 as per FAN. Run 3 l3v3ls of QC. All tests are successful. | Re classified RMED | Within RMED Specs
ABL825 FLEX | looked through the logs and error started on June 8 intermittently 8 to 10 times a day | ran a complete oxy test and all test are passing; performed the software upgrade will document on a Separate SWO for tracking | D2B10E0DC21117E4 | No | looked through the logs and error started on June 8 intermittently 8 to 10 times a day | ran a complete oxy test and all test are passing; performed the software upgrade will document on a Separate SWO for tracking | No fault found | HW | Oxi module | Thermistor | 27
ABL825 FLEX |  | B | D2B253F0380C154C | No |  |  | Defect | HW | FlexQ Module | FlexQ Cover
AQT90 FLEX | Gerät geprüft, Verbrauchsmaterial (LP und Kassetten) ersetzt, Funktionskontrolle ok | Gerät geprüft, Verbrauchsmaterial (LP und Kassetten) ersetzt, Funktionskontrolle ok | SWO-200803 | No | Device checked, consumables (LP and cassettes) replaced, function control ok | Device checked, consumables (LP and cassettes) replaced, function control ok | Defect | Consumables | Test Kit | Cartridge | D-Dimer
AQT90 FLEX |  | B | D2B0DA42C20666D6 | Yes
AQT90 FLEX |  | B | D2B0F8CDC2152F39 | No |  |  | Defect | HW | Process Module | Optical Unit
AQT90 FLEX |  | B | D2B21C1C38257601 | Yes |  |  | Re classified RMED | See next upSWO | Test Kit | TnT | Calibration
AQT90 FLEX |  | B | D2B0F8FBC24152D2 | No |  |  | Defect | HW | Process Module | Optical Unit
ABL90 FLEX | Fehlersuche : Druckertreiber überprüft, sämtliche Einstellungen weg. Treiber Einstellungen korrigiert, 1x Kopie Treiber gelöscht. Patres Dateien im Odin und Language Ordner kopiert. Einstellungen mit Commit Befehl gespeichert. Startphase abgewartet, Testdrucke gemacht. | B | SWO-201200 | No | Troubleshooting: Printer driver checks, all settings gone. Driver settings corrected, 1x copy driver deleted. Patres files copied to the Odin and Language folder. Settings saved with commit command. start-up phase waited, test prints made. |  | Setup | SW | External Printer | External Printer
ABL90 | Verified CPU and wireless usb adapter were malfunctioning. | Replaced CPU and wireless usb adapter. Verified instrument operation. | D2B271B73822CCBC | No | Verified CPU and wireless usb adapter were malfunctioning. | Replaced CPU and wireless usb adapter. Verified instrument operation. | Defect | HW | DMS_PC | CPU
ABL90 FLEX PLUS | Gerät demontiert, Schlauchsystem gespült, SC 90 neu installiert, Funktionskontrolle ok | Gerät demontiert, Ventileinheit demontiert und gereinigt, Ventilschlauch ersetzt, SC 90 erneuert, Funktionskontrolle ok | SWO-201569 | No | Device dismantled, hose system flushed, SC 90 reinstalled, function control ok | Device dismantled, valve unit disassembled and cleaned, valve hose replaced, SC 90 renewed, function control ok | Cleaned | Fluidic pathway | Oxi module | Hemolyzer
ABL90 | KIT ANUAL REVISIÓN PREVENTIVA | KIT ANUAL REVISIÓN PREVENTIVA | D2B220873853DCC6 | No | ANNUAL KIT PREVENTIVE REVIEW | ANNUAL KIT PREVENTIVE REVIEW | Defect | HW | PM Kit | PM Kit
AQT90 FLEX | 更换902-924后修复 by jesson | 更换902-924后修复 by jesson | SWO-200701 | No | Repair by jesson after replacement 902-924 | Repair by jesson after replacement 902-924 | Defect | HW | Analyzer Control Module | Barcode reader, display
AQT90 FLEX | 更换903-198后修复 by andrew | 更换903-198后修复 by andrew | SWO-200727 | No | Repair by andrew after replacement 903-198 | Repair by andrew after replacement 903-198 | Defect | HW | Analyzer Control Module | PCB, Main Board | 5766
TCM5 FLEX |  | B | D2B21E27380DAE58 | No
ABL825 FLEX | Aufgef. Verschleißteile ersetzt, dazu Analysator runtergefahren, Flüssigkeit war in Sensor gelaufen, Analysator wieder hochgefahren , LS-Justage im Serviceprogramm i.O., Nachfüllprogramm, Lac-Elektrode nach Radiometer Verfahren gereinigt und mit neuer Lot (3526) membraniertKal, QC | Aufgef. Verschleißteile ersetzt, dazu Analysator runtergefahren, Flüssigkeit war in Sensor gelaufen, Analysator wieder hochgefahren , LS-Justage im Serviceprogramm i.O., Nachfüllprogramm, Lac-Elektrode nach Radiometer Verfahren gereinigt und mit neuer Lot (3526) membraniertKal, QC | SWO-201208 | No | Aufgef. Replacement of wear parts, with analyser shut down, liquid had run into sensor, analyzer restarted, LS-Adjustment in the service program i.O., refill program, Lac electrode cleaned according to radiometer procedure and membraned with new solder (3526)Kal Qc | Aufgef. Replacement of wear parts, with analyser shut down, liquid had run into sensor, analyzer restarted, LS-Adjustment in the service program i.O., refill program, Lac electrode cleaned according to radiometer procedure and membraned with new solder (3526)Kal Qc | Adjustment | HW | El/Met Module | LS
ABL825 FLEX | CLEANED AND DECONTAMINATED HEMOLYZER WITH NO RESULTS. | NONE | D2B2C7873855B0C6 | No | CLEANED AND DECONTAMINATED HEMOLYZER WITH NO RESULTS. | NONE | Cleaned | HW | Oxi module | Oxi module | Revisit?
ABL825 FLEX |  | B | D2B25454384D1550 | No |  |  | Defect | HW | FlexQ Module | FlexQ Cover | LEH Email
ABL90 FLEX PLUS | Gerät demontiert, Ventileinheit demontiert und gereinigt, Ventilschlauch ersetzt, SC 90 erneuert, Funktionskontrolle ok | Gerät demontiert, Ventileinheit demontiert und gereinigt, Ventilschlauch ersetzt, SC 90 erneuert, Funktionskontrolle ok | SWO-201568 | No | Device dismantled, valve unit disassembled and cleaned, valve hose replaced, SC 90 renewed, function control ok | Device dismantled, valve unit disassembled and cleaned, valve hose replaced, SC 90 renewed, function control ok | Defect | HW | PM Kit | PM Kit
ABL90 FLEX | 更换试剂包后修复 | 更换试剂包后修复 | SWO-201893 | No | Repair after replacement of reagent kit | Repair after replacement of reagent kit | Assisted Replacement | SP | SP | FT
AQT90 FLEX | 更换新机器后修复 by andrew | 更换新机器后修复 by andrew | SWO-200908 | No | Fix by andrew after replacing a new machine | Fix by andrew after replacing a new machine | Defect | HW | Process Module | Optical Unit | 159
AQT90 FLEX | 更换新机器后修复 by andrew | 更换新机器后修复 by andrew | SWO-200921 | Yes | Fix by andrew after replacing a new machine | Fix by andrew after replacing a new machine | Re classified RMED | Double entry
AQT90 FLEX |  | B | SWO-201259 | No |  |  | Assisted Replacement | Consumables | Test Kit | BNP Calibration
ABL825 FLEX | 更换902-813后修复 by andrew | 更换902-813后修复 by andrew | SWO-200990 | No | Repair by andrew after replacement 902-813 | Repair by andrew after replacement 902-813 | Defect | HW | Inlet module | PCB, Flap Detector
ABL825 FLEX | Confirmed reported issue with customer. | Suggested customer replace inlet gasket and waste pump tubes, problem resolved. | D2B31562AE407867 | Yes | Confirmed reported issue with customer. | Suggested customer replace inlet gasket and waste pump tubes, problem resolved. | Re classified RMED | Phone support
ABL90 | Confirmed issue. Someone prior to calling for assistance had begun the replacement process but stopped, with cassette bay door open. | Advised customer to continue with replacement process. Instructed tech to save credit form printout and provide it with SC to the manage for submission of replacement. | D2B483812446F077 | Yes | Confirmed issue. Someone prior to calling for assistance had begun the replacement process but stopped, with cassette bay door open. | Advised customer to continue with replacement process. Instructed tech to save credit form printout and provide it with SC to the manage for submission of replacement. | Assisted Replacement | SC | SC | Reference
ABL90 FLEX | 更换测试卡,试剂包正常。 | 更换测试卡,试剂包正常。 | SWO-201334 | No | Replace the test card and the reagent package is OK. | Replace the test card and the reagent package is OK. | Assisted Replacement | SC | SC | Glu
AQT90 FLEX | 重新定标正常。(by tim) | 重新定标正常。(by tim) | SWO-201265 | No | Recalibration is normal. (by tim) | Recalibration is normal. (by tim) | Assisted Replacement | Consumables | Test Kit | Myo Calibration
AQT90 FLEX | 重新定标后修复 by andrew | 重新定标后修复 by andrew | SWO-200932 | No | Fix by andrew after re-calibration | Fix by andrew after re-calibration | Assisted Replacement | Consumables | Test Kit | BNP Calibration
ABL80 FLEX Basic version |  | B | D2B604DB9A4A8A0E | No
ABL80 | Checked the data logs. | Replaced the Sensor board. | D2B5E0859A1075FA | No | Checked the data logs. | Replaced the Sensor board.
ABL90 FLEX | Aktivitäten durchgesehen - Problem heute Nacht war, dass ein Ableitstrom im LP erkannt wurde, danach wurde er gewechselt Ableitstrom aktuell überprüft war bei -37mV und damit i.O. Oxi Schlauch, Einkassdichtungen überprüft i.O. Serviceprogramme durchlaufen lassen i.O. | Aktivitäten durchgesehen - Problem heute Nacht war, dass ein Ableitstrom im LP erkannt wurde, danach wurde er gewechselt Ableitstrom aktuell überprüft war bei -37mV und damit i.O. Oxi Schlauch, Einkassdichtungen überprüft i.O. Serviceprogramme durchlaufen lassen i.O. | SWO-201565 | No | Activities reviewed - Problem tonight was that a leakage stream was detected in the LP, after which it was changed Discharge current was currently checked at -37mV and thus i.O. Oxi hose, cash seals checked in i.O. Let service programs go through as a. | Activities reviewed - Problem tonight was that a leakage stream was detected in the LP, after which it was changed Discharge current was currently checked at -37mV and thus i.O. Oxi hose, cash seals checked in i.O. Let service programs go through as a. | Assisted Replacement | SP | SP | Leak current
ABL90 FLEX PLUS | Gerät überprüft, Touchsreen erneuert | Gerät überprüft, Touchsreen erneuert | SWO-202066 | No | Device checked, touchsreen renewed | Device checked, touchsreen renewed | Defect | HW | Basic Unit | Touch screen
ABL90 FLEX PLUS | Allgemeine Überprüfung - Erneuerung Sensorkassette - Funktionsüberprüfung etc. | Allgemeine Überprüfung - Erneuerung Sensorkassette - Funktionsüberprüfung etc. | SWO-201766 | Yes | General inspection - Renewal sensor cassette - Function check etc. | General inspection - Renewal sensor cassette - Function check etc. | Re classified RMED | Logistical upSWO
ABL90 FLEX PLUS | Replaced Oxi pinch valve tube. No actual fault found. | Replaced Oxi pinch valve tube. No actual fault found. | SWO-201802 | No | Replaced Oxi pinch valve tube. No current fault found. | Replaced Oxi pinch valve tube. No current fault found. | Adjustment | HW | Oxi module | tHb Cal
ABL835 FLEX | measuring chamber değişti | measuring chamber değişti | SWO-201665 | Yes | measuring chamber değişti | measuring chamber değişti
ABL825 FLEX | error 270 in the manual instructs to replace hemolyzer | replaced hemolyzer and optical fiber cable; cable was crimped; after replacing parts ran complete oxy test and neon lamp failed erro code 258 neon lamp too low neon lamp was at 45% should be > 50 %; replaced the lamp unit | D2B5FF529A5C2FCB | No | error 270 in the manual instructs to replace hemolyzer | replaced hemolyzer and optical fiber cable; cable was crimped; after replacing parts ran complete oxy test and neon lamp failed erro code 258 neon lamp too low neon lamp was at 45% should be > 50 %; replaced the lamp unit | Defect | HW | Oxi module | Hemolyzer | 27
ABL835 FLEX | Demonøgler er nu slået fra og fejl 0600 Demoperiode udløber snart skulle gerne være væk i morgen. | Demonøgler er nu slået fra og fejl 0600 Demoperiode udløber snart skulle gerne være væk i morgen. | SWO-201690 | Yes | Demo keys are now turned off and Error 0600 Demo period expires soon should be like to be gone tomorrow. | Demo keys are now turned off and Error 0600 Demo period expires soon should be like to be gone tomorrow. | Re classified RMED
ABL825 FLEX | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer. | Replaced hemolyzer, adjusted lamp, did complete oximetry test. Did THB in service and as user and as sample. | D2B64A459A50D7A2 | No | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer. | Replaced hemolyzer, adjusted lamp, did complete oximetry test. Did THB in service and as user and as sample. | Defect | HW | Oxi module | Hemolyzer | Spectrum mismatch
ABL825 FLEX | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer. | Replaced hemolyzer, adjusted lamp, did complete oximetry test. Did THB in service and as user and as sample. | D2B739FA1049ED55 | No | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer. | Replaced hemolyzer, adjusted lamp, did complete oximetry test. Did THB in service and as user and as sample. | Revisit required | HW | Oxi module | Hemolyzer
AQT90 FLEX | Comprobar bombas y válvulas rellenar los tubos de paso del reactivo, calibrar y probar con un blanco. Revisión del equipo y puesta en marcha despues de estar un largo tiempo parado. | B | D2B5EF739A08F999 | No | Check pumps and valves fill the reagent passage tubes, calibrate and test with a white. Equipment review and commissioning after being stopped for a long time. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway | 3414
AQT90 FLEX | İĞNE DEĞİŞTİRİLDİ. | İĞNE DEĞİŞTİRİLDİ. | SWO-201919 | Yes | NEEDLE CHANGED. | NEEDLE CHANGED.
AQT90 FLEX |  | B | D2B651039A5E9A99 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | 3412
AQT90 FLEX |  | B | D2B62CE79A149D7E | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | 3412
ABL80 FLEX BASIC | 连续更换2张测试卡后修复 | 连续更换2张测试卡后修复 | SWO-202075 | Yes | Repair after 2 consecutive test cards | Repair after 2 consecutive test cards
ABL80 FLEX BASIC | Instrument was rebooting, check the CF care it was OK, need to replace with new CF card and check awaiting for Spares. | Instrument was rebooting, check the CF care it was OK, need to replace with new CF card and check awaiting for Spares. | SWO-202541 | No | Instrument was rebooting, check the CF care it was OK, need to replace with new CF card and check awaiting for Spares. | Instrument was rebooting, check the CF care it was OK, need to replace with new CF card and check awaiting for Spares.
ABL80 FLEX BASIC |  | B | SWO-201880 | No
ABL80 FLEX BASIC | Sharing. Check the instrument found blood got clotted inside the Sensor cassette, flushed the SC with water reinstalled the SC found SC was working instrument is working satisfactorily. | Sharing. Check the instrument found blood got clotted inside the Sensor cassette, flushed the SC with water reinstalled the SC found SC was working instrument is working satisfactorily. | SWO-202561 | No | Sharing. Check the instrument found blood got clotted inside the Sensor cassette, flushed the SC with water reinstalled the SC found SC was working instrument is working satisfactorily. | Sharing. Check the instrument found blood got clotted inside the Sensor cassette, flushed the SC with water reinstalled the SC found SC was working instrument is working satisfactorily.
ABL80 FLEX Basic version |  | B | D2B789DD104F8A18 | No
ABL80 FLEX BASIC | Sensor Board Replaced. SC and SP installed. Zeroing calibration done. QC check and found not passing. SP replaced but same problem found. SC replaced and found some parameters not calibrating. Need to replace both SC and SP. Analyser Serviced and Calibrated properly. Handed over in a fine working condition. | Sensor Board Replaced. SC and SP installed. Zeroing calibration done. QC check and found not passing. SP replaced but same problem found. SC replaced and found some parameters not calibrating. Need to replace both SC and SP. Analyser Serviced and Calibrated properly. Handed over in a fine working condition. | SWO-203337 | No | Sensor Board Replaced. SC and SP installed. Zeroing calibration done. QC check and found not passing. SP replaced but same problem found. SC replaced and found some parameters not calibrating. Need to replace both SC and SP. Analyser Serviced and Calibrated properly. Handed over in a fine working condition. | Sensor Board Replaced. SC and SP installed. Zeroing calibration done. QC check and found not passing. SP replaced but same problem found. SC replaced and found some parameters not calibrating. Need to replace both SC and SP. Analyser Serviced and Calibrated properly. Handed over in a fine working condition.
ABL90 FLEX | Fließwege gespült, neues LP eingelegt, Gerät kalibriert und QC durchgeführt, | Fließwege gespült, neues LP eingelegt, Gerät kalibriert und QC durchgeführt, | SWO-202226 | No | Flow paths rinsed, new LP inserted, device calibrated and QC performed, | Flow paths rinsed, new LP inserted, device calibrated and QC performed, | Assisted Replacement | SP | SP | pO2 QC C High
ABL90 | pCO2 calibration was in but QC results were out of range. Lactate calibration sensitivity was low and controls out of range. Customer has changed several SC's & SP's but each time the failed parameters change for different reasons. | Changed SP & SC and instrument came up but had some inhomogeneous cal solution 3. replaced inlet probe and gasket. Calibrated and ran QC. Instrument operates normally. | D2B6FF85105CA17F | No | pCO2 calibration was in but QC results were out of range. Lactate calibration sensitivity was low and controls out of range. Customer has changed several SC's & SP's but each time the failed parameters change for different reasons. | Changed SP & SC and instrument came up but had some inhomogeneous cal solution 3. replaced inlet probe and gasket. Calibrated and ran QC. Instrument operates normally. | Assisted Replacement | SC | SC | Lac | and pCO2
ABL90 FLEX | 更换903-183后修复(by tim) | 更换903-183后修复(by tim) | SWO-201419 | No | Repair after replacement 903-183 (by tim) | Repair after replacement 903-183 (by tim) | Defect | HW | Basic Unit | Sensor Interface PCB
ABL826 FLEX |  | B | D2BB02517212A0EE | No |  |  | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | checked replacement logs, Gas1 changed on 5 29, 6 5, 6 12, 6 19, | checked for leaks, entered Gas program and recalibrated Gas valves | D2B7BD3610248F03 | No | checked replacement logs, Gas1 changed on 5 29, 6 5, 6 12, 6 19, | checked for leaks, entered Gas program and recalibrated Gas valves | Revisit required | HW | Gas module | Gas valve
ABL835 FLEX | Lampe justeret tHb kal. udført | Lampe justeret tHb kal. udført | SWO-202229 | No | Lampe justeret tHb kal. udført | Lampe justeret tHb kal. udført | Adjustment | HW | Oxi module | Lamp unit | Halogen
ABL825 FLEX | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer. | Replaced hemolyzer, adjusted lamp, did complete oximetry test. Did THB in service and as user and as sample | D2B73D45100549DC | No | All Cals and QC passed but Cal Ver level III would not pass, would get oxi spectrum mismatch error code 0581, 0688, 0094. Suspect bad Hemolyzer. | Replaced hemolyzer, adjusted lamp, did complete oximetry test. Did THB in service and as user and as sample | Defect | HW | Oxi module | Hemolyzer
ABL90 FLEX | Einstellung Timeserver deaktiviert | Nach Haken bei Daylight saving rein raus stimmt Uhrzeit mit Update Now nicht mehr, Haken gestetzt, Timer Sever Test Uhrzeit okay | SWO-202413 | No | Setting Timeserver disabled | After hooking at Daylight saving in out time is no longer correct with Update Now, hook stung, Timer Sever test time okay | Setup | SW | Time setting | Time setting
ABL90 FLEX PLUS | Installed new UPS and removed the faulty UPS | Installed new UPS and removed the faulty UPS | SWO-204074 | No | Installed new UPS and removed the faulty UPS | Installed new UPS and removed the faulty UPS | No fault found | HW | UPS | UPS
ABL90 FLEX PLUS | Sensorcassette udskiftet, startet op og slutafprøvet. | Sensorcassette udskiftet, startet op og slutafprøvet. | SWO-202621 | No | Sensorcassette replaced, booted and finished tested. | Sensorcassette replaced, booted and finished tested. | Assisted Replacement | SC | SC | Lac
ABL90 FLEX PLUS | Startphase QC und Kalibration. | Startphase QC und Kalibration. | SWO-202481 | No | Start phase QC and calibration. | Start phase QC and calibration. | Assisted Replacement | SP | SP | pO2
ABL90 FLEX | 更换测试卡 | 更换测试卡 | SWO-203542 | No | Replace the test card | Replace the test card | Assisted Replacement | SC | SC | Lac
AQT90 FLEX |  | B | SWO-202339 | No |  |  | Defect | HW | Storage Module | Barcode reader, Storage wheel | br1
AQT90 FLEX | 更换903-329备件一个(by tim) | 更换903-329备件一个(by tim) | SWO-201918 | No | Replace 903-329 spare parts one (by tim) | Replace 903-329 spare parts one (by tim) | Defect | HW | Aspiration Module | Needle Wash Unit | 3412
AQT90 FLEX | 取出测试卡后清理修复(by tim) | 取出测试卡后清理修复(by tim) | SWO-202449 | No | Clean up the repair after removing the test card (by tim) | Clean up the repair after removing the test card (by tim) | Defect | Consumables | Test Kit | Cartridge | Broken
AQT90 FLEX | 更换903-364后修复 | 更换903-364后修复 | SWO-202445 | No | Repair after replacement 903-364 | Repair after replacement 903-364 | Defect | HW | Process Module | Optical Unit | 124
ABL825 FLEX | Drift in the range of 20 out. Tried complete oxi test and lamp unit adjust. Problem goes away and comes back. | Replaced oxi module PCB. Performed all necessary calibrations.Performed multiple 1 and 2pt calibrations. Drift in normal range. Problem seems to have gone away. Will have customer monitor over next week or so since this was an intermittent problem. | D2B8645E8644B4D3 | No | Drift in the range of 20 out. Tried complete oxi test and lamp unit adjust. Problem goes away and comes back. | Replaced oxi module PCB. Performed all necessary calibrations. Performed multiple 1 and 2pt calibrations. Drift in normal range. Problem seems to have gone away. Will have customer monitor over next week or so since this was an intermittent problem. | Defect | HW | Oxi module | PCB, Oxi module
AQT90 FLEX | Gerät überprüft, keine weiteren Fehlermeldungen vorhanden, Daten gelöscht, Filter gereinigt, Probenmessungen o.k. | Gerät überprüft, keine weiteren Fehlermeldungen vorhanden, Daten gelöscht, Filter gereinigt, Probenmessungen o.k. | SWO-202918 | No | Device checked, no further error messages available, data deleted, filter cleaned, sample measurements ok. | Device checked, no further error messages available, data deleted, filter cleaned, sample measurements ok. | No fault found | No fault found | No fault found | No fault found | 167
AQT90 FLEX | Mangelhaft justierten Barcodeleser ausgerichtet. | Mangelhaft justierten Barcodeleser ausgerichtet. | SWO-202678 | No | Poorly adjusted barcode reader aligned. | Poorly adjusted barcode reader aligned. | Adjustment | HW | Storage Module | Barcode reader, Storage wheel
ABL80 FLEX | 备件接错，现需要重新借备件924-738，故此单先关单(by tim) | 更换924-738一个正常。(by tim) | SWO-202853 | No | Spare parts are wronganded and need to be re-borrowed for spare parts 924-738, so the single first clearance (by tim) | Replace 924-738 one normal. (by tim)
ABL90 FLEX | Startphase QC und Kalibration. | Startphase QC und Kalibration. | SWO-203031 | No | Start phase QC and calibration. | Start phase QC and calibration. | Assisted Replacement | SP | SP | Leakage
ABL90 FLEX PLUS | Sensor cassette vervangen. | Sensor cassette vervangen. | SWO-202937 | No | Replace the Sensor cassette. | Replace the Sensor cassette. | Assisted Replacement | SC | SC | Clot
ABL90 FLEX PLUS | Service visit required. Passed toColin Maunder. | Service visit required. Passed toColin Maunder. | SWO-202836 | Yes | Service visit required. Passed toColin Maunder. | Service visit required. Passed toColin Maunder. | Re classified RMED | Phone support
ABL90 FLEX PLUS | Cleaned gasket and flushed cassette. Cleaned sensor cassette contact pins. Removed solution pack and checked it was activated correctly. Clicks could be heard when pressure was applied to the top. Replaced sensor cassette. Analyser is now passing calibrations. I will call POCT 23 06 am to see if QC is now passing. Both cassettes have been removed and will claimed for. | Service visit required. Passed toColin Maunder. | SWO-203438 | No | Cleaned gasket and flushed cassette. Cleaned sensor cassette contact pins. Removed solution pack and checked it was activated correctly. Clicks could be heard when pressure was applied to the top. Replaced sensor cassette. Analyser is now passing calibrations. I will call POCT 23 06 am to see if QC is now passing. Both cassettes have been removed and will claimed for. | Service visit required. Passed toColin Maunder. | Assisted Replacement | SC | SC | Lac
AQT90 FLEX | 更换测试卡后修复 | 更换测试卡后修复 | SWO-202868 | No | Repair after replacing test card | Repair after replacing test card | Defect | Consumables | Test Kit | Cartridge | Broken
AQT90 FLEX | 更换903-329后修复(by tim) | 更换903-329后修复(by tim) | SWO-202451 | No | Repair after replacement 903-329 (by tim) | Repair after replacement 903-329 (by tim) | Defect | HW | Aspiration Module | Needle Wash Unit
ABL837 FLEX |  | B | D2B9A65CFC08F92C | No |  |  | Adjustment | HW | DMS_PC | Touch screen
ABL825 FLEX | Remotely connected to analyzer and found that the Roma.ini file and AU board stop communicating or not responding. Replaced the roma.ini. and still happening. Reloaded SW 6.16v and roma.ini up and running but customer went to run sample and touch screen is not working results not crossing over. Arrived onsite and found that the touch screen cable was loose and not properly attached. Found that the USB port was broken and TCP IP was not responding. Results not crossing over to cerner. | Replaced the CPU and booted analyzer without errors. Tightened down calbe for touch screen. Cofigured the LIS HIS settings to communicate with cerner. Pinged the HIS and is responding. | D2B9F7DCFC01E971 | No | Remotely connected to analyzer and found that the Roma.ini file and AU board stop communicating or not responding. Replaced the roma.ini. and still happening. Reloaded SW 6.16v and roma.ini up and running but customer went to run sample and touch screen is not working results not crossing over. Arrived onsite and found that the touch screen cable was loose and not properly attached. Found that the USB port was broken and TCP IP was not responding. Results not crossing over to cerner. | Replaced the CPU and booted analyzer without errors. Tightened down calbe for touch screen. Cofigured the LIS HIS settings to communicate with cerner. Pinged the HIS and is responding. | Defect | HW | DMS_PC | CPU
ABL835 FLEX | Calibrations all ok. No Ph errors reported. Nothing physical showing. | Replaced PH BG air and rinse valves thinking they may be randomly failing. Performed software upgrade to version 6.17. Performed all necessary calibrations. | D2B9C12AFC5EF14E | No | Calibrations all ok. No Ph errors reported. Nothing physical showing. | Replaced PH BG air and rinse valves thinking they may be randomly failing. Performed software upgrade to version 6.17. Performed all necessary calibrations. | Defect | HW | pH/BG Module | Valve, air
ABL835 FLEX | OXI modul skiftet ud med et fra de nye ABL'ere Hjemtaget til reparation | OXI modul skiftet ud med et fra de nye ABL'ere Hjemtaget til reparation | SWO-203384 | No | OXI module replaced with one from the new Abls Home taken for repair | OXI module replaced with one from the new Abls Home taken for repair | Defect | HW | Oxi module | Oxi module | Neon lamp
ABL90 FLEX | Replaced touch screen. | Replaced touch screen. | SWO-203457 | No | Replaced touch screen. | Replaced touch screen. | Defect | HW | DMS_PC | Touch screen
ABL90 FLEX PLUS | Fluid transport errors - resolved by customer actions, but unknown which action taken was the solution. Subsequent pCO2 errors - possibly needs new Sensor Cassette of a different lot. | Fluid transport errors - resolved by customer actions, but unknown which action taken was the solution. Subsequent pCO2 errors - possibly needs new Sensor Cassette of a different lot. | SWO-203484 | No | Fluid transport errors - resolved by customer actions, but unknown which action taken was the solution. Subsequent pCO2 errors - possibly needs new Sensor Cassette of a different lot. | Fluid transport errors - resolved by customer actions, but unknown which action taken was the solution. Subsequent pCO2 errors - possibly needs new Sensor Cassette of a different lot. | No fault found | SC | SC | pCO2
ABL90 FLEX PLUS | Swapped phase 0 Analyser display unit with faulty ED Analyser. Touch screen panel needs to be replaced, due to liquid damage. Analyser has passed calibration and QC. | Swapped phase 0 Analyser display unit with faulty ED Analyser. Touch screen panel needs to be replaced, due to liquid damage. Analyser has passed calibration and QC. | SWO-203635 | No | Swapped phase 0 Analyser display unit with faulty ED Analyser. Touch screen panel needs to be replaced, due to liquid damage. Analyser has passed calibration and QC. | Swapped phase 0 Analyser display unit with faulty ED Analyser. Touch screen panel needs to be replaced, due to liquid damage. Analyser has passed calibration and QC. | Revisit required | HW | DMS_PC | Touch screen
ABL90 FLEX PLUS |  | B | D2BB16BF72086C9F | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB
AQT90 FLEX | 更换903-364后修复(by tim) | 更换903-364后修复(by tim) | SWO-202446 | Yes | Repair after replacement 903-364 (by tim) | Repair after replacement 903-364 (by tim) | Re classified RMED | OOB |  |  | 124
AQT90 FLEX | 重新定标正常。(by tim) | 重新定标正常。(by tim) | SWO-202851 | No | Recalibration is normal. (by tim) | Recalibration is normal. (by tim) | User error | Consumables | Test Kit | BNP Calibration
ABL827 FLEX | Removed auto check module cleaned dust from unit. Refitted and test ok. | user has blown into the scanner slot to try and remove dust on the lens, still misses some barcodes and the positions it misses change each time the analyser rescans. checked in service laser is on all the time. FSE visit required to resolve. | SWO-203494 | No | Removed auto check module cleaned dust from unit. Refitted and test ok. | user has blown into the scanner slot to try and remove dust on the lens, still misses some barcodes and the positions it misses change each time the analyser rescans. checked in service laser is on all the time. FSE visit required to resolve. | Cleaned | HW | Autocheck module | Barcode reader, Autocheck Module
ABL827 FLEX | Removed auto check module cleaned dust from unit. Refitted and test ok. | Removed auto check module cleaned dust from unit. Refitted and test ok. | SWO-203494 | No | Removed auto check module cleaned dust from unit. Refitted and test ok. | Removed auto check module cleaned dust from unit. Refitted and test ok. | Cleaned | HW | Autocheck module | Barcode reader, Autocheck Module
ABL80 FLEX | 更换924-738一个正常。(by tim) | 更换924-738一个正常。(by tim) | SWO-203144 | No | Replace 924-738 one normal. (by tim) | Replace 924-738 one normal. (by tim)
ABL90 FLEX PLUS |  | B | D2BE8784D42A1338 | No |  |  | Assisted Replacement | SC | SC | pCO2
ABL835 FLEX |  | B | SWO-203518 | No |  |  | Defect | HW | Basic Unit | Waste detector
ABL90 FLEX |  | B | SWO-203540 | No |  |  | Defect | HW | Basic Unit | PCB, Wet section Control
AQT90 FLEX | 更换903-364，重新定标正常。(by tim) | 更换903-364，重新定标正常。(by tim) | SWO-203248 | No | Replace the 903-364 and re-mark normally. (by tim) | Replace the 903-364 and re-mark normally. (by tim) | Defect | HW | Process Module | Optical Unit | 131
ABL80 FLEX BASIC |  | B | SWO-203765 | No
ABL90 FLEX | Thb Kal durchgeführt, Werte im Level 3 0,1 unter Oberem Grenzwerte, in Level 2 0,2 mmol L kpl. Oxy Test durchgeführt i.O. Hemolyzer Unit bestellt -> Info an Kunden | Thb Kal durchgeführt, Werte im Level 3 0,1 unter Oberem Grenzwerte, in Level 2 0,2 mmol L kpl. Oxy Test durchgeführt i.O. Hemolyzer Unit bestellt -> Info an Kunden | SWO-203859 | Yes | Thb Kal performed, values at level 3 0.1 below upper limit values, in level 2 0.2 mmol L Compl. Oxy test performed in the first place Hemolyzer Unit ordered -> Info to customers | Thb Kal performed, values at level 3 0.1 below upper limit values, in level 2 0.2 mmol L Compl. Oxy test performed in the first place Hemolyzer Unit ordered -> Info to customers | Re classified RMED | HW | Oxi module | Hemolyzer
ABL90 FLEX PLUS |  | B | D2C17F55C03B4976 | Yes
ABL90 FLEX PLUS | Schlauchsystem geprüft, Schlauch ausgetauscht siehe SWO Jahreswartung | Kd. bittet um Behebung der Störung durch einen KDB. | SWO-204395 | No | Hose system tested, hose replaced see SWO annual maintenance | Kd. asks for a correction of the fault by a KDB. | Defect | HW | Oxi module | Pinch valve tube
ABL90 FLEX PLUS | Replaced touch screen panel on the phase 0 Analyser and then swapped back to original displays. Both Analysers are ready for use. | Swapped phase 0 Analyser display unit with faulty ED Analyser. Touch screen panel needs to be replaced, due to liquid damage. Analyser has passed calibration and QC. | SWO-204501 | No | Replaced touch screen panel on the phase 0 Analyser and then swapped back to original displays. Both Analysers are ready for use. | Swapped phase 0 Analyser display unit with faulty ED Analyser. Touch screen panel needs to be replaced, due to liquid damage. Analyser has passed calibration and QC. | User error | HW | DMS_PC | Touch screen
ABL90 FLEX PLUS |  | B | D2BF0EF3D407053D | Yes
ABL90 FLEX PLUS |  | B | D2BF02A0D45C77AF | No |  |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
TCM4 |  | B | SWO-203710 | No
TCM4 |  | B | SWO-203713 | No
TCM4 |  | B | SWO-203709 | No
ABL90 FLEX | 更换三张测试卡正常。 | 更换三张测试卡正常。 | SWO-203966 | No | Replacing the three test cards is OK. | Replacing the three test cards is OK. | Assisted Replacement | SC | SC | Thermistor
ABL90 FLEX | 连续更换2张测试卡后修复 | 连续更换2张测试卡后修复 | SWO-203794 | Yes | Repair after 2 consecutive test cards | Repair after 2 consecutive test cards
AQT90 FLEX |  | B | SWO-203665 | No |  |  | Defect | HW | Aspiration Module | Needle Wash Unit | 1999
ABL815 FLEX | Überprüfung Dateiarchiv => immer wieder sporadisch Fehler 0367 pH BG Modul nicht zeitgerecht gefüllt und 0878 keine Spüllösung während Kal Überprüfung, Ausbau und Reinigung alles Messventile => Membran vom pH BG Airventil defekt => Austausch Mehrmals Kal ohne Abbrüche durchgeführt | Überprüfung Dateiarchiv => immer wieder sporadisch Fehler 0367 pH BG Modul nicht zeitgerecht gefüllt und 0878 keine Spüllösung während Kal Überprüfung, Ausbau und Reinigung alles Messventile => Membran vom pH BG Airventil defekt => Austausch Mehrmals Kal ohne Abbrüche durchgeführt | SWO-203825 | No | Checking file archive => again and again sporadically error 0367 pH BG module not filled on time and 0878 no flushing solution during Kal Inspection, removal and cleaning of all measuring valves => membrane of the pH BG air valve defective => Replacement Several times Kal performed without aborts | Checking file archive => again and again sporadically error 0367 pH BG module not filled on time and 0878 no flushing solution during Kal Inspection, removal and cleaning of all measuring valves => membrane of the pH BG air valve defective => Replacement Several times Kal performed without aborts | Defect | HW | pH/BG Module | Valve, air | Membrane
ABL805 FLEX | The analyzer keeps shutting itself off and, then it won't startup windows thus it would not load the AU software. | Replaced the power supply and, the CPU unit. I shutdown the analyzer several times and, the AU software loaded ok. We will monitor the instrument for any reoccurrence of the problem. | D2BEA516D44A4027 | No | The analyzer keeps shutting itself off and, then it won't startup windows thus it would not load the AU software. | Replaced the power supply and, the CPU unit. I shutdown the analyzer several times and, the AU software loaded ok. We will monitor the instrument for any reoccurrence of the problem. | Defect | HW | DMS_PC | Power Supply
ABL835 FLEX | Continued work from last visit same problem. | Replaced PH electrode. Performed all necessary calibrations. Old electrode under warranty: 945-614 LOT XR-01 | D2BF07D5D4121CC2 | No | Continued work from last visit same problem. | Replaced PH electrode. Performed all necessary calibrations. Old electrode under warranty: 945-614 LOT XR-01 | Assisted Replacement | Electrode | Electrode | pH
ABL827 FLEX | Inspected chamber and electrodes, all ok. Checked valves in service mode, checked gas flow in service, both ok. Ph chamber filling with fluid and not draining. Ph bg air valve stuck, unable to pass air through. Stripped and cleaned valve which was very dirty. | FSE visit required to investigate this. | SWO-203843 | No | Inspected chamber and electrodes, all ok. Checked valves in service mode, checked gas flow in service, both ok. Ph chamber filling with fluid and not draining. Ph bg air valve stuck, unable to pass air through. Stripped and cleaned valve which was very dirty. | FSE visit required to investigate this. | Cleaned | HW | pH/BG Module | Valve, air
ABL835 FLEX | Gerät überprüft. OXI Modul PCB erneuert, Fehler bleibt bestehen. Absprache mit Kundin getroffen,dass ein Ersatzteil ( OXI Modul ) bestellt wird. | OXI Modul zur Reparatur Überbrückung eingesetzt. | SWO-205189 | No | device checked. OXI module PCB renewed, error remains. agreement with customer,that a spare part ( OXI module ) is ordered. | OXI module used for repair bridging. | Revisit required | HW | Oxi module | Oxi module
ABL825 FLEX | checked data | replaced hemolyzer,ran, firmware update for screen. | D2BF0F5CD41CCB8A | No | checked data | replaced hemolyzer,ran, firmware update for screen. | Defect | HW | Oxi module | Hemolyzer
ABL825 FLEX | checked opertaion screen very slow | installed firmware update and permanent keys. | D2BF1213D42CCB91 | No | checked opertaion screen very slow | installed firmware update and permanent keys. | Defect | HW | DMS_PC | Touch screen | Firmware
AQT90 FLEX |  | B | D2BE96BCD42B8F5B | No |  |  | Defect | HW | Storage Module | PCB, Cartridge Lock | 167
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-204847 | No | Repair after replacing test card | Repair after replacing test card
ABL90 |  | B | D2BFEF374A5095CF | No |  |  | Defect | HW | Basic Unit | Barcode reader, display
ABL90 FLEX | Battery unit replaced. The LED indicator is indicating that the battery is charging and the battery percentage is raising. The battery values are updating in service. | Battery unit replaced. The LED indicator is indicating that the battery is charging and the battery percentage is raising. The battery values are updating in service. | SWO-204243 | No | Battery unit replaced. The LED indicator is indicating that the battery is charging and the battery percentage is raising. The battery values are updating in service. | Battery unit replaced. The LED indicator is indicating that the battery is charging and the battery percentage is raising. The battery values are updating in service. | Defect | HW | Basic Unit | Battery
ABL90 FLEX | Einlass demontiert und den Rest der Dichtung entfernt. Einlassnadel und Dichtung erneuert (aus Kundenbestand) und Gerät neu gestartet (Kaltstart). Funktionskontrolle: I.O. | Einlass demontiert und den Rest der Dichtung entfernt. Einlassnadel und Dichtung erneuert (aus Kundenbestand) und Gerät neu gestartet (Kaltstart). Funktionskontrolle: I.O. | SWO-204362 | No | inlet and remove scan the rest of the seal. Inlet needle and seal renewed (from customer stock) and device restarted (cold start). Functional control: I.O. | inlet and remove scan the rest of the seal. Inlet needle and seal renewed (from customer stock) and device restarted (cold start). Functional control: I.O. | Defect | HW | Inlet module | Inlet gasket | Broken
ABL90 | The analyzer had no power | I replaced the power supply and it came up. I upgraded the software to 3.3 | D2BFCD444A44C32B | No | The analyzer had no power | I replaced the power supply and it came up. I upgraded the software to 3.3 | Defect | HW | Power Supply | Power Supply
ABL90 | PO2 was failing external qc during 20 day qc study. Level 1 was low, Levels 2 and 3 were high. Suspect is silicone contamination or malfunctioning sensor cassettes. Customer is concerned about instrument performance and would like instrument replaced if problem is not resolved by FSE. | Performed protein removal, replaced all tubing using pm kit. Installed new sensor cassette and ran multiple waste bloods to hydrate sensor cassette. Installed new sensor cassette and solution pack. Verified instrument operation. | D2C014BB4A0A4B8B | No | PO2 was failing external qc during 20 day qc study. Level 1 was low, Levels 2 and 3 were high. Suspect is silicone contamination or malfunctioning sensor cassettes. Customer is concerned about instrument performance and would like instrument replaced if problem is not resolved by FSE. | Performed protein removal, replaced all tubing using pm kit. Installed new sensor cassette and ran multiple waste bloods to hydrate sensor cassette. Installed new sensor cassette and solution pack. Verified instrument operation. | Assisted Replacement | SC | SC | Wetability
ABL90 FLEX PLUS |  | B | D2C013714A61D74C | No |  |  | Defect | HW | Basic Unit | Detector plate
TCM5 BASIC | Changed sensor, still same Took logs and service dump to create TechLine case | Changed sensor, still same Took logs and service dump to create TechLine case | SWO-204247 | No | Changed sensor, still same Took logs and service dump to create TechLine case | Changed sensor, still same Took logs and service dump to create TechLine case
AQT90 FLEX |  | B | SWO-203946 | No |  |  | Defect | HW | Process Module | Optical Unit | 131
AQT90 FLEX |  | B | SWO-203944 | No |  |  | Defect | HW | Process Module | Optical Unit | 159
AQT90 FLEX | 更换944-216后修复 by jesson | 更换944-216后修复 by jesson | SWO-203557 | No | Repair by jesson after replacement 944-216 | Repair by jesson after replacement 944-216 | Cleaned | HW | Process Module | Process wheel | 1999
ABL825 FLEX | PH Elektrode (E777 XB-02) geprüft und ausgetauscht, Funktionskontrolle ok | PH Elektrode (E777 XB-02) geprüft und ausgetauscht, Funktionskontrolle ok | SWO-204393 | No | PH electrode (E777 XB-02) tested and replaced, function control ok | PH electrode (E777 XB-02) tested and replaced, function control ok | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Examined data logs. Examined pCO2 and Cl electrodes. | Replaced Cl, pCO2, pO2, and Ref membranes, with customer stock. Advised customer to replace pCO2 electrode. Verified operation of ABL825. | D2C043AE4A18BFDB | No | Examined data logs. Examined pCO2 and Cl electrodes. | Replaced Cl, pCO2, pO2, and Ref membranes, with customer stock. Advised customer to replace pCO2 electrode. Verified operation of ABL825. | Assisted Replacement | Electrode | Membrane | pCO2
ABL827 FLEX | This only started after a gas2 and cal solution change. Customer has replaced electrode. Tried new membranes and gas, both from other lot numbers. Replaced ref membrane. Chambers cleaned and inspected on previous visit, so swapped chambers BG for el met, also swapped over gas valve assembly. Replaced solutions. Replaced black tubing. After all this still the QC reading is almost identical to original error values. | Fitted new flow regulators for gas 1 & gas 2 as well as a new pressure transducer chassis. The analysers was left to start up and then the QCs run, for all levels the pCO2 & pO2 were back to the levels prior to the problem starting. The firmware for the touchscreen was also updated to solve the issue with the touchscreen becoming unresponsive. | SWO-204300 | No | This only started after a gas2 and cal solution change. Customer has replaced electrode. Tried new membranes and gas, both from other lot numbers. Replaced ref membrane. Chambers cleaned and inspected on previous visit, so swapped chambers BG for el met, also swapped over gas valve assembly. Replaced solutions. Replaced black tubing. After all this still the QC reading is almost identical to original error values. | Fitted new flow regulators for gas 1 & gas 2 as well as a new pressure transducer chassis. The analysers was left to start up and then the QCs run, for all levels the pCO2 & pO2 were back to the levels prior to the problem starting. The firmware for the touchscreen was also updated to solve the issue with the touchscreen becoming unresponsive. | Revisit required | HW | Gas module | Gas valve | 1 & 2
ABL835 FLEX | OXI Modul zur Reparatur Überbrückung eingesetzt. | OXI Modul zur Reparatur Überbrückung eingesetzt. | SWO-205186 | No | OXI module used for repair bridging. | OXI module used for repair bridging. | Defect | HW | Oxi module | Oxi module
ABL827 FLEX | Tried restarting the unit and it failed inlet calibration. Ran the inlet calibration from the service screen and failed. | Replaced the friction detector unit. Replaced the LAC membrane. | D2C0289A4A0B0269 | No | Tried restarting the unit and it failed inlet calibration. Ran the inlet calibration from the service screen and failed. | Replaced the friction detector unit. Replaced the LAC membrane. | Defect | HW | Inlet module | Friction detector unit
AQT90 FLEX |  | B | D2C054D04A42D0C5 | No |  |  | Cleaned | HW | Cover and Chassis | Inlet Door
ABL80 FLEX BASIC | Cl Parameter failed. At the time of re-installation V C fail msg coming. Found SC failed. SC replaced. QC run found all with in range. Sampling perform and found values not reliable as per clinical condition of Patient. Ask to conduct a Pre-analytical training. Dr given time for next day. Analyser serviced and calib rated properly. Handed over in a fine working condition. | Cl Parameter failed. At the time of re-installation V C fail msg coming. Found SC failed. SC replaced. QC run found all with in range. Sampling perform and found values not reliable as per clinical condition of Patient. Ask to conduct a Pre-analytical training. Dr given time for next day. Analyser serviced and calib rated properly. Handed over in a fine working condition. | SWO-205740 | No | Cl Parameter failed. At the time of re-installation V C fail msg coming. Found SC failed. SC replaced. QC run found all with in range. Sampling perform and found values not reliable as per clinical condition of Patient. Ask to conduct a Pre-analytical training. Dr given time for next day. Analyser serviced and calib rated properly. Handed over in a fine working condition. | Cl Parameter failed. At the time of re-installation V C fail msg coming. Found SC failed. SC replaced. QC run found all with in range. Sampling perform and found values not reliable as per clinical condition of Patient. Ask to conduct a Pre-analytical training. Dr given time for next day. Analyser serviced and calib rated properly. Handed over in a fine working condition.
ABL80 FLEX BASIC | 更换902-971后修复(by tim) | 更换902-971后修复(by tim) | SWO-204341 | No | Repair after replacement 902-971 (by tim) | Repair after replacement 902-971 (by tim)
ABL90 FLEX | Funktion im Serviceprogramm überprüft, Fließwege gespült, Einlassdichtung aus Kundenbestand ersetzt. Funktion z. Zt. OK. Dem Anwender empfohlen das LP zu erneuern, wenn das Problem nochmals auftritt. | Funktion im Serviceprogramm überprüft, Fließwege gespült, Einlassdichtung aus Kundenbestand ersetzt. Funktion z. Zt. OK. Dem Anwender empfohlen das LP zu erneuern, wenn das Problem nochmals auftritt. | SWO-204597 | No | Function checked in the service program, flow paths flushed, inlet seal replaced from customer stock. Function currently OK. The user is advised to renew the LP if the problem occurs again. | Function checked in the service program, flow paths flushed, inlet seal replaced from customer stock. Function currently OK. The user is advised to renew the LP if the problem occurs again. | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL90 FLEX | Meßbereich Ca von 0,4 auf 0,1 abgesenkt | Meßbereich Ca von 0,4 auf 0,1 abgesenkt | SWO-204628 | No | Measuring range Ca lowered from 0.4 to 0.1 | Measuring range Ca lowered from 0.4 to 0.1 | Setup | SW | Measuring Range | Dialysis mode (range)
ABL827 FLEX | New gas regulators have been ordered. Radiometer to return to complete the repair once the spares become available. | Fitted new flow regulators for gas 1 & gas 2 as well as a new pressure transducer chassis. The analysers was left to start up and then the QCs run, for all levels the pCO2 & pO2 were back to the levels prior to the problem starting. The firmware for the touchscreen was also updated to solve the issue with the touchscreen becoming unresponsive. | SWO-204737 | No | New gas regulators have been ordered. Radiometer to return to complete the repair once the spares become available. | Fitted new flow regulators for gas 1 & gas 2 as well as a new pressure transducer chassis. The analysers was left to start up and then the QCs run, for all levels the pCO2 & pO2 were back to the levels prior to the problem starting. The firmware for the touchscreen was also updated to solve the issue with the touchscreen becoming unresponsive. | Revisit required | HW | Gas module | Gas valve | 1 & 2
ABL825 FLEX | 0636, Gas 1 low and 0326, gas 1 empty, errors are present. Gas 1 has been replaced repeatedly as the customer indicated. | Replaced gas 1 regulator and verified gas flow. Observed several successful calibrations. Analyzer is within acceptable operating specifications. | D2C18DFEC01B2776 | No | 0636, Gas 1 low and 0326, gas 1 empty, errors are present. Gas 1 has been replaced repeatedly as the customer indicated. | Replaced gas 1 regulator and verified gas flow. Observed several successful calibrations. Analyzer is within acceptable operating specifications. | Defect | HW | Gas module | Gas valve | 636
ABL80 FLEX BASIC | 更换测试卡后修复 | 更换测试卡后修复 | SWO-204997 | No | Repair after replacing test card | Repair after replacing test card
ABL80 FLEX BASIC | 更换903-334后未修复，要重新借903-334和902-972各一个 by jesson | 更换902-972后修复 by jesson | SWO-204392 | No | Replace 903-334 after not repaired, to borrow 903-334 and 902-972 each by jesson | Repair by jesson after replacement 902-972
ABL80 CO-OX | checked data logs for errors, checked diagnostics, checked oxi pinch valve, | removed wet section and repositioned oxi tubing from outside valve to though the valve. | D2C310A236018F53 | No | checked data logs for errors, checked diagnostics, checked oxi pinch valve, | removed wet section and repositioned oxi tubing from outside valve to though the valve.
ABL90 FLEX PLUS |  | B | D2C275A636328D47 | No |  |  | Assisted Replacement | SP | SP | pO2
ABL90 FLEX PLUS | Inletmodule heb vervangen. | Inletmodule heb vervangen. | SWO-205216 | No | Have replaced the inletmodule. | Have replaced the inletmodule. | Defect | HW | Auto Inlet module | Auto Inlet module | Cap.
ABL90 FLEX PLUS |  | B | D2C29BAF3663D125 | Yes
ABL90 FLEX PLUS |  | B | D2C29B2D361915E7 | No |  |  | Defect | HW | Basic Unit | Sensor Interface PCB | Glu/Lac
ABL90 FLEX | 更换903-139备件后修复(By tim) | 更换903-139备件后修复(By tim) | SWO-204036 | No | Repair after replacement of 903-139 spare parts (By tim) | Repair after replacement of 903-139 spare parts (By tim) | Defect | HW | Inlet module | Inlet module
ABL825 FLEX | INTERMITTENT PROBLEM PROBABLY SPECTROMETER OR HEMOLYZER | REPLACED SPECTROMETWER AND HEMOLYZER, INSPECTED LAMP AND COMPLETED ALIGNMENT AND CALIBRATION | D2C28CEE36583C97 | No | INTERMITTENT PROBLEM PROBABLY SPECTROMETER OR HEMOLYZER | REPLACED SPECTROMETWER AND HEMOLYZER, INSPECTED LAMP AND COMPLETED ALIGNMENT AND CALIBRATION | Defect | HW | Oxi module | Hemolyzer
ABL827 FLEX | Fitted new flow regulators for gas 1 & gas 2 as well as a new pressure transducer chassis. The analysers was left to start up and then the QCs run, for all levels the pCO2 & pO2 were back to the levels prior to the problem starting. The firmware for the touchscreen was also updated to solve the issue with the touchscreen becoming unresponsive. | Fitted new flow regulators for gas 1 & gas 2 as well as a new pressure transducer chassis. The analysers was left to start up and then the QCs run, for all levels the pCO2 & pO2 were back to the levels prior to the problem starting. The firmware for the touchscreen was also updated to solve the issue with the touchscreen becoming unresponsive. | SWO-205104 | No | Fitted new flow regulators for gas 1 & gas 2 as well as a new pressure transducer chassis. The analysers was left to start up and then the QCs run, for all levels the pCO2 & pO2 were back to the levels prior to the problem starting. The firmware for the touchscreen was also updated to solve the issue with the touchscreen becoming unresponsive. | Fitted new flow regulators for gas 1 & gas 2 as well as a new pressure transducer chassis. The analysers was left to start up and then the QCs run, for all levels the pCO2 & pO2 were back to the levels prior to the problem starting. The firmware for the touchscreen was also updated to solve the issue with the touchscreen becoming unresponsive. | Defect | HW | Gas module | Gas valve | 1 & 2
ABL80 FLEX BASIC | 更换902-972和903-334后修复 by jesson | 更换902-972和903-334后修复 by jesson | SWO-204779 | No | Repair by jesson after replacing 902-972 and 903-334 | Repair by jesson after replacing 902-972 and 903-334
ABL80 FLEX BASIC | 更换902-972后修复 by jesson | 更换902-972后修复 by jesson | SWO-204999 | No | Repair by jesson after replacement 902-972 | Repair by jesson after replacement 902-972
ABL90 FLEX | Dem Kunden erklärt das es Probleme gibt. SC getauscht. Kunde bestellt neu. | Dem Kunden erklärt das es Probleme gibt. SC getauscht. Kunde bestellt neu. | SWO-205538 | Yes | Explains to the customer that there are problems. SC. Customer reorders. | Explains to the customer that there are problems. SC. Customer reorders. | Re classified RMED | Phone support
ABL90 FLEX PLUS | Sample Mixer Unit ersetzt, Gerät kalibriert und mit QC überprüft Defekt: R0946N005 | Sample Mixer Unit ersetzt, Gerät kalibriert und mit QC überprüft Defekt: R0946N005 | SWO-205662 | No | Sample Mixer Unit replaced, device calibrated and checked with QC Defect: R0946N005 | Sample Mixer Unit replaced, device calibrated and checked with QC Defect: R0946N005 | Defect | HW | Basic Unit | Sampler Mixer
ABL90 FLEX PLUS | Mit Herrn Samuel und Herrn Kozelj Rücksprache gehalten und den Analysator richtig an Radiance angebunden. ABL nun mit Radiance verbunden. | Mit Herrn Samuel und Herrn Kozelj Rücksprache gehalten und den Analysator richtig an Radiance angebunden. ABL nun mit Radiance verbunden. | SWO-205636 | No | Consultation with Mr Samuel and Mr Kozelj and the analyzer properly connected to Radiance. ABL is now connected to Radiance. | Consultation with Mr Samuel and Mr Kozelj and the analyzer properly connected to Radiance. ABL is now connected to Radiance. | Setup | SW | His/lis | Radiance
ABL90 FLEX PLUS | Pfl. Simon und Pfl. Christoph nicht im Dienst Beim Eintreffen Kal und alle QC Level fehlerhaft bei Crea Urea => LP Wechsel (von Lot YN-01 auf YN-01) von gestern brachte keine Besserung Reinstallierung SC ServiceDump und DataLogs erstellt => Daten und Info an Ingo Später nochmal Rückruf auf Station durch RenRo | Pfl. Simon und Pfl. Christoph nicht im Dienst Beim Eintreffen Kal und alle QC Level fehlerhaft bei Crea Urea => LP Wechsel (von Lot YN-01 auf YN-01) von gestern brachte keine Besserung Reinstallierung SC ServiceDump und DataLogs erstellt => Daten und Info an Ingo Später nochmal Rückruf auf Station durch RenRo | SWO-205505 | No | Pfl. Simon and Pfl. Christoph not on duty On arrival Kal and all QC levels faulty at Crea Urea => LP change (from Lot YN-01 to YN-01) from yesterday brought no improvement Reinstall SC ServiceDump and DataLogs created => Data and info to Ingo Later callback to station by RenRo | Pfl. Simon and Pfl. Christoph not on duty On arrival Kal and all QC levels faulty at Crea Urea => LP change (from Lot YN-01 to YN-01) from yesterday brought no improvement Reinstall SC ServiceDump and DataLogs created => Data and info to Ingo Later callback to station by RenRo | Assisted Replacement | SC | SC | Crea/Urea
TCM4 | 更换备用电池431-018C后修复 by jesson | 更换备用电池431-018C后修复 by jesson | SWO-204898 | No | Repair by jesson after replacing backup battery 431-018C | Repair by jesson after replacing backup battery 431-018C
ABL90 | Chequeo completo de equipo y comprobación de registro de datos. Sustitución de SP y SC Calibraciones y pasado qc | Sustitución SP y SC | D2C42138AC531834 | Yes | Complete computer check and data logging check. Replacing SP and SC Calibrations and past qc | SP and SC substitution
ABL90 FLEX | 更换910-285后修复 by jesson | 更换910-285后修复 by jesson | SWO-204475 | No | Repair by after replacement 910-285 jesson | Repair by after replacement 910-285 jesson | Defect | HW | Power Supply | Power Supply
ABL90 FLEX | 液体传感器错误，更换试剂包，冲洗管路后，暂时解决。未更换配件 | 液体传感器错误，更换试剂包，冲洗管路后，暂时解决。未更换配件 | SWO-205368 | No | Liquid sensor error, replace the reagent package, rinse the line, temporarily solve. Unreplaced accessories | Liquid sensor error, replace the reagent package, rinse the line, temporarily solve. Unreplaced accessories | Revisit required | HW | Basic Unit | Cassette retainer
ABL90 | Limpieza del circuito neumatico. Cambio de inlet gasket. Realizacion del test de servicio. Cal, qc. | B | D2C40878AC499E0B | No | Cleaning the tire circuit. Change of inlet gasket. Performing the service test. Cal, qc. |  | Cleaned | Fluidic pathway | Fluidic pathway | Fluidic pathway
ABL835 FLEX | pH電極交換。pCO2電極交換。（手持ち） | pH電極交換。pCO2電極交換。（手持ち） | SWO-205499 | No | pH electrode replacement. PCO2 electrode replacement. (Hand held) | pH electrode replacement. PCO2 electrode replacement. (Hand held) | Assisted Replacement | Electrode | Electrode | pH
ABL825 FLEX | Changed the oxi module PCB, power supply, oxi module completed. | Replaced the oxi module complete and ran tHb calibration, calibration and all passed. QC was out. Ran tHb cal in service followed by 2 point calibration and all passed including QC. | D2C4B38822498B2F | No | Changed the Oxygen module PCB, power supply, oxygen module completed. | Replaced the oxi module complete and ran tHb calibration, calibration and all passed. QC was out. Ran tHb cal in service followed by 2 point calibration and all passed including QC. | Defect | HW | Oxi module | PCB, Oxi module
ABL825 FLEX | モジュール返却し代替えモジュール引き上げ コンプリートテスト・ｔHbCalOK | B | SWO-205423 | No | Module return and replacement module pull Complete Test tHbCalOK |  | Defect | HW | Oxi module | Hemolyzer
ABL800 FLEX | USB stick ready but found BIOS setting not to be update as per requirement. Inform to Manager for further investigation and updates. Touch calibration done. Analyzer calibrated and found parameters well with in range. External Key board and Mouse connected and handed over in fine working condition. | USB stick ready but found BIOS setting not to be update as per requirement. Inform to Manager for further investigation and updates. Touch calibration done. Analyzer calibrated and found parameters well with in range. External Key board and Mouse connected and handed over in fine working condition. | SWO-205745 | No | USB stick ready but found BIOS setting not to be update as per requirement. Inform to Manager for further investigation and updates. Touch calibration done. Analyzer calibrated and found parameters well with in range. External Key board and Mouse connected and handed over in fine working condition. | USB stick ready but found BIOS setting not to be update as per requirement. Inform to Manager for further investigation and updates. Touch calibration done. Analyzer calibrated and found parameters well with in range. External Key board and Mouse connected and handed over in fine working condition. | Defect | HW | DMS_PC | Touch screen | Firmware
ABL825 FLEX | Mit Herrn Samuel und Herrn Kozelj Rücksprache gehalten und den Analysator richtig an Radiance angebunden. ABL nun mit Radiance verbunden. | Mit Herrn Samuel und Herrn Kozelj Rücksprache gehalten und den Analysator richtig an Radiance angebunden. ABL nun mit Radiance verbunden. | SWO-205637 | No | Consultation with Mr Samuel and Mr Kozelj and the analyzer properly connected to Radiance. ABL is now connected to Radiance. | Consultation with Mr Samuel and Mr Kozelj and the analyzer properly connected to Radiance. ABL is now connected to Radiance. | Setup | SW | Roma.ini | His/lis
AQT90 FLEX | checked and found scratches on bar code scanner . | checked and found scratches on bar code scanner . | SWO-205409 | No | checked and found scratches on bar code scanner . | checked and found scratches on bar code scanner . | Defect | HW | Storage Module | Barcode reader, Storage wheel
AQT90 FLEX | 更换903-364后修复 by jesson | 更换903-364后修复 by jesson | SWO-205511 | No | Repair by jesson after replacing 903-364 | Repair by jesson after replacing 903-364
AQT90 FLEX | 更换903-364后修复 by jesson | 更换903-364后修复 by jesson | SWO-205498 | No | Repair by jesson after replacing 903-364 | Repair by jesson after replacing 903-364
AQT90 FLEX | 更换903-030后修复 by jesson | 更换903-030后修复 by jesson | SWO-205448 | No | Repair by jesson after replacement 903-030 | Repair by jesson after replacement 903-030
AQT90 FLEX | 更换903-364备件后修复(By Ran) | 更换903-364备件后修复(By Ran) | SWO-205710 | No | Repair after replacement of 903-364 spare parts (By Ran) | Repair after replacement of 903-364 spare parts (By Ran)